## Foglio1
High-flow nasal cannula (HFNC)
Published on: 19 November 2020 - Version: 2
Version 1: September 2019 (WHO-UNICEF)
Technical Specification | Minimum requirements
General technical requirements | Capability to generate a high flow of mixed room air and oxygen. Capability to use oxygen from an oxygen concentrator or cylinder. In-built air compressor/turbine/piston or similar. Easy to operate user interface, with displayed parameters clearly visible. The mixed room air and oxygen should be warmed up to 37 °C and 100% RH. Controls to be easy to operate, numbers and displays to be clearly visible. It should have a humidity compensation system. Noise level < 35 dB at mid pressure range. Protection IP21 required (IP22 preferable).
Monitored and controlled parameters (by clinical user) | FiO2: 21–100% (preferable). Flow up to: 50 L/min (minimum).
Displayed parameters (colour and graphics preferable) | Display easily readable in low ambient light and sunlight. Gas temperature (°C). Flow (L/min). Air leak (%) (preferable). FiO2 (%) (preferable).
Alarms, related to gas delivered (visual and audible) | Incorrect temperature/humidity. System leakage or blockage. High/low FiO2 (preferable).
Alarms, related to equipment operation (visual and/or audible) | Lack of water. System failure. Air filter to be replaced. Power failure. Low battery (if applicable).
Consumables, labelled “single use” (included and mentioned in a disaggregated list) | Housing and patient interface for adult and paediatric: at least 30 per equipment (different sizes). Inlet bacteria filters, if applicable: 30 per equipment. Expiratory filters: high efficiency.
Accessories, reusable (included and mentioned in a disaggregated list) | Flowmeter, graduated in L/min, as required to operate. Humidifier, as required to operate. Water chamber, as required to operate. Standard hoses and connectors (i.e. DISS/NIST, barb, as applicable for the country) for oxygen wall outlets and cylinder, as required to operate. Mains power cable ≥ 2 m.
Spare parts (included and mentioned in a disaggregated list) | 1 year’s spare parts kit as per preventive maintenance programme
Portability | Mounting tray and support stand with at least 2 castors fitted with breaks.
Power supply (voltage, frequency and plug vary across the countries) | Operates from AC power electric line: 100–240 V AC ± 10% / 50–60 Hz ± 10%. In-built rechargeable battery (preferable); if the equipment does not have an internal battery, an external battery or uninterruptible power supply should be included to provide battery back up in the case of AC power failure. Automatic switch from AC power electric-line mode to battery operating mode and vice versa.
Documentation (included) | Instruction for use; service manual and product information to be provided in English language, at least.
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent, (including the technical tests for safety and performance from accredited laboratory or third party), for (latest version recommended but compliance to previous standards versions could be accepted): IEC 60601-1:2005 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-2:2014 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. ISO 18562-1:2017 Biocompatibility evaluation of breathing gas pathways in healthcare applications – Part 1: Evaluation and testing within a risk management process. ISO 17510:2015 Medical devices – Sleep apnoea breathing therapy – Masks and application accessories.
Warranty | Minimum 2 years. Availability of accessories, consumables and spare parts for at least 2 years.