## Foglio1
Syringe, luer, sterile, single use (Several sizes) (The specification detailed below is for general reference only and it refers to a "Syringe, 10 mL, three pieces, Luer type, sterile, single use")
Published on: 19 November 2020 - Version: 3
Version 2: 16/09/2013 (modification of the initial version)
Version 1: 25/06/2013 (initial version)
WHO Item name: | Required technical specifications: | Dimensions / Sizes / Presentation: | Material: | Certificates for product safety and quality assurance, as requested or equivalent
Syringe, 10 mL, three pieces, Luer type, sterile, single use | Three pieces syringe: barrel with Luer nozzle, piston, and stopper. Barrel: permanent and legible graduated scale in mL with intervals of 0.20 or 0.50 mL. Increment of each mL to be numbered. Concentric or eccentric Luer lock or Luer slip nozzle. Length with a maximum usable capacity of at least 10% more than the nominal capacity. Plunger stopper with backstop. Double sealing ring on plunger. Initial sterilization method: ethylene oxide gas or equivalent. Individually peel-packed in paper and/or plastic. | Barrel nominal capacity: 10 mL. | Barrel: polyethylene (PE) or polypropylene (PP) or polystyrene (PS); sufficiently transparent. Piston: polypropylene (PP) or equivalent. Latex-free/PVC-free. | Barrel graduation: ISO 7886-1:2017: Sterile hypodermic syringe for single use: part 1: syringe for manual use. Needle, Luer type: ISO 80369-7:2016: Small-bore connectors for liquids and gases in healthcare applications – Part 7: Connectors for intravascular or hypodermic applications. Sterilization method: ISO 11135:2014: Sterilization of health-care products – Ethylene oxide – Requirements for the development, validation and routine control of a sterilization process for medical devices.