## Foglio1
Patient monitor multiparametric, advanced
Published on: 19 November 2020 - Version: 3
Version 2: 18/06/2014 (modification of the initial version)
Version 1: 13/06/2012 (initial version)
Technical Specification | Minimum requirements
General technical requirements | Continuous display on screen of patient ECG, respiration rate and heart rate, invasive and non-invasive blood pressure, body temperature, SpO2 and CO2 mainstream or side stream. Dynamic digital display that can show all active parameters. Unwanted parameters can be deselected from display. Operator can set audio-visual alarm levels for low or high levels of each parameter independently. Operates from mains voltage or from internal rechargeable battery. ECG patient connectors that are sterilizable and reusable are preferred. Hard copy printout of traces will not be required. • Multichannel (up to 12 leads) ECG measurement and selectable display; an extra option for simple five-lead connection preferred. • Heart rate measurement range to be at least 30–250 bpm, with accuracy better than ± 5 bpm and minimum gradation 1 bpm. • SpO2 measurement range at least 70–99%, with accuracy better than ± 3% and minimum gradation 1%. • Blood pressure monitoring range at least 30–270 mmHg, minimum gradation 1 mmHg. Cuff sizes neonatal/paediatric and adult. User selectable measurement intervals. • Internal pump for cuff inflation for non-invasive blood pressure measurement, with over pressure protection. • Temperature probe to be reusable, external skin contact type, but consumable to protect between patients or disinfection method explained. • Temperature range at least 30–40 °C, minimum gradation 0.1 °C. • Respiration rate measurement range at least 0–100 bpm, minimum gradation 1 bpm. • CO2 monitoring capabilities. • Invasive blood pressure (IBP) monitoring capabilities. Automatic and programmable memory. Storage of continuous monitoring data. Trace signal velocity of at least 25 mm/s. LCD or TFT screen with: • analogue shape signals and numerical values visualization; • settable limits for the measured variables; • not less than 10”wide. Design must enable use in demanding environments, e.g. shock, vibration and free fall tests as per tests. Protections of all the functions against defibrillator discharges and electrosurgical units. Pace-maker detection. Data management functions (preferable). Capable of operating continuously in ambient temperature of 10–40 °C and relative humidity of 15–85% (90% preferable). Enclosure to have ingress protection level IPX1 or better.
Displayed parameters | Trend display of each parameter.
User adjustable settings | Alarm override and temporary silence facility to be included. Audio-visual alarms required: high and low levels for each parameter (operator variable settings), sensor/wire/probe disconnected, low battery.
Alarms | Alarm override and temporary silence facility to be included. Audio-visual alarms required: high and low levels for each parameter (operator variable settings), sensor/wire/probe disconnected, low battery, cuff leak, cuff disconnect, hose leak, inflation/deflation errors, failure to take successful reading, low battery notice. Power failure.
Consumables | ECG electrodes (if applicable): 100 sets.
Accessories | All the cables, sensors and connectors needed for full monitor functionality are to be included in the bid. Lead ECG cable: 2 per equipment. Lead ECG cable (if option offered): 2 per equipment. Sets of ECG connection electrodes (if reusable type): 5 sets. Tubes electrode gel (if required): 5 tubes. Reusable SpO2 probes adult: 2 probes. Reusable SpO2 probes paediatric use: 2 probes. Blood pressure – non-invasive: 3 paediatric reusable cuffs; 3 adult reusable cuffs. Blood pressure – invasive: 1 sensor for each channel offered. External skin temperature probes: 2 probes. If CO2 mainstream technology: tube adapter: 3 per equipment; sensor: 3 per equipment. If CO2 side stream technology: sample lines: 100 lines; water traps: 10 per equipment. Battery: 1 set.
Spare parts | 1-year spare parts kit as per preventive maintenance programme including but not exclusively, sets of spare fuses (if non-resettable fuses used) and battery.
Power supply, voltage, frequency and plug vary across the countries | Operated by line electrical power supply with internal replaceable rechargeable battery backup allows operation for at least 1 hour in the event of power failure. Operates from AC power electric line: 100–240 V~/50–60 Hz. Main power cable and plug adapted for various countries. Mains power cable length ≥ 2.5 m. Protections against over-voltage and over-current line conditions. Protection against defibrillator discharges and electrosurgical units. Automatic switch between battery and mains powered modes, when recharging or in mains power failure. The display shall show which power source is in use. Compliance with electrical standards and regulations.
Documentation requirements (English language mandatory) | Set: user and maintenance manuals, hard and soft copies, in English (mandatory) and other languages (preferable). Certificate of calibration and inspection. Troubleshooting, calibration and routine maintenance. List of all spare parts and accessories, with part numbers and contact details for parts supply. Document with contact details of manufacturer, supplier and local service agent.
Primary packaging label | Name and/or trademark of the manufacturer. Electrical power input requirements (voltage, frequency and socket type) and safety use and storage (keep away from oil, grease and petroleum-based or flammable products as well as smoking or open flames). Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: IEC 60601-1 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-1 Medical electrical equipment – Part 1-1: General requirements for safety – Collateral standard: Safety requirements for medical electrical systems. IEC 60601-1-2 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. IEC 60601-1-8 General requirements for basic safety and essential performance – Collateral standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems). IEC 80601-2-49 Medical electrical equipment – Part 2-49: Particular requirements for the basic safety and essential performance of multifunction patient monitoring equipment. IEC 80601-2-30 Particular requirements for the basic safety and essential performance of automated non-invasive sphygmomanometer. IEC 60601-2-34 Particular requirements for the basic safety and essential performance of invasive blood pressure monitoring equipment). ISO 80601-2-55 Particular requirements for the basic safety and essential performance of respiratory gas monitors). ISO 80601-2-61 Medical electrical equipment – Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment. IEC 60601-2-27 Particular requirements for the basic safety and essential performance of electrocardiographic monitoring equipment. Preferable if tested for: IEC 62366-1 Medical devices – Part 1: Application of usability engineering to medical devices. IEC 60068-1:2013 Environmental testing – Part 1: General and guidance. IEC 60068-2-31 Environmental testing – Part 2-31: Tests: Rough handling shocks, primarily for equipment-type specimens.
Warranty | 5 years with regards efficiency and quality of the product (software upgrades included). Availability of accessories and spare parts for at least 5 years.