## Foglio1
Mask and tubing for oxygen
Published on: 19 November 2020 - Version: 1
Technical Specification | Minimum requirements
General technical requirements | Non-rebreather mask with reservoir bag, used to deliver medical oxygen directly to the upper airway of the patient; single use. It includes two unidirectional valves, one that closes during inspiration to prevent room air mixing with oxygen in a reservoir bag; and one that closes during exhalation to prevent exhaled respiratory gases from entering the reservoir bag (non-rebreathing oxygen face mask). Mask is soft, transparent, well-fitting moulded, with two side vents. The nose clip is soft, malleable and adjustable. The tubing (oxygen line) is non-kinking, well-fitted. Tubing compatibility with standard oxygen connection tubing, 3–5 mm internal diameter and 7–8 mm external diameter, and 15/22 mm diameter ventilation tubing; available with “standard” and “universal” hose end connector. Individually packed. Non-sterile. Box of 50 or 100 units.
Sizes | Adult. Paediatric: tube length: 1.5–2 m.
Material | Mask and tubing PVC or other material, FDA Title 21/USP VI compliant and certified for medical use, hardness > 60 Shore A (ASTM D-2240).
Primary packaging label | Single use. Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: ISO 11712:2009 Anaesthetic and respiratory equipment – Supralaryngeal airways and connectors. ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen. ISO 18562 Biocompatibility evaluation of breathing gas pathways in healthcare applications. ISO 18190 Anaesthetic and respiratory equipment – General requirements for airways and related equipment. ISO 18562-1 Biocompatibility evaluation of breathing gas pathways in healthcare applications Part 1: Evaluation and testing within a risk management process. ISO/DIS 23368 Anaesthetic and respiratory equipment – Low flow nasal cannula for oxygen therapy. ISO/DIS 17256 Anaesthetic and respiratory equipment – Respiratory therapy tubing and connectors. ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements.