## Foglio1
Syringe, feeding, catheter tip, 50ml, sterile, single use (The specification detailed below is for general reference only and it refers to a "Syringe, feeding, ENFit, sterile, single use, different sizes")
Published on: 19 November 2020 - Version: 1
WHO Item name: | Required technical specifications: | Dimensions / Sizes / Presentation: | Material: | Certificates for product safety and quality assurance, as requested or equivalent
Syringe, feeding, 10 mL, ENFit, sterile, single use | Syringe used for the administration of oral medication or delivery of enteral nutrition by connection to an enteral administration device. It consists of a calibrated hollow barrel (cylinder) made of plastic and a moveable plunger. The tip is designed to mate only with enteral administration devices and is specifically incompatible with Luer (6%) connectors and intravenous devices. Initial sterilization method: ethylene oxide gas. Individually peel-packed in paper and/or plastic. | Barrel: 10 mL; graduations every 0.5 mL. | Barrel: transparent polypropylene, ending with a female ENFit tip. Plunger: polypropylene or polyethylene coloured to be distinguished from injection syringes. Gasket: synthetic elastomer (latex free). | ISO 80369-3:2016: Small-bore connectors for liquids and gases in healthcare applications – Part 3: Connectors for enteral applications. ISO 18250-3:2018: Medical devices – Connectors for reservoir delivery systems for healthcare applications – Part 3: Enteral applications.
Syringe, feeding, 20 mL, ENFit, sterile, single use | Syringe used for the administration of oral medication or delivery of enteral nutrition by connection to an enteral administration device. It consists of a calibrated hollow barrel (cylinder) made of plastic and a moveable plunger. The tip is designed to mate only with enteral administration devices and is specifically incompatible with Luer (6%) connectors and intravenous devices. Initial sterilization method: ethylene oxide gas. Individually peel-packed in paper and/or plastic. | Barrel: 20 mL; graduations every 0.5 mL. | Barrel: transparent polypropylene, ending with a female ENFit tip. Plunger: polypropylene or polyethylene coloured to be distinguished from injection syringes. Gasket: synthetic elastomer (latex free). | ISO 80369-3:2016: Small-bore connectors for liquids and gases in healthcare applications – Part 3: Connectors for enteral applications. ISO 18250-3:2018: Medical devices – Connectors for reservoir delivery systems for healthcare applications – Part 3: Enteral applications.
Syringe, feeding, 60 mL, ENFit, sterile, single use | Syringe used for the administration of oral medication or delivery of enteral nutrition by connection to an enteral administration device. It consists of a calibrated hollow barrel (cylinder) made of plastic and a moveable plunger. The tip is designed to mate only with enteral administration devices and is specifically incompatible with Luer (6%) connectors and intravenous devices. Initial sterilization method: ethylene oxide gas. Individually peel-packed in paper and/or plastic. | Barrel: 60 mL; graduations every 1 mL. | Barrel: transparent polypropylene, ending with a female ENFit tip. Plunger: polypropylene or polyethylene coloured to be distinguished from injection syringes. Gasket: synthetic elastomer (latex free). | ISO 80369-3:2016: Small-bore connectors for liquids and gases in healthcare applications – Part 3: Connectors for enteral applications. ISO 18250-3:2018: Medical devices – Connectors for reservoir delivery systems for healthcare applications – Part 3: Enteral applications.