## Oxygen analyzer, ultrasonic
Oxygen analyzer, electrochemical (galvanic cell), handheld
Published on:
Version 3:
Version 2:
Version 1:
Technical Specification | The oxygen analyser measures the oxygen concentration in a flow of gas from a medical gas source or, with adapters, through a medical gas flow device such as a ventilator or anaesthesia system, or within an environment such as oxygen hood and infant incubator.}Handheld oxygen analyser, displays the oxygen concentration measured in the flow of gas or the environment. Electrochemical cell oxygen sensing technology. Some products or versions are suitable for pressure and flow measurements. Operates from commercial batteries.ç
General technical requirements | Handheld oxygen analyser for spot check and/or continuous measurement of the oxygen concentration from a medical gas source and in an environment (depending on the configuration or version of the analyser). Galvanic fuel cell (electro-chemical) oxygen sensing technology. Supplied with connectors and/or adapters suitable for measurement of various medical gas supply sources, for example (but not limited to) oxygen concentrators, ventilators/anaesthesia machines and patient circuits (T-piece and/or in-line adapters), wall/column/cylinder supplies (compliance with ISO 7396-1). Oxygen measurement to include the range: 15–99%. Oxygen resolution: 0.1%. Oxygen accuracy: within ± 3%. Suitable for measuring gas supply with pressure up to 345 kPa (3.5 bar, 50 psi). Performance and calbration requirements at different pressures to be stated. Response time ≤ 20 s. Warm-up time < 10 s. Replaceable galvanic fuel cell (oxygen sensor), nominal operating life ≥ 1.5 years or 600 000 %O2-hours, whichever is greater. Calibration and self-test mode, two point calibration at ambient and 100% oxygen concentration. Internal calibration timer, with reminder (alarm and/or display message). . . Automatic power-off when not in use. Enclosure to have ingress protection level IPX1 or better.
Displayed parameters | Display visualizing oxygen concentration, system messages and battery status
User adjustable settings | N/A
Alarms | Low and high oxygen concentration audible and visual alarms required
Accessories (included and mentioned in a disaggregated list) | Carry case. Adapters for measuring various medical gas supply sources/ambient air (if applicable, depending on the models). Adapters for all available standards for fittings, including T-pieces and/or in-line adapters for various types and sizes of breathing circuits and adapters for central supply systems and cylinders.
Sterilization process for accessories (if relevant) | Suitable for cleaning and disinfection.
Consumables / reagents (if relevant) | Oxygen galvanic fuel cell
Spare parts (included and mentioned in a disaggregated list) | Batteries (disposable and/or rechargeable). Oxygen cell/sensor, sample line (if applicable). Sample line (if applicable) Set of spare fuses (if non-resettable fuses are used), display, connectors, battery holder, control panel, casing, battery charger.
Mobility, portability | Portable
Power supply, voltage, frequency and plug vary across countries | For the battery charger (if applicable). External or built-in AC battery charger, if rechargeable type. Charger, if used, to have protection against over-voltage and over-current line conditions, and be certified to IEC 60601-1. The requirements for power input voltage/frequency and plug type of the equipment must be chosen according to the local electrical supply. Depending on the local electrical supply availability and quality, voltage corrector/stabilizer/UPS can be recommended, in order to allow operation at ± 30% of local rated voltage, providing also protection for over current events. Electrical protection by resettable circuit breakers in both live and neutral supply lines.
Documentation requirements (English language mandatory) | User and maintenance manuals, hard and soft copies, to be supplied in local language, or agreed other language if local not available. English version must also be available. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance. List to be provided of common spares and accessories, with part numbers. Contact details of manufacturer, supplier and local service agent to be provided.
Environmental requirements | Capable of being stored in ambient temperature of at least 5–50 °C, relative humidity of at least 15–95% non condensing. Suitable for continuous operation in ambient temperature of at least 5–45 °C, relative humidity of at least 15–90% non condensing. Specific requirements for altitude may be required, depending on the installation site.
Primary packaging label | Primary packaging: Unit of use: one (1) oxygen analyser in a box or case or bag with manufacturer’s instruction for use, spare parts and accessories (when applicable). Labelling on the primary packaging: Name and/or trademark of the manufacturer; manufacturer’s product reference; type of product and main characteristics; lot number prefixed by the word “LOT” (or equivalent harmonized symbol) (if applicable); information for particular storage conditions(temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonized symbol); information for handling, if applicable (or equivalent harmonized symbol). Over packaging: Packaging unit. Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: number of units.
Standards, for the manufacturer | ISO 13485 Medical devices – Quality management systems – Requirements for regulatory purposes or ISO 9001 Quality management systems – Requirements. ISO 14971 Medical devices – Application of risk management to medical devices
Regulatory approval/ certification | Free sales certificate (FSC). Certified quality management system for medical devices (e.g. ISO 13485, or good manufacturing practice (GMP)). Certificate for exportation of medical device provided by the authority in the manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen. IEC 62133 – Secondary cells and batteries containing alkaline or other non-acid electrolytes – Safety requirements for portable sealed secondary cells. Part 1: Nickel, Part 2: Lithium.
Estimated Life Span | Not less than 5 years.
Shelf life (if relevant) | > 3 years for the oxygen galvanic fuel cell.
Spare parts availability post-warranty | 8 years at least, starting from the installation.
Warranty | 2 years recommended (1 year at least), the product shall be in production and fully supported when procured.
Maintenance tasks | Preventive maintenance, functionality tests and calibration as per manufacturer's specifications and at least once per year.
Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided upon request. Training of technical staff in advanced maintenance tasks shall be provided uopn request.