## Introduction
Technical specifications for selection of essential in vitro diagnostics (WHO EDL)
Introduction
Technical specifications enable procurement officers and agencies to adopt a systematic, unbiased approach to the selection of medical products, resulting in procurement, leasing or donation of appropriate commodities in line with the requirements of each level of the health-care system and local setting. Technical specifications for in vitro diagnostics (IVDs) constitute a set of criteria and baseline requirements to ensure good quality, safety, performance and efficacy. They are a decisive companion to the WHO Model list of Essential In Vitro Diagnostics (WHO EDL) that will allow Member States, donor agencies and nongovernmental organizations to select products within each test category of the WHO-EDL and will guide procurement decisions. The present publication defines the basic generic technical characteristics of IVDs listed in the WHO EDL. WHO will review and update this document regularly.
Objective
To provide information to support decisions on procurement of good-quality, safe in vitro diagnostics to improve in vitro diagnostic testing.
Target audience
This tool is intended for programme managers, laboratory managers, procurement officers and planning officers in ministries of health. It will also inform biomedical engineering professionals, private health sector, medical device industry and intergovernmental and international agencies about the characteristics of the products required.
Methods
The method used to prepare the technical specifications consisted of a review of existing WHO guidance and specifications for the selection, procurement, donation or lease of good-quality IVDs and medical devices in order to prepare a common template of generic specifications. The guidance documents analysed were: Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (1), WHO Guidelines on hepatitis B and C testing (2), the Third WHO Model List of Essential In Vitro Diagnostics (3) and WHO internal procurement templates.
In the second phase of the project, WHO commissioned an external team of consultants to develop technical specifications for the IVDs listed in the 3rd WHO-EDL. The process included: a market survey and a horizon scan of existing tests and platforms; analysis of products in the market after approval by regulatory agencies; analysis of international, regional and country standards; review of WHO guidance and guidelines, review of peer-reviewed publications, textbooks and published qualitative and quantitative comparisons of tests and platforms; comparison of test specifications (i.e. accuracy, precision, positive and negative predictive value, likelihood ratios) from catalogues and previous tenders by United Nations agencies with generic formulations; and data from external quality assessments and technical expert opinion. Each template was drafted and reviewed by at least two qualified technical experts. The data presented represents a universal template, generally applicable to a broad range of analysers and methods. All the consultants provided declarations of conflicts of Interest, and no conflicts were found. The draft technical specifications were reviewed internally by experts before being posted on the WHO website and shared with relevant networks for public consultation. The comments received were reviewed and included as appropriate, and the technical specifications were finalized.
References
1. Guidance for procurement of in vitro diagnostics and related laboratory items and equipment, Geneva: World Health Organization; 2017 (https://apps.who.int/iris/handle/10665/255577). 2. Guidelines on hepatitis B and C testing. Geneva: World Health Organization; 2017 (https://www.who.int/publications/i/item/9789241549981). 3. Third WHO Model List of Essential In Vitro Diagnostics, Geneva:World Health Organization; 2021 (https://www.who.int/publications/i/item/9789240019102).

## Content
Technical specifications for selection of essential in vitro diagnostics (WHO EDL)
The present tool includes technical specifications for 13 in vitro diagnostic tests for HCV listed in the WHO EDL
Each technical specification is presented in a standard template, which is organized as follows: name of the IVD test or laboratory equipment at the top, and 88 characteristics of technical specifications organized into 13 categories: Name, category and coding; Intended use; Performance characteristics; Technical and operational characteristics; Instrument physical and technical characteristics; Infrastructure requirements; Accessories, consumables, spare parts and other components; Documentation; Environmental and safety requirements; Training, installation and use; Warranty and maintenance; Decommissioning; Quality and registration.
To review technical specifications for a specific IVD category, go to the test menu tab, and click on the name of the relevant diagnostic test category. The selected item will be displayed on the screen. Scroll down to review the basic generic technical characteristics of the item.

## Tests menu
Technical specifications for selection of essential in vitro diagnostics: WHO EDL
No. | Page in the TRS 1031/EDL 3 | GMDN code/Term name | EMDN code/Term name | Diagnostic test name | Test purpose | Assay format
1 | P282 for no laboratory | 30829/ Hepatitis C virus total antibody IVD, kit, rapid ICT, clinical | W0105090202/ HCV ANTIBODY -RT & POC | Antibodies to hepatitis C virus (anti-HCV) | To screen for or to aid in the diagnosis of viraemic HCV infection: children, adolescents and adults | Rapid diagnostic test (RDT)
2 | P327 for clinical laboratory | 30829/ Hepatitis C virus total antibody IVD, kit, rapid ICT, clinical | W0105090202/ HCV ANTIBODY -RT & POC | Antibodies to hepatitis C virus (anti-HCV) | To screen for HCV infection and to aid in the diagnosis of viraemic HCV infection: children, adolescents and adults | Rapid diagnostic test (RDT)
3 | P327 for clinical laboratory | 48365/ Hepatitis C virus total antibody IVD, kit, enzyme immunoassay (EIA) | W0105020303/ HCV ANTIBODY (TOTAL) | Antibodies to hepatitis C virus (anti-HCV) | To screen for HCV infection and to aid in the diagnosis of viraemic HCV infection: children, adolescents and adults | Immunoassay Manual method
4 (Auto) | P327 for clinical laboratory | 48365/ Hepatitis C virus total antibody IVD, kit, enzyme immunoassay (EIA) | W0105020303/ HCV ANTIBODY (TOTAL) | Antibodies to hepatitis C virus (anti-HCV) | To screen for HCV infection and to aid in the diagnosis of viraemic HCV infection: children, adolescents and adults | Immunoassay Automated method
5 | P327 for clinical laboratory | 48381/ Hepatitis C virus antibody/antigen IVD, kit, enzyme immunoassay (EIA) | W0105020308/ HEPATITIS C VIRUS (HCV) ANTIGEN/ANTIBODY COMBO TEST | Combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg) | To screen for HCV infection and to aid in the diagnosis of viraemic HCV infection: children, adolescents and adults | Immunoassay Manual method
6 (Auto) | P327 for clinical laboratory | 48354/ Hepatitis C virus core antigen IVD, kit, enzyme immunoassay (EIA) | W0105020302/ HCV ANTIGEN - CONFIRMATORY | HCV core antigen (HCVcAg) | To aid in the diagnosis of viraemic HCV infection | Immunoassay Automated method
7 | P327 for clinical laboratory | 48374/ Hepatitis C virus nucleic acid IVD, kit, nucleic acid technique (NAT) | W0105020307/ HEPATITIS C VIRUS - NA REAGENTS | Qualitative or quantitative HCV nucleic acid test | To diagnose viraemic HCV and as a test of cure | Nucleic acid test (NAT) Manual method
8 (Semi auto) | P327 for clinical laboratory | 48374/ Hepatitis C virus nucleic acid IVD, kit, nucleic acid technique (NAT) | W0105020307/ HEPATITIS C VIRUS - NA REAGENTS | Qualitative or quantitative HCV nucleic acid test | To diagnose viraemic HCV and as a test of cure | Nucleic acid test (NAT) Semi-automated method
9 (Fully auto) | P327 for clinical laboratory | 48374/ Hepatitis C virus nucleic acid IVD, kit, nucleic acid technique (NAT) | W0105020307/ HEPATITIS C VIRUS - NA REAGENTS | Qualitative or quantitative HCV nucleic acid test | To diagnose viraemic HCV and as a test of cure | Nucleic acid test (NAT) Fully automated method
10 | P347 for blood-screening laboratory | 30829/ Hepatitis C virus total antibody IVD, kit, rapid ICT, clinical | W0105090202/ HCV ANTIBODY -RT & POC | Antibodies to hepatitis C virus (anti-HCV) | To screen donated blood for HCV | Rapid diagnostic test (RDT)
11 | P347 for blood-screening laboratory | 48365/ Hepatitis C virus total antibody IVD, kit, enzyme immunoassay (EIA) | W0105020303/ HCV ANTIBODY (TOTAL) | Antibodies to hepatitis C virus (anti-HCV) | To screen donated blood for HCV | Immunoassay Manual method
12 (Auto) | P347 for blood-screening laboratory | 48365/ Hepatitis C virus total antibody IVD, kit, enzyme immunoassay (EIA) | W0105020303/ HCV ANTIBODY (TOTAL) | Antibodies to hepatitis C virus (anti-HCV) | To screen donated blood for HCV | Immunoassay Automated method
13 | P347 for blood-screening laboratory | 48381/ Hepatitis C virus antibody/antigen IVD, kit, enzyme immunoassay (EIA) | W0105020308/ HEPATITIS C VIRUS (HCV) ANTIGEN/ANTIBODY COMBO TEST | Combined antibodies to HCV (anti-HCV) and HCV core antigen (HCV cAg) | To screen donated blood for HCV | Immunoassay Manual method
Excel cells coloured in ivory are for disease-specific IVDs for use in community settings and health facilities without laboratories.
Excel cells coloured in green are for disease-specific IVDs for use in clinical laboratories.
Excel cells coloured in pink are for disease-specific IVDs for blood-screening laboratories.

## 1
Antibodies to hepatitis C virus (anti-HCV)
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category/MTE code
2 | GMDN code/ Term name | 30829/ Hepatitis C virus total antibody IVD, kit, rapid ICT, clinical
3 | EMDN code/ Term name | W0105090202/ HCV ANTIBODY -RT & POC
4 | Generic name | Antibodies to hepatitis C virus (anti-HCV) rapid diagnostic test
5 | Specific type or variation (optional) | Qualitative, immunochromatographic assay
6 | Alternative name/s (optional) | Rapid diagnostic test (RDT) for antibodies to hepatitis C virus (anti-HCV)
7 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV antibody, hepatitis, rapid diagnostic test, RDT, HCV
Intended use
8 | Detection target | Hepatitis C virus antibodies
9 | Test purpose | To screen for or to aid in the diagnosis of viraemic HCV infection: children, adolescents and adults
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Assay format | Rapid diagnostic test (RDT)
12 | Specimen type | Oral fluid, capillary whole blood, venous whole blood
13 | Testing population | Children, adolescents and adults
14 | Intended user | Trained health-care providers, trained lay and peer workers
15 | Level of the health-care system | For use in communities and health facilities without laboratories
Performance characteristics
16 | Clinical sensitivity | > 99%
17 | Clinical specificity | > 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Limit of detection | Currently, there are no international units or international reference materials for measuring HCV antibodies, unlike for those of HBV. Two types of unit are used to express measurement results: signal-to-cut-off (s/co) ratios and the cut-off index (COI). The s/co ratio is obtained by dividing the signal strength for optical density of a sample by the signal strength for optical density of an internal cut off. As the s/co ratio is a ratio to the cut off and does not reflect the actual amount of antibody. In many test kits in which s/co ratios are used, the s/co is set between 0.75 and 1.1. The COI is obtained by converting the cut-off value (threshold value) obtained to 1.0 and indexing it. The COI standardizes measurements. As the COI is a ratio to the positive standard, it does not reflect the actual amount of antibody. In many test kits in which COI is used, it is set to 1.0.
20 | Rate of invalid, error and unreturnable results
21 | Precision (repeatability/reproducibility) | Not applicable
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Lateral flow immunoassay. Can detect antibodies to HCV genotypes 1 to 6.
24 | Specimen(s) stability | Venous whole blood can be stored at 2–8 °C for 2 days.
25 | Specimen(s) volume | 5–100 µL
26 | Type of result | Qualitative (reactive or non-reactive)
27 | Time to result | ≤ 20 min
28 | End-point stability | ≤ 20 min
29 | Ease of use for point-of-care tests only: number of steps that require precision | ≤ 3 steps
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 3 specimens per hour per operator
31 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
32 | Internal quality control | A built-in quality control line that must appear when a specimen is added plays a role of Internal quality control. Absence of the control line invalidates the test. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected from humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected from humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected from humidity and direct sunlight. Any specific requirement for reagents once they are opened should be clearly stated.
38 | Transport stability for specimen collection(temperature and humidity) | Venous whole blood can be stored at 2–8 °C for 2 days.
39 | Storage stability for specimen collection (temperature and humidity) | Venous whole blood can be stored at 2–8 °C for 2 days.
40 | In-use stability for specimen collection (temperature and humidity) | Venous whole blood can be stored at 2–8 °C for 2 days.
41 | Transport stability for controls/calibrators (temperature and humidity) | Not applicable
42 | Storage stability for controls/calibrators (temperature and humidity) | Not applicable
43 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability for controls/calibrators (temperature and humidity) | Not applicable
45 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months from the date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time to charge battery | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Displayed parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability/availability of a test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Not required
62 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature will require electricity.
63 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits.
64 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
65 | Kit components – test, reagents, consumables (if relevant) | Test strips or devices or cassette and/or buffer/diluent
66 | Reagent/kit size (number of tests) | One test or more
67 | Consumables required but not provided in the test kit | For collecting finger-prick or capillary whole blood, lancets are required. If capillary tubes or disposable pipettes are not provided, tips for micropipettes are required. For venous whole blood collection, blood collection tubes and needles are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if disposable pipettors for venous whole blood are not provided, tips for micropipette are required. Sodium hypochlorite (bleach) for sterilization before disposal of infectious waste is required.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection or for collecting finger-prick or capillary whole blood is required. For testing, a timer is required. If multiple tests are being performed more than 1 timer would be an advantage. If disposable pipettors are not provided, a calibrated micropipette and a micropipette stand are required.
69 | Spare parts (if relevant) | Not applicable
Docymentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Required. The instructions for use submitted must correspond to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. The certificate should be retained by the testing centre with the dates on which the lot was used.
73 | Material safety data sheet | Required. The technical document should provide detailed, comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | Potential source of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This information is usually provided in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of in-country technical support | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification status | This assay in a RDT format was within the scope of WHO Prequalification as of March 2023. There is a WHO-prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 2
Antibodies to hepatitis C virus (anti-HCV)
Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | GMDN code/ Term name | 30829/ Hepatitis C virus total antibody IVD, kit, rapid ICT, clinical
3 | EMDN code/ Term name | W0105090202/ HCV ANTIBODY -RT & POC
4 | Generic name | Antibodies to hepatitis C virus (anti-HCV) rapid diagnostic test
5 | Specific type or variation (optional) | Qualitative immunochromatographic assay
6 | Alternative name/s (optional) | Rapid diagnostic test (RDT) for antibodies to hepatitis C virus (anti-HCV)
7 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV antibody, hepatitis, rapid diagnostic test, HCV
Intended use
8 | Detection target | Hepatitis C virus antibodies
9 | Test purpose | To screen for or to aid in the diagnosis of viraemic HCV infection: children, adolescents and adults
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Assay format | Rapid diagnostic test (RDT)
12 | Specimen type | Capillary whole blood, venous whole blood, plasma and serum
13 | Testing population | Children, adolescents and adults
14 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
15 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
16 | Clinical sensitivity | > 99%
17 | Clinical specificity | > 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HCV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Limit of detection | Currently, there are no international units or international reference materials for measuring HCV antibodies, unlike for those of HBV. Two types of unit are used to express measurement results: signal-to-cut-off (s/co) ratios and the cut-off index (COI). The s/co ratio is obtained by dividing the signal strength for optical density of a sample by the signal strength for optical density of an internal cut off. As the s/co ratio is a ratio to the cut off and does not reflect the actual amount of antibody. In many test kits in which s/co ratios are used, the s/co is set between 0.75 and 1.1. The COI is obtained by converting the cut-off value (threshold value) obtained to 1.0 and indexing it. The COI standardizes measurements. As the COI is a ratio to the positive standard, it does not reflect the actual amount of antibody. In many test kits in which COI is used, it is set to 1.0.
20 | Rate of invalid, error and unreturnable results | < 5%
21 | Precision (repeatability/reproducibility) | Not applicable
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Lateral flow immunoassay. The assay can detect antibodies to HCV genotypes 1 to 6.
24 | Specimen(s) stability | Venous whole blood can be stored at 2–8 °C for 2 days. Serum and plasma can be stored at 2–8 °C for 5 days or at –20 °C or colder for > 5 days.
25 | Specimen(s) volume | 5–100 µL
26 | Type of result | Qualitative
27 | Time to result | ≤ 20 min
28 | End-point stability | ≤ 20 min
29 | Ease of use for point-of-care tests only: number of steps that require precision | ≤ 3 steps
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 3 specimens per hour per operator
31 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
32 | Internal quality control | A built-in quality control line that must appear when a specimen is added is part of internal quality control. Absence of the control line invalidates the test. Re-testing should be recommended in the procedure in case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected from humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected from humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected from humidity and direct sunlight. Any specific requirement once the reagents are opened should be clearly stated.
38 | Transport stability for specimen collection(temperature and humidity) | Venous whole blood can be stored at 2–8 °C for 2 days. Serum and plasma can be stored at 2–8 °C for 5 days or at –20 °C or colder for > 5 days.
39 | Storage stability for specimen collection (temperature and humidity) | Venous whole blood can be stored at 2–8 °C for 2 days. Serum and plasma can be stored at 2–8 °C for 5 days or at –20 °C or colder for > 5 days.
40 | In-use stability for specimen collection (temperature and humidity) | Venous whole blood can be stored at 2–8 °C for 2 days. Serum and plasma can be stored at 2–8 °C for 5 days or at –20 °C or colder for > 5 days.
41 | Transport stability for controls/calibrators (temperature and humidity) | Not applicable
42 | Storage stability for controls/calibrators (temperature and humidity) | Not applicable
43 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability for controls/calibrators (temperature and humidity) | Not applicable
45 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months from date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time to charge battery | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Displayed parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability/avaialability of a test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Not required
62 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature will require electricity.
63 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits
64 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
65 | Kit components – test, reagents, consumables (if relevant) | Test strips, devices or cassette and/or buffer/diluent
66 | Reagent/kit size (number of tests) | One test or more
67 | Consumables required but not provided in the test kit | For collecting finger-prick or capillary whole blood, lancets are required. If capillary tubes or disposable pipetters are not provided, tips for micropipette are required. For venous whole blood collection, blood collection tubes and needles are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if disposable pipetters for venous whole blood are not provided, tips for micropipettes are required. Sodium hypochlorite (bleach) for sterilization before disposal of infectious waste is required.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection or for collecting finger-prick or capillary whole blood is required. For testing, a timer is required. If several tests are being performed. more than one timer would be an advantage. If disposable pipetters are not provided, a calibrated micropipette and a micropipette stand are required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Required. The instructions for use submitted must correspond to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. The certificate should be retained by the testing centre with the dates on which the lot was used.
73 | Material safety data sheet | Required. The technical document should provide detailed, comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This information is usually available in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of in-country technical support | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification status | This assay with an RDT format was within the scope of WHO Prequalification as of March 2023. There is a WHO-prequalified product, and prequalification is recommended. See WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 3
Antibodies to hepatitis C virus (anti-HCV)
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | GMDN code/ Term name | 48365/ Hepatitis C virus total antibody IVD, kit, enzyme immunoassay (EIA)
3 | EMDN code/ Term name | W0105020303/ HCV ANTIBODY (TOTAL)
4 | Generic name | Antibodies to hepatitis C virus (anti-HCV) immunoassay
5 | Specific type or variation (optional) | Qualitative assay
6 | Alternative name/s (optional) | Enzyme immunoassay (EIA) for antibodies to hepatitis C virus (anti-HCV)
7 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV antibody, hepatitis, microtitre plate enzyme immunoassay, HCV
Intemded use
8 | Detection target | Hepatitis C virus antibodies
9 | Test purpose | To screen for or to aid in the diagnosis of viraemic HCV infection: children, adolescents and adults
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Assay format | Immunoassay (manual method)
12 | Specimen type | Plasma and serum
13 | Testing population | Children, adolescents and adults
14 | Intended user | Trained laboratory scientist, technologist, technician, specialist
15 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
16 | Clinical sensitivity | > 99%
17 | Clinical specificity | > 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HCV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Limit of detection | Currently, there are no international units or international reference materials for measuring HCV antibodies, unlike for those of HBV. Two types of unit are used to express measurement results: signal-to-cut-off (s/co) ratios and the cut-off index (COI). The s/co ratio is obtained by dividing the signal strength for optical density of a sample by the signal strength for optical density of an internal cut off. As the s/co ratio is a ratio to the cut off and does not reflect the actual amount of antibody. In many test kits in which s/co ratios are used, the s/co is set between 0.75 and 1.1. The COI is obtained by converting the cut-off value (threshold value) obtained to 1.0 and indexing it. The COI standardizes measurements. As the COI is a ratio to the positive standard, it does not reflect the actual amount of antibody. In many test kits in which COI is used, it is set to 1.0.
20 | Rate of invalid, error and unreturnable results | < 5%
21 | Precision (repeatability/reproducibility) | Intra-assay (reproducibility): CV ≤ 13% Inter-assay (reproducibility at one site): CV ≤ 15%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Enzyme immunoassay with microwells or strips coated with recombinant HCV antigens. The test can detect antibodies to HCV genotypes 1–6 in human serum or plasma.
24 | Specimen(s) stability | At 2–8 °C ≤ 7 for 2 h and at –20 °C for > 72 h
25 | Specimen(s) volume | 10–100 µL
26 | Type of result | Qualitative (reactive or non-reactive)
27 | Time to result | 180–240 min
28 | End-point stability | < 60 min
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 40 specimens per day per operator
31 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
32 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify that the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement for reagents once they are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
38 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for ≤ 72 h and at –20 °C for > 72 h
39 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for ≤ 72 h and at –20 °C for > 72 h
40 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for ≤ 72 h and at –20 °C for > 72 h
41 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
42 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
43 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement for reagents once they are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
45 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months for kits and reagents and 6 months for controls and calibrators from date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment https://apps.who.int/iris/handle/10665/255577
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time to charge battery | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Displayed parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or cosed system | Not applicable
60 | Multi-disease testing capability/availability of a test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Distilled or deionized water of adequate quality for enzyme immunoassays and enzyme-linked immunosorbent assays. See the instruction for use or package insert of the kit/reagents.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit components – test, reagents, consumables (if relevant) | Coated plate with wells, conjugate and positive and negative samples, wash buffer, substrate concentrate, stop solution and package insert or instructions for use
66 | Reagent/kit size (number of tests) | ≥ 20 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents that the kits do not contain are required. See the package insert or instructions for use of the kit or reagent.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader/enzyme immunoassay spectrophotometer, an incubator or water bath, a microwell plate or strip washer/enzyme immunoassay washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with hood and a cover for calibrating the micropipettes that can weigh ≤ 0.001 g (1 µL), a micropipette stand, a timer, a plate cover, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | The instructions for use submitted must correspond to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. The certificate should be retained by the testing centre with the dates on which the lot was used.
73 | Material safety data sheet | The technical document should provide detailed, comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | Potential source of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This information is usually in material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of in-country technical support | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification status | This assay in an RDT format was within the scope of WHO Prequalification as of March 2023. There is a WHO-prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 4 (Auto)
Antibodies to hepatitis C virus (anti-HCV) |  | This technical specification is for one specific test to be run on an automated analyser. As automated analysers are usually used to run a number of different tests, however, the selection of equipment should also depend on the other tests performed in the facility and not only the test described below.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | GMDN code/ Term name | 48365/ Hepatitis C virus total antibody IVD, kit, enzyme immunoassay (EIA)
3 | EMDN code/ Term name | W0105020303/ HCV ANTIBODY (TOTAL)
4 | Generic name | Antibodies to hepatitis C virus (anti-HCV) immunoassay
5 | Specific type or variation (optional) | Qualitative assay
6 | Alternative name/s (optional) | Chemiluminescent immunoassay for antibodies to hepatitis C virus (anti-HCV) Chemiluminescent enzyme immunoassay for antibodies to hepatitis C virus (anti-HCV)
7 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV, hepatitis, chemiluminescent immunoassay, CLIA, chemiluminescent enzyme immunoassay, HCV
Intended use
8 | Detection target | Hepatitis C virus antibodies
9 | Test purpose | To screen for or to aid in the diagnosis of viraemic HCV infection: children, adolescents and adults
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Assay format | Immunoassay (automated method)
12 | Specimen type | Plasma and serum
13 | Testing population | Children, adolescents and adults
14 | Intended user | Trained laboratory scientist, technologist, technician, specialist
15 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 99%
17 | Clinical specificity | ≥ 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HCV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Limit of detection | Currently, there are no international units or international reference materials for measuring HCV antibodies, unlike for those of HBV. Two types of unit are used to express measurement results: signal-to-cut-off (s/co) ratios and the cut-off index (COI). The s/co ratio is obtained by dividing the signal strength for optical density of a sample by the signal strength for optical density of an internal cut off. As the s/co ratio is a ratio to the cut off and does not reflect the actual amount of antibody. In many test kits in which s/co ratios are used, the s/co is set between 0.75 and 1.1. The COI is obtained by converting the cut-off value (threshold value) obtained to 1.0 and indexing it. The COI standardizes measurements. As the COI is a ratio to the positive standard, it does not reflect the actual amount of antibody. In many test kits in which COI is used, it is set to 1.0.
20 | Rate of invalid, error and unreturnable results | Not applicable
21 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV < 10%; inter-assay precision (reproducibility): CV < 15%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Chemiluminescent immunoassay or chemiluminescent enzyme immunoassay. The test can detect HCV genotypes 1–6.
24 | Specimen(s) stability | At 2–8° C for 48 h and at –20 °C for < 48 h
25 | Specimen(s) volume | ≤ 200 µL
26 | Type of result | Qualitative
27 | Time to result | ≤ 60 min
28 | End-point stability | Not applicable
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | Low sample volume laboratory Specimens per day per analyser: ≤ 500 (Sample throughput per hour per analyser: ≤ 62 ) Middle sample volume laboratory Specimens per day per analyser: 500–1000 (Sample throughput per hour: 63–125) High sample volume laboratory Specimens per day per analyser: ≥ 1000 (Sample throughput per hour: ≥ 126)
31 | Test limitations | The manufacturer should state any test limitations including conditions that have not been validated with the test, e.g. testing population, sample types, users.
32 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify that the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how positive, negative and invalid results are read and determined.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
36 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of the analyser's reagent unit or reagent storage
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources. Any specific requirements once the reagents are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
38 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 48 h, at –20 °C or lower for > 48 h
39 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 48 h, at –20 °C or lower for > 48 h
40 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 48 h, at –20 °C or lower for > 48 h
41 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources
42 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources
43 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature of analyzer's measuring unit
44 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources. Any specific requirements once the reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 9 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Depends on the brand or model and whether bench-top or floor-standing
48 | Weight of the device (kg) | Depends on the brand or model and whether bench-top or floor-standing
49 | Power requirements | 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz
50 | Time to charge battery | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Ideally, an uninterruptible power source
53 | Operating conditions (temperature, humidity and altitude) | Temperature range: 18–30 °C; humidity: < 80%; altitude: ≤ 2000 m
54 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle ware
55 | Displayed parameters | Digital display of test results, electronic and printed readout
56 | Display languages | English or local language
57 | Built-in memory storage capacity | Analyser for low sample volume laboratory: ≥ 800 tests Analyser for middle sample volume laboratory: ≥ 2000 tests Analyser for high sample volume laboratory: ≥ 10 000 tests
58 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programme database for sharing information. Capability for data transmission, e.g. via USB, electronic health records and health programme databases
59 | Open or closed system | Both open and closed systems
60 | Multi-disease testing capability/availability of a test menu | Multi-disease testing
Infrastructure requirements
61 | Water requirements | For analysers with a wet system, deionized water of adequate quality. See the analyser's operational manual, instructions for use or package insert of kit/reagents. For analysers with a dry system, water is not required.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit components – test, reagents, consumables (if relevant) | In general, buffer/diluent, antigen and conjugate If only cups, trays, cartridges or wells for the test are included in the test kit component, other necessary reagents specified by the manufacturer are required. See the package insert or instructions for use of the test kit.
66 | Reagent/kit size (number of tests) | ≥ 42 tests
67 | Consumables required but not provided in the test kit | For collecting venous whole blood, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection, alcohol swabs, tissues, permanent markers/pens, biohazard bags and disposable gloves are required. For testing, controls, calibrator and other reagents and consumables specified by the manufacturer, disposable gloves, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste and biohazard bags are required. See the package insert or instructions for use of the test kit.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyser (if specified by the kits/reagent, that specified is required), a timer, micropipettes and a micropipette stand, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
69 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer
Documentation
70 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must include recommended procedures for disposal and any probable hazards to the environment and/or community.
71 | Instructions for use | Instructions for use submitted in the local language. Must correspond to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. Th certificate should be retained by the testing centre with the dates on which the lot was used.
73 | Material safety data sheet | Required. The manufacturer should supply detailed,comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazardous classification | The manufacturer should clearly document the potential source of harm in kit inserts. Classification should be according to the Global Harmonized System, e.g. health, physical or environmental hazard. More details are usually included in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification, e.g. simple waste disposal, hazardous waste, incineration.
Training, installation and use
76 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range should be assessed.
77 | Requirements for installation and calibration | The manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
78 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
79 | Warranty | Manufacturer must provide a minimum of 1 year’s warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty must be specified.
80 | Preventive maintenance | The minimal maintenance expected to be conducted by user and by the supplier should be stated. The frequency of maintenance should be either fixed or based on the number of tests the instrument processes, which will de dictated by the workload.
81 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or supplier will describe in detail the maintenance support they will supply.
82 | Type of service contract (including leasing or reagent rental) | Must cover labour, repair, spare parts, shipping and logistics costs and training. Possibility of leasing an instrument for smaller devices.
83 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
84 | Availability of in-country technical support | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
85 | Estimated lifespan | 5 years after the start of use or installation on the condition of the regular maintenance and inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures performed under normal use and operating conditions
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification status | This assay in an RDT format was within the scope of WHO Prequalification as of March 2023. There is a WHO-prequalified product, and prequalification is recommended. See WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 5
Combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg)
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | GMDN code/ Term name | 48381/ Hepatitis C virus antibody/antigen IVD, kit, enzyme immunoassay (EIA)
3 | EMDN code/ Term name | W0105020308/ HEPATITIS C VIRUS (HCV) ANTIGEN/ANTIBODY COMBO TEST
4 | Generic name | Combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg) immunoassay
5 | Specific type or variation (optional) | Qualitative assay
6 | Alternative name/s (optional) | Enzyme immunoassay (EIA) for combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg)
7 | Keywords (optional) | Hepatitis C virus antibody and core antigen, hepatitis, microtitre plate enzyme immunoassay, HCV
Intended use
8 | Detection target | Hepatitis C virus antibodies and core antigen
9 | Test purpose | To screen for HCV infection and to aid in the diagnosis of viraemic HCV infection: children, adolescents and adults
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Assay format | Immunoassay (manual method)
12 | Specimen type | Plasma and serum
13 | Testing population | Children, adolescents and adults
14 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
15 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 99%
17 | Clinical specificity | ≥ 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HCV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Limit of detection | Currently, there are no international units or international reference materials for measuring HCV antibodies, unlike for those of HBV. Two types of unit are used to express measurement results: signal-to-cut-off (s/co) ratios and the cut-off index (COI). The s/co ratio is obtained by dividing the signal strength for optical density of a sample by the signal strength for optical density of an internal cut off. As the s/co ratio is a ratio to the cut off and does not reflect the actual amount of antibody. In many test kits in which s/co ratios are used, the s/co is set between 0.75 and 1.1. The COI is obtained by converting the cut-off value (threshold value) obtained to 1.0 and indexing it. The COI standardizes measurements. As the COI is a ratio to the positive standard, it does not reflect the actual amount of antibody. In many test kits in which COI is used, it is set to 1.0. On HCV core antigen: < 1700 IU/mL
20 | Rate of invalid, error and unreturnable results | < 5%
21 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV ≤ 10%; inter-assay precision (reproducibility): CV ≤ 15%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Enzyme immunoassay, microwells coated with recombinant HCV antigens and anti-HCVcAg monoclonal antibody. The test should detect HCV genotypes 1–3, at least. Ideally, the test detects HCV subtypes 1–6.
24 | Specimen(s) stability | At 2–8 °C for 72 h, at –20 °C beyond 72 h
25 | Specimen(s) volume | 50–100 µL
26 | Type of result | Qualitative (reactive or non-reactive)
27 | Time to result | 120–180 min
28 | End-point stability | 10–20 min
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 40 specimens per day per operator
31 | Test limitations | The manufacturer should state any test limitations including conditions that have not been validated with the test, e.g. testing population, sample types, users.
32 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify that the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how positive, negative and invalid results are read and determined.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirements once the reagents are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
38 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h, at –20 °C beyond 72 h
39 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h, at –20 °C beyond 72 h
40 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h, at –20 °C beyond 72 h
41 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
42 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
43 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement once controls are opened should be clearly stated. The instructions should indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time to charge battery | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Displayed parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability/availability of a test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Distilled or deionized water of adequate quality for enzyme immunoassay or enzyme-linked immunosorbent assay. See the instructions for use or package insert of the kit/reagents.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit components – test / reagents / consumables (if relevant) | Coated plate with wells, conjugate and positive and negative controls, wash buffer, substrate concentrate, stop solution and package insert or instructions for use.
66 | Reagent/kit size (number of tests) | ≥ 96 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents that the kits do not contain are required. See the package insert or instructions for use of the kit/reagent.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader/enzyme immunoassay spectrophotometer, an incubator or water bath, a microwell plate or strip washer/enzyme immunoassay washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with hood and a cover for calibrating the micropipettes that can weigh ≤ 0.001 g (1 µL), a micropipette stand, a timer, a plate cover, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | The instructions for use submitted must correspond to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. The certificate should be retained by the testing centre with the dates on which the lot was used.
73 | Material safety data sheet | The technical document should provide detailed, comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazardous classification | Potential source of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This information is usually contained in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of in-country technical support | The availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification status | This assay in an RDT format is within the scope of WHO Prequalification as of March 2023. There is a WHO-prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 6 (Auto)
HCV core antigen (HCVcAg) |  | This technical specification is for one specific test to be run on an automated analyser. As automated analysers are usually used to run a number of different tests, however, the selection of equipment should also depend on the other tests performed in the facility and not only the test described below.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | GMDN code/ Term name | 48354/ Hepatitis C virus core antigen IVD, kit, enzyme immunoassay (EIA)
3 | EMDN code/ Term name | W0105020302/ HCV ANTIGEN - CONFIRMATORY
4 | Generic name | HCV core antigen (HCVcAg) immunoassay
5 | Specific type or variation (optional) | Qualitative assay
6 | Alternative name/s (optional) | Chemiluminescent immunoassay (CLIA) for HCV core antigen (HCVcAg) Chemiluminescent enzyme immunoassay (CLEIA) for HCV core antigen (HCVcAg)
7 | Keywords (optional) | Hepatitis C virus core antigen, chemiluminescent immunoassay, CLIA, chemiluminescent enzyme immunoassay, CLEIA, qualitative immunoassay, HCV
Intended use
8 | Detection target | Hepatitis C virus core antigen
9 | Test purpose | To aid in the diagnosis of viraemic HCV infection
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Assay format | Immunoassay (automated method)
12 | Specimen type | Plasma and serum
13 | Testing population | Adults and children (0–18 years)
14 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
15 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance chracteristics
16 | Clinical sensitivity | ≥ 97%
17 | Clinical specificity | ≥ 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HCV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Limit of detection | ≤ 3 fmol/L
20 | Rate of invalid, error and unreturnable results | Not applicable
21 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV < 10%; inter-assay precision (reproducibility): CV < 15%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Chemiluminescent immunoassay or chemiluminescent enzyme immunoassay can detect HCV genotypes 1–6.
24 | Specimen(s) stability | At 2–8 °C for 24 h, –20 °C or lower beyond 24 h
25 | Specimen(s) volume | ≤ 150 µL
26 | Type of result | Qualitative
27 | Time to result | ≤ 60 min
28 | End-point stability | Not applicable
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | Low sample volume laboratory Specimens per day per analyser: ≤ 500 (Sample throughput per hour per analyser: ≤ 62 ) Middle sample volume laboratory Specimens per day per analyser: 500–1000 (Sample throughput per hour: 63–125) High sample volume laboratory Specimens per day per analyser: ≥ 1000 (Sample throughput per hour: ≥ 126)
31 | Test limitations | The manufacturer should state any test limitations including conditions that have not been validated with the test, e.g. testing population, sample types, users.
32 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify that the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how positive, negative and invalid results are read and determined.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
36 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of the analyser's reagent unit or reagent storage
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
38 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h, –20 °C or lower beyond 24 h
39 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h, –20 °C or lower beyond 24 h
40 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h, –20 °C or lower beyond 24 h
41 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources.
42 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources
43 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature of analyser's measuring unit
44 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources. Any specific requirement forcontrols or calibrators once opened should be clearly stated. The instructions should indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months
46 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Depends on the brand or model and whether bench-top or floor-standing
48 | Weight of the device (kg) | Depends on the brand or model and whether bench-top or floor-standing
49 | Power requirements | 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz
50 | Time to charge battery | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Ideally, uninterruptible power source
53 | Operating conditions (temperature, humidity and altitude) | Temperature range: 18–30 °C; humidity: < 80%; altitude: ≤ 2000 m
54 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle ware
55 | Displayed parameters | Digital display of test results, electronic and printed readout.
56 | Display languages | English or local language
57 | Built-in memory storage capacity | Analyser for low sample volume laboratory: ≥ 800 tests Analyser for middle sample volume laboratory: ≥ 2000 tests Analyser for high sample volume laboratory: ≥ 10 000 tests
58 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programme database for sharing information. Capability for data transmission, e.g. via USB, electronic health records and health programme databases.
59 | Open or closed system | Closed system
60 | Multi-disease testing capability/availability of a test menu | Multi-disease testing
Infrastructure requirements
61 | Water requirements | For analysers with a wet system, deionized water of adequate quality. See the analyser's operational manual, instructions for use or package insert of kit/reagents. For analysers with a dry system, water is not required.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit components – test, reagents, consumables (if relevant) | In general, buffer/diluent, antigen and conjugate If only cups, trays, cartridges or wells for the test are included in the test kit component, other reagents specified by the manufacturer are required. Please see the package insert or instructions for use of the test kit.
66 | Reagent/kit size (number of tests) | ≥ 40 tests
67 | Consumables required but not provided in the test kit | For collecting venous whole blood, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection, alcohol swabs, tissues, permanent markers/pens, biohazard bags and disposable gloves are required. For testing, controls, calibrator and other necessary reagents and consumables specified by the manufacturer, disposable gloves, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste and biohazard bags are required. Please see the package insert or instructions for use of the test kit.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyser (if specified by the kit/reagent, that specified is required), a timer, micropipettes and a micropipette stand, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
69 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer
Documentation
70 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and a list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
71 | Instructions for use | Instructions for use submitted in the local language. Must correspond to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing centre with the dates on which the lot was used.
73 | Material safety data sheet | Required. The manufacturer should supply detailed, comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation related to handling, storage, use and disposal of the product; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | The manufacturer should clearly document potential sources of harm in kit inserts. Classification should be according to the Global Harmonized System, e.g. health, physical or environmental hazard. More details are usually provided in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g. simple waste disposal, hazardous waste, incineration.
Training, installation and use
76 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range should be assessed.
77 | Requirements for installation and calibration | The manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
78 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
79 | Warranty | Manufacturer must provide a minimum of 1 year’s warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
80 | Preventive maintenance | The minimal maintenance expected to be conducted by the user and the supplier should be stated. The frequency of maintenance should be either fixed or based on the number of tests the instrument processes, which will de dictated by the workload.
81 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed in the kit insert. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or supplier will describe in detail the maintenance support they will supply.
82 | Type of service contract (including leasing or reagent rental) | Must cover labour, repair, spare parts, shipping and logistics costs and training. Possibility of leasing instrument for smaller devices.
83 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
84 | Availability of in-country technical support | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
85 | Estimated lifespan | 5 years after the start of use or installation on the condition that the regular maintenance and inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures are performed under normal use and operating conditions.
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification status | This assay ina RDT format is within the scope of WHO Prequalification as of March 2023. There is a WHO-prequakified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 7
Qualitative or quantitative HCV nucleic acid test+B2:D34D82B2:D29B2:D44B2:D59D82B2:D29B2:D72D82B2:D29B2:D81D82B2:D29B2:D88D82B2:D29B2:D97D82B2:D29B2:D100D82B2:D29B2:D101B2:D100B2:D99B2:D98B2B2:D100
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | GMDN code/ Term name | 48374/ Hepatitis C virus nucleic acid IVD, kit, nucleic acid technique (NAT)
3 | EMDN code/ Term name | W0105020307/ HEPATITIS C VIRUS - NA REAGENTS
4 | Generic name | HCV nucleic acid test
5 | Specific type or variation (optional) | Quantitative or qualitative test
6 | Alternative name/s (optional) | Quantitative HCV nucleic acid test Qualitative HCV nucleic acid test
7 | Keywords (optional) | Hepatitis C virus, HCV, nucleic acid test (NAT), nucleic acid amplification test (NAAT), quantitative or qualitative test, reverse transcription polymerase chain reaction, RT-PCR, quantitative PCR, real-time polymerase chain reaction, real-time PCR, qPCR, real-time reverse transcription polymerase chain reaction, real-time RT-PCR, qRT-PCR
Intended use
8 | Detection target | Hepatitis C virus nucleic acid (HCV RNA)
9 | Test purpose | To diagnose viraemic HCV, to monitor response to treatment and as a test of cure
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Test format | Nucleic acid test (manual method)
12 | Specimen type | Plasma and/or serum
13 | Testing population | Adults and children (0–18 years)
14 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
15 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
16 | Clinical sensitivity | > 99%
17 | Clinical specificity | > 98%
18 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HCV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Linear range | 65–1 000 000 IU/mL
20 | Limit of detection | ≤ 65 IU/mL
21 | Rate of invalid, error and unreturnable results | < 5%
22 | Precision (repeatability/reproducibility) | Data on basis of the quantitative results (in IU/mL) Intra-assay variability/precision (low positive samples, 250 IU/mL): < 23% Inter-assay variability/precision (low positive samples, 250 IU/mL): < 30%
23 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
24 | Principle of the assay | Real-time reverse transcription polymerase chain reaction (qRT- PCR) for amplification of specific regions of HCV nucleic acid. The PCR product should be detected with fluorescent oligonucleotide probes that bind specifically to HCV nucleic acid. The test can detect HCV genotypes 1–6.
25 | Specimen(s) stability | At 2–8 °C ≥ 3 days, at –20 °C for several weeks and at –70 °C for several months
26 | Specimen(s) volume | 50–1000 µL
27 | Type of result | Quantitative
28 | Time to result | Approximately 4 h, including RNA extraction step
29 | End-point stability | No information
30 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
31 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | 24–48 tests per 8-h day
32 | Test limitations | The manufacturer should state any test limitations including conditions that have not been validated with the test, e.g. testing population, sample types, users.
33 | Internal quality control | Internal control should be provided for sample preparation (nucleic acid extraction) and the subsequent real-time reverse transcription PCR. This allows identification of sample-specific reaction inhibition and systemic malfunction due to reagent failure for consideration in interpreting the results. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how positive, negative and invalid results are read and determined.
34 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
35 | Storage stability of kit/reagents (temperature and humidity) | At –25 °C to –15 °C or at 2–8 °C
36 | Transport stability of kit/reagents (temperature and humidity) | At –25 °C to –15 °C or at 2–8 °C
37 | On-board stability of kit/reagents (temperature and humidity) | At the temperature inside the instrument
38 | In-use stability of kit/reagents (temperature and humidity) | At 0–2°C or at 2–8 °C
39 | Transport stability for specimen (temperature and humidity) | Specimens can be stored at 2–8 °C for ≤ 1 day and at –20 °C to –70 °C for several months.
40 | Storage stability for specimen (temperature and humidity) | Specimens can be stored at 2–8 °C for ≤ 1 day and at –20 °C to –70 °C for several months.
41 | In-use stability for specimen (temperature and humidity) | Specimens can be stored at 2–8 °C for ≤ 1 day and at –20 °C to –70 °C for several months.
42 | Transport stability for controls/calibrators (temperature and humidity) | At –25 °C to –15 °C or at 2–8 °C protected against humidity and direct sunlight
43 | Storage stability for controls/calibrators (temperature and humidity) | At –25 °C to –15 °C or at 2–8 °C protected against humidity and direct sunlight
44 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature inside the instrument
45 | In-use stability for controls/calibrators (temperature and humidity) | At –25 °C to –15 °C or at 2–8 °C protected against humidity and direct sunlight
46 | Shelf-life of kit/reagents upon manufacture (months) | 12 months
47 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see p, 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
48 | Size of device (height x width x depth) | Nucleic acid extraction, isolation and purification: not applicable Any commercially available nucleic acid extraction kit approved by regulatory authorities can be used. If the testing kit specifies a particular nucleic acid extraction kit, please use that product. Real-time PCR instrument: depends on the brand or model Any commercially available real-time PCR instrument approved by regulatory authorities can be used for amplification and detection of extracted nucleic acids. If the testing kit specifies a particular real-time PCR instrument, please use that product.
49 | Weight of the device (kg) | As above
50 | Power requirements | Real-time PCR instrument: 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz
51 | Time to charge battery | Ideally, uninterruptible power source
52 | Battery duration | Nucleic acid extraction, isolation and purification: not applicable Any commercially available nucleic acid extraction kit approved by regulatory authorities can be used. If the testing kit specifies a particular nucleic acid extraction kit, please use that product. Real-time PCR instrument: depends on the brand or model Any commercially available real-time PCR instrument approved by regulatory authorities can be used for amplification and detection of extracted nucleic acids. If the testing kit specifies a particular real-time PCR instrument, please use that product.
53 | Alternative charging options | As above
54 | Operating conditions (temperature and humidity) | As above
55 | User interface | As above
56 | Displayed parameters | As above
57 | Display languages | As above
58 | Built-in memory storage capacity | As above
59 | Diagnostic connectivity | As above
60 | Open or closed system | Open system
61 | Multi-disease testing capability/availability of a test menu | Analysis of multiple analytes
Infrastructure requirements
62 | Water requirements | Molecular-grade water
63 | Electricity requirements | Yes
64 | Refrigeration or cold chain for storage of kit/reagents | Yes
65 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
66 | Kit components – test, reagents / consumables (if relevant) | Reagents required for real-time reverse transcription PCR, including internal control, quality standards and controls, primers, probes, polymerase enzyme and package insert or instructions for use
67 | Reagent/kit size (number of tests) | ≥ 24 tests
68 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, a commercially available nucleic acid extraction, isolation and purification kit approved by regulatory authorities, HCV control and calibrator kits if not provided, 96-well plates, tubes of PCR, covers, seals or foils for sealing the plates, aerosol barrier pipette tips, containers or tubes for specimens, disposable powder-free gloves, 5–7% (0.7–1.0 M) sodium hypochlorite solution for sterilization before disposal of infectious waste, 70% ethanol and biohazard bags are required.
69 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a biological safety cabinet approved for working with infectious materials, a laminar flow hood for control and reagent preparation, a desktop microcentrifuge for tubes, a plate centrifuge, a vortex mixer, a plate-shaking platform, a biohazard waste container, a refrigerator, a freezer, an automated nucleic acid extraction, isolation and purification instrument, a real-time PCR instrument or thermal cycler (specified for the test kit) with software, calibrated adjustable micropipettes, a cooling block, a timer, an electronic scale with hood and a cover for calibrating micropipettes that can weigh ≤ 0.001 g (1 µL, a micropipette stand and an autoclave are required. (The autoclave cannot be used for nucleic products due to the possibility of explosion and diffusion of nucleic products in the laboratory.) If the the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
70 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer
Documentation
71 | Instrument operator manual | For the real-time PCR instrument: Operating and service manuals in the local language, including lists of important spares and accessories with their reference numbers and a list of equipment and procedures required for calibration and routine maintenance. Documentation must include recommended procedures for disposal and any probable hazards to the environment and/or community.
72 | Instructions for use | Instructions for use in the local language. Must correspond to the regulatory version registered for sale and use in the country of supply.
73 | Certificate of analysis | A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
74 | Material safety data sheet | Required. The manufacturer should supply detailed, comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
75 | Hazard classification | The manufacturer should clearly document potential sources of harm on kit inserts. Harm should be classified according to the Global Harmonized System, e.g. health, physical or environmental hazard. More details are usually be provided in the material safety data sheet.
76 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g., simple waste disposal, hazardous waste, incineration
Training, installation and use
77 | Pre-installation requirements (if relevant) | For real-time PCR instrument, access, space, water, power, waste, refrigeration and room temperature range should be assessed.
78 | Requirements for installation and calibration | For real-time PCR instrument, the manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
79 | Training of user/s (if relevant) | For real-time PCR instrument, manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
80 | Warranty | For real-time PCR instrument, the manufacturer must provide a minimum 1-year warranty from the date of handover and specify the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty.
81 | Preventive maintenance | For real-time PCR instrument, the minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either fixed or according to the number of tests processed, which will be dictated by the workload.
82 | Corrective maintenance | For real-time PCR instrument, specific requirements for calibration and test procedures must be described in the kit insert. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or the supplier will specify the maintenance support they will provide.
83 | Type of service contract (including leasing or reagent rental) | For real-time PCR instrument, must cover the costs of labour, repair, spare parts, shipping, logistics and training. Possible leasing of small instruments
84 | Availability of spare parts post-warranty | For the real-time PCR instrument used, spare parts should be available for the life of the system, which should be a minimum of 5 years.
85 | Availability of technical support in the country | For real-time PCT instrument, required. The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
86 | Estimated lifespan | For real-time PCR instrument, in general, 5 years after the start of use or installation on the condition of regular maintenance and inspection specified in the instruction manual, replacement of parts and necessary repairs or maintenance under normal use and operating conditions
Quality and registration
87 | Global regulatory approval | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
88 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
89 | WHO prequalification status | This assay is within the scope of WHO Prequalification as of March 2023. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
90 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
91 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 8 (Semi auto)
Quantitative or qualitative HCV nucleic acid test |  | This technical specification is for one specific test to be run on an automated analyser. As automated analysers are usually used to run a number of different tests, however, the selection of equipment should also depend on the other tests performed in the facility and not only the test described below.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | GMDN code/ Term name | 48374/ Hepatitis C virus nucleic acid IVD, kit, nucleic acid technique (NAT)
3 | EMDN code/ Term name | W0105020307/ HEPATITIS C VIRUS - NA REAGENTS
4 | Generic name | HCV nucleic acid test
5 | Specific type or variation (optional) | Quantitative or qualitative test
6 | Alternative name/s (optional) | Quantitative HCV nucleic acid test Qualitative HCV nucleic acid test
7 | Keywords (optional) | Hepatitis C virus, HCV, nucleic acid test, NAT, nucleic acid amplification test, NAAT, quantitative or qualitative test, reverse transcription polymerase chain reaction, RT-PCR, quantitative PCR, real-time polymerase chain reaction, real-time PCR, qPCR, real-time reverse transcription polymerase chain reaction, real-time RT-PCR, qRT-PCR and transcription meditated amplification, TMA
Intended use
8 | Detection target | Hepatitis C virus nucleic acid
9 | Test purpose | To diagnose viraemic HCV, to monitor response to treatment and as a test of cure
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Test format | Nucleic acid test (semi-automated method)
12 | Specimen type | Plasma and /or serum or Plasma, serum and dried blood spots
13 | Testing population | Adults and children (0–18 years)
14 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
15 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Performance characteristics
16 | Clinical sensitivity | > 96%
17 | Clinical specificity | > 95%
18 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HCV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Linear range | 50–100 000 000 IU/mL
20 | Limit of detection | ≤ 50 IU/mL
21 | Rate of invalid, error and unreturnable results | < 5%
22 | Precision (repeatability/reproducibility) | Data on basis of the quantitative results (in IU/mL) Intra-assay variability/precision: <35% Inter-assay variability/precision: <35%
23 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
24 | Principle of the assay | Real-time reverse transcription polymerase chain reaction (qRT-PCR) for amplification of specific regions of HCV nucleic acid. The PCR product should be detected with fluorescent oligonucleotide probes that bind specifically to HCV nucleic acid. The test should be able to detect HCV genotypes 1–6 at least.
25 | Specimen(s) stability | At 2–8 °C up to 3 days At –20 °C or colder for ≤ 6 weeks
26 | Specimen(s) volume | 200–1000 µL
27 | Type of result | Quantitative
28 | Time to result | ≤ 5 h
29 | End-point stability | Not applicable
30 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
31 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 72 specimens per 8-h day
32 | Test limitations | The manufacturer should state any test limitations including conditions that have not been validated with the test, e.g. testing population, sample types, users.
33 | Internal quality control | Internal control should be provided for sample preparation (nucleic acid extraction) and the subsequent real-time reverse transcription PCR. This allows identification of sample-specific reaction inhibition and systemic malfunction due to reagent failure for consideration in interpreting the results. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how positive, negative and invalid results are read and determined.
34 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
35 | Storage stability of kit/reagents (temperature and humidity) | At –25 °C to –15 °C or at 2–8°C protected from humidity and direct sunlight
36 | Transport stability of kit/reagents (temperature and humidity) | At –25 °C to –15 °C or at 2–8°C protected from humidity and direct sunlight
37 | On board stability of kit/reagents (temperature and humidity) | At the temperature of the analyser's measuring unit
38 | In-use stability of kit/reagents (temperature and humidity) | At 0–2 °C or at 2–8 °C
39 | Transport stability for specimen (temperature and humidity) | At 2–8 °C for ≤ 3 days At –20 °C or colder for ≤ 6 weeks
40 | Storage stability for specimen (temperature and humidity) | At 2–8 °C ≤ 3 days At –20 °C or colder for ≤ 6 weeks
41 | In-use stability for specimen (temperature and humidity) | At 2–8 °C for ≤ 3 days At –20 °C or colder for ≤ 6 weeks
42 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C or at –15 °C to –25 °C Repeated thawing and freezing should be avoided.
43 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C or at –15 °C to –25 °C Repeated thawing and freezing should be avoided.
44 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature inside the analyser
45 | In-use stability for controls/calibrators (temperature and humidity) | At –25 °C to –15 °C or at 2–8 °C protected from humidity and direct sunlight Any specific requirement once reagents are opened should be clearly stated.
46 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months
47 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
48 | Size of device (height x width x depth) | Automated instrument for nucleic acid extraction, isolation and purification: Depends on the brand or model Any commercially available automated instrument for extraction, isolation ad purification approved by regulatory authorities can be used. If the testing kit specifies a particular automated instrument, please use that product. Real-time PCR instrument: Depends on the brand or model Any commercially available real-time PCR instrument approved by regulatory authorities can be used for amplification and detection of extracted nucleic acids. If the testing kit specifies a particular real-time PCR instrument, please use that product.
49 | Weight of the device (kg) | As above
50 | Power requirements | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, 100–240 VAC/50–60 Hz. or 200–240 VAC/50–60 Hz
51 | Time to charge battery | Not applicable
52 | Battery duration | Not applicable
53 | Alternative charging options | Ideally, an uninterruptible power source
54 | Operating conditions (temperature and humidity) | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument: operating temperature range: 15–30 °C; humidity: 30–80%; altitude: ≤ 2000 m
55 | User interface | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, simple test menu, simple PC as provided by manufacturer or built-in PC, option for connection to laboratory computer and middle ware
56 | Displayed parameters | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, digital display of test results, electronic and printed readout
57 | Display languages | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, English or local language
58 | Built-in memory storage capacity | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, two types of memory storage for instruments: those containing the compute ring system (built-in) and those that do not contain the computing system (not built-in), depending on the brand or model. Please ask the manufacturer.
59 | Diagnostic connectivity | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, data connectivity compatible with laboratory information system and programme database for sharing information. Capability for data transmission, e.g. via USB, electronic health records and health programme databases.
60 | Open or closed system | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, closed or open system
61 | Multi-disease testing capability/availability of a test menu | Real-time PCR instrument: analysis of multiple analytes
Infrastructure requirements
62 | Water requirements | Yes, molecular biology-grade water
63 | Electricity requirements | Yes
64 | Refrigeration or cold chain for storage of kit/reagents | Yes
65 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
66 | Kit components – test, reagents, consumables (if relevant) | Reagents required for real-time reverse transcription PCR, including internal control, quality standards and controls, primers, probes, polymerase enzyme and package Insert or instructions for use
67 | Reagent/kit size (number of tests) | ≥ 24 tests
68 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, consumables depend on the kits. In general, controls, calibrators, reaction tubes, specimen diluent, lysis reagent and buffer solution that are specified for the kits are required. Please see the package insert or instructions for use. Common consumables for testing (sterile filter pipette tips, sterile microtubes, tube racks and disposable powder-free gloves, sodium hypochlorite solution at the concentration required on the instructions for use or package insert for sterilization before disposal of infectious waste, 70% ethanol and biohazard bags) are required.
69 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a biological safety cabinet approved for working with infectious materials, a laminar flow hood for control and reagent preparation, a desktop microcentrifuge for tubes, a plate centrifuge, a vortex mixer, a plate-shaking platform, a biohazard waste container, a refrigerator, a freezer, an automated nucleic acid extraction, isolation and purification instrument, a real-time PCR instrument or thermal cycler (specified for the test kit) with software, calibrated adjustable micropipettes, a cooling block, a timer, an electronic scale with hood and a cover for calibrating micropipettes that can weigh ≤ 0.001 g (1 µL), a micropipette stand and an autoclave are required (The autoclave cannot be used for nucleic products because of possible explosion and diffusion of nucleic products in the laboratory.) If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
70 | Spare parts (if relevant) | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer.
Documentation
71 | Instrument operator manual | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, operating and service manuals in the local language, including lists of important spares and accessories with their reference numbers and a list of equipment and procedures required for calibration and routine maintenance. Documentation must include recommended procedures for disposal and any probable hazards to the environment and/or community.
72 | Instructions for use | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, instructions for use in the local language. Must correspond to the regulatory version registered for sale and use in the country of supply.
73 | Certificate of analysis | A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
74 | Material safety data sheet | Required. The manufacturer should supply detailed,comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
75 | Hazardous classification | The manufacturer should clearly document potential sources of harm on kit inserts. Harm should be classified according to the Global Harmonized System, e.g., health, physical or environmental hazard. More details are usually be provided in the material safety data sheet.
76 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g., simple waste disposal, hazardous waste, incineration
Training, installation and use
77 | Pre-installation requirements (if relevant) | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, access, space, water, power, waste, refrigeration and room temperature range should be assessed.
78 | Requirements for installation and calibration | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, the manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
79 | Training of user/s (if relevant) | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, the manufacturer should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
80 | Warranty | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, the manufacturer must provide a minimum 1-year warranty from the date of handover and specify the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty.
81 | Preventive maintenance | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, the minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either fixed or according to the number of tests processed, which will be dictated by the workload.
82 | Corrective maintenance | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, specific requirements for calibration and test procedures must be described in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or the supplier will specify the maintenance support they will provide.
83 | Type of service contract (including leasing or reagent rental) | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, must cover the costs of labour, repair, spare parts, shipping, logistics and training. Possible leasing of small instruments
84 | Availability of spare parts post-warranty | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, spare parts should be available for the life of the system, which should be a minimum of 5 years.
85 | Availability of technical support in the country | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, required. The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
86 | Estimated lifespan | For automated instrument for nucleic acid extraction, isolation and purification and real-time PCR instrument, 5 years after the start of use or installation on the condition of regular maintenance and inspection specified in the instruction manual, replacement of parts and necessary repairs or maintenance under normal use and operating conditions
Quality and registration
87 | Global regulatory approval | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
88 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
89 | WHO prequalification status | This assay is within the scope of WHO Prequalification as of March 2023. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
90 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
91 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 9 (Full auto)
Quantitative or qualitative HCV nucleic acid test |  | This technical specification is for one specific test to be run on an automated analyser. As automated analysers are usually used to run a number of different tests, however, the selection of equipment should also depend on the other tests performed in the facility and not only the test described below.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | GMDN code/ Term name | 48374/ Hepatitis C virus nucleic acid IVD, kit, nucleic acid technique (NAT)
3 | EMDN code/ Term name | W0105020307/ HEPATITIS C VIRUS - NA REAGENTS
4 | Generic name | HCV nucleic acid test
5 | Specific type or variation (optional) | Quantitative or qualitative test
6 | Alternative name/s (optional) | Quantitative HCV nucleic acid test Qualitative HCV nucleic acid test
7 | Keywords (optional) | Hepatitis C virus, HCV, nucleic acid test, NAT, nucleic acid amplification test, NAAT, quantitative or qualitative test, reverse transcription polymerase chain reaction, RT-PCR, quantitative PCR, real-time polymerase chain reaction, real-time PCR, qPCR, real-time reverse transcription polymerase chain reaction, real-time RT-PCR, qRT-PCR, transcription mediated amplification, TMA
Intended use
8 | Detection target | Hepatitis C virus nucleic acid
9 | Test purpose | To diagnose viraemic HCV, to monitor response to treatment and as a test of cure
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Test format | Nucleic acid test (fully automated method)
12 | Specimen type | Plasma, serum or Capillary whole blood, venous whole blood, plasma and serum
13 | Testing population | Adults and children (0–18 years)
14 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
15 | Level of the health-care system | To be used in health-care facilities with clinical laboratories
Technical and operational characteristics
16 | Clinical sensitivity | Serum and plasma: > 99% Capillary whole blood and venous whole blood: > 96%
17 | Clinical specificity | Serum and plasma: > 99% Capillary whole blood and venous whole blood: > 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases, especially those presenting with signs and symptoms similar to those of HCV, or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Linear range | Serum and plasma: 50–100 000,000 IU/mL Capillary whole blood and venous whole blood: 100–100 000,000 IU/mL
20 | Limit of detection | Serum and plasma: ≤ 50 IU/mL Capillary whole blood and venous whole blood: ≤ 100 IU/mL
21 | Rate of invalid, error and unreturnable results | Serum and plasma:<5 %
22 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV ≤3 0%, inter-assay precision (reproducibility): CV ≤3 5%
23 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
24 | Principle of the assay | Real-time reverse transcription polymerase chain reaction (RT-PCR) or transcription-meditated amplification (TMA) for amplification of specific regions of HCV nucleic acid. The PCR or TMA product should be detected with fluorescent oligonucleotide probes that bind specifically to HCV nucleic acid. The test should be able to detect HCV genotypes 1–6.
25 | Specimen(s) stability | Capillary or venous whole blood can be stored at 2–25 ºC for ≤ 6 h. Plasma and serum can be stored at 2–8 ºC for ≥1 day. Plasma and serum can be stored at –18 ºC to –20 ºC for ≥ 30 days. Plasma and serum can be stored at –60 ºC or colder for > 30 days.
26 | Specimen(s) volume | Serum and plasma: 300–1200 µL Capillary whole blood and venous whole blood: ≥ 100 µL
27 | Type of result | Quantitative
28 | Time to result | Serum and plasma:105–165 min Capillary whole blood and venous whole blood: 60 min
29 | End-point stability | Not applicable
30 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
31 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | For low sample volume laboratories: 4–16 tests per 8-h day For middle sample volume laboratories: 16–300 tests per 8-h day For high sample volume laboratories: > 300 tests per 8-h day
32 | Test limitations | The manufacturer should state any test limitations including conditions that have not been validated with the test, e.g. testing population, sample types, users.
33 | Internal quality control | Internal control should be provided for sample preparation (nucleic acid extraction) and the subsequent real-time reverse transcription PCR. This allows identification of sample-specific reaction inhibition and systemic malfunction due to reagent failure for consideration in interpreting the results. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how positive, negative and invalid results are read and determined.
34 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
35 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C protected from humidity and direct sunlight in an upright position Ideally, > 8 °C
36 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C protected from humidity and direct sunlight in an upright position Ideally, > 8 °C
37 | On board stability of kit/reagents (temperature and humidity) | At the temperature of the analyser's reagent unit or reagent storage
38 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C Ideally, > 8 °C
39 | Transport stability for specimen (temperature and humidity) | Venous whole blood can be stored at 2–25 ºC for ≥ 6h. Plasma and serum can be stored at 2–8 ºC for ≥ 1 day. Plasma and serum can be stored at –18 ºC to –20 ºC for ≥ 30 days. Plasma and serum can be stored at–60 ºC or colder for > 30 days.
40 | Storage stability for specimen (temperature and humidity) | Venous whole blood can be stored at 2–25 ºC for ≥ 6h. Plasma and serum can be stored at 2–8 ºC for ≥ 1 day. Plasma and serum can be stored at –18 ºC to –20 ºC for ≥ 30 days. Plasma and serum can be stored at–60 ºC or colder for > 30 days.
41 | In-use stability for specimen (temperature and humidity) | Venous whole blood can be stored at 2–25 ºC for ≥ 6h. Plasma and serum can be stored at 2–8 ºC for ≥ 1 day. Plasma and serum can be stored at –18 ºC to –20 ºC for ≥ 30 days. Plasma and serum can be stored at–60 ºC or colder for > 30 days.
42 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C or at –15 °C to –25 °C Ideally, ≥ 8 °C The freeze–thaw cycle depends on the test kit and system.
43 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C or at –15 °C to –25 °C Ideally, ≥ 8 °C The freeze–thaw cycle depends on the test kit and system
44 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature inside the analyser
45 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 ºC or at 15–28 ºC Ideally, > 28 ºC
46 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months
47 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
48 | Size of device (height x width x depth) | Depends on the brand or model and whether bench-top or floor-standing
49 | Weight of the device (kg) | Depends on the brand or model and whether bench-top or floor-standing
50 | Power requirements | 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz
51 | Time to charge battery | Not applicable
52 | Battery duration | Not applicable
53 | Alternative charging options | Idealy, an uninterruptible power source
54 | Operating conditions (temperature and humidity) | Operating temperature range: 18–28 °C at least; humidity: 30–65% at least
55 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle ware
56 | Displayed parameters | Digital display of test results, electronic and printed readout
57 | Display languages | English or local language
58 | Built-in memory storage capacity | There are two types of memory storage for analysers, containing (built in) or not containing the computer ring system (not built in), depending on the brand or model. Please ask the manufactures.
59 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programme database for sharing information. Capability for data transmission. e.g. via USB, electronic health records and health programme databases.
60 | Open or closed system | Closed or semi-closed system
61 | Multi-disease testing capability/availability of a test menu | Multi-disease testing
Infrastructure requirements
62 | Water requirements | No
63 | Electricity requirements | Yes
64 | Refrigeration or cold chain for storage of kit/reagents | Yes
65 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
66 | Kit components – test, reagents, consumables (if relevant) | Reagents required for real-time PCR or TMA (RNA extraction and DNA amplification reagents) or Test trays or test cartridges with the reagents required for real-time PCR or TMA (RNA extraction and DNA amplification reagents), package inserts or instructions for use.
67 | Reagent/kit size (number of tests) | ≥ 10 tests
68 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, the consumables depend on the kits. In general, the controls, calibrators, reaction tubes, specimen diluent, lysis reagent and buffer solution specified for the kits are required. See the package insert or instructions for use. Common consumables for testing (sterile filter pipette tips, sterile microtubes, tube racks and disposable powder-free gloves, sodium hypochlorite solution at the concentration required in the instructions for use for sterilization before disposal of infectious waste and biohazard bags ) are required.
69 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum and plasma, a centrifuge, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, an analyser and specification file or software for the analyser are required. A biological safety cabinet approved for working with infectious materials is desirable. Common equipment (such as a timer, a vortex mixer, a centrifuge, calibrated micropipettes (10–1000 μL), a micropipette stand, an electronic scale with hood and a cover for calibrating the micropipettes (that can weigh ≤ 0.001 g (1 µL)), a refrigerator, a freezer and an autoclave) for testing are required. The autoclave cannot be used for nucleic products because of possible scattering of nucleic products. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
70 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer.
Documentation
71 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their reference numbers and a list of equipment and procedures required for calibration and routine maintenance. Documentation must include recommended procedures for disposal and any probable hazards to the environment and/or community.
72 | Instructions for use | Instructions for use in the local language. Must correspond to the regulatory version registered for sale and use in the country of supply.
73 | Certificate of analysis | A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
74 | Material safety data sheet | Required. The manufacturer should supply detailed,comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Ennvironmental and safety requirements
75 | Hazard classification | The manufacturer should clearly document potential sources of harm on kit inserts. Harm should be classified according to the Global Harmonized System, e.g. health, physical or environmental hazard. More details are usually provided in the material safety data sheet.
76 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g. simple waste disposal, hazardous waste, incineration
Training, installation and use
77 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range should be assessed.
78 | Requirements for installation and calibration | The manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
79 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
80 | Warranty | The manufacturer must provide a minimum 1-year warranty from the date of handover and specify the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty.
81 | Preventive maintenance | The minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either fixed or according to the number of tests processed, which will be dictated by the workload.
82 | Corrective maintenance | Specific requirements for calibration and test procedures must be described in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or the supplier will specify the maintenance support they will provide.
83 | Type of service contract (including leasing or reagent rental) | Must cover costs of labour, repair, spare parts, shipping, logistics and training. Possible leasing of small instruments
84 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 7 years.
85 | Availability of technical support in the country | Required. The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
86 | Estimated lifespan | 7 years after the start of use or installation on the condition of regular maintenance and inspection specified in the instruction manual, replacement of parts and necessary repairs or maintenance under normal use and operating conditions
Quality and registration
87 | Global regulatory approval | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
88 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
89 | WHO prequalification status | This assay is within the scope of WHO Prequalification as of March 2023. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
90 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
91 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 10
Antibodies to hepatitis C virus
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and cding
1 | WHO category
2 | GMDN code/ Term name | 30829/ Hepatitis C virus total antibody IVD, kit, rapid ICT, clinical
3 | EMDN code/ Term name | W0105090202/ HCV ANTIBODY -RT & POC
4 | Generic name | Antibodies to hepatitis C virus (anti-HCV) rapid diagnostic test
5 | Specific type or variation (optional) | Qualitative, immunochromatographic assay
6 | Alternative name/s (optional) | Rapid diagnostic test (RDT) for antibodies to hepatitis C virus (anti-HCV)
7 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV antibody, hepatitis, rapid diagnostic test, HCV
Intended use
8 | Detection target | Hepatitis C virus antibodies
9 | Test purpose | To screen donated blood for HCV
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Assay format | Rapid diagnostic test (RDT)
12 | Specimen type | Capillary whole blood, venous whole blood, plasma and serum
13 | Testing population | Adolescents and adults
14 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
15 | Level of the health-care system | For use in blood-screening laboratories
Performance characteristics
16 | Clinical sensitivity | > 99%
17 | Clinical specificity | > 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Limit of detection | Currently, there are no international units or international reference materials for measuring HCV antibodies, unlike for those of HBV. Two types of unit are used to express measurement results: signal-to-cut-off (s/co) ratios and the cut-off index (COI). The s/co ratio is obtained by dividing the signal strength for optical density of a sample by the signal strength for optical density of an internal cut off. As the s/co ratio is a ratio to the cut off and does not reflect the actual amount of antibody. In many test kits in which s/co ratios are used, the s/co is set between 0.75 and 1.1. The COI is obtained by converting the cut-off value (threshold value) obtained to 1.0 and indexing it. The COI standardizes measurements. As the COI is a ratio to the positive standard, it does not reflect the actual amount of antibody. In many test kits in which COI is used, it is set to 1.0.
20 | Rate of invalid, error and unreturnable results | < 5%
21 | Precision (repeatability/reproducibility) | Not applicable
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Lateral flow immunoassay. The assay can detect antibodies to HCV genotypes 1–6.
24 | Specimen(s) stability | Venous whole blood can be stored at 2–8 °C for 2 days. Serum and plasma can be stored at 2–8 °C for 5 days or at –20 °C or colder for > 5 days.
25 | Specimen(s) volume | 5–100 µL
26 | Type of result | Qualitative
27 | Time to result | ≤ 20 min
28 | End-point stability | ≤ 20 min
29 | Ease of use for point-of-care tests only: number of steps that require precision | ≤ 3 steps
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 3 specimens per hour per operator
31 | Test limitations | The manufacturer should state any test limitations including conditions that have not been validated with the test, e.g. testing population, sample types, users.
32 | Internal quality control | A built-in quality control line that must appear when a specimen is added plays a role of Internal quality control. Absence of the control line invalidates the test. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected from humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected from humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected from humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated.
38 | Transport stability for specimen collection(temperature and humidity) | Venous whole blood can be stored at 2–8 °C for 2 days. Serum and plasma can be stored at 2–8 °C for 5 days or at –20 °C or colder for > 5 days.
39 | Storage stability for specimen collection (temperature and humidity) | Venous whole blood can be stored at 2–8 °C for 2 days. Serum and plasma can be stored at 2–8°C for 5 days or at –20 °C or colder for > 5 days.
40 | In-use stability for specimen collection (temperature and humidity) | Venous whole blood can be stored at 2–8 °C for 2 days. Serum and plasma can be stored at 2–8°C for 5 days or at –20 °C or colder for > 5 days.
41 | Transport stability for controls/calibrators (temperature and humidity) | Not applicable
42 | Storage stability for controls/calibrators (temperature and humidity) | Not applicable
43 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability for controls/calibrators (temperature and humidity) | Not applicable
45 | Shelf-life of kit/reagents upon manufacture (months) | ≤ 12 months from date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | ≤ 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment ªhttps://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time to charge battery | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Displayed parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or Closed system | Not applicable
60 | Multidisease testing capabilities / Test menu availability | Not applicable
Infrastructure requirements
61 | Water requirements | Not required
62 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature will require electricity.
63 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits.
64 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
65 | Kit components – test, reagents, consumables (if relevant) | Test strips or devices or cassette and/or buffer or diluent
66 | Reagent/kit size (number of tests) | ≥ 1 test
67 | Consumables required but not provided in the test kit | For collecting finger-prick or capillary whole blood, lancets are required. If capillary tubes or disposable pipettes are not provided, tips for micropipettes are required. For venous whole blood collection, blood collection tubes and needles are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens, biohazard bags and disposable gloves) are required. If disposable pipettes for venous whole blood are not provided, tips for micropipettes are required. Sodium hypochlorite (bleach) for sterilization before disposal of infectious waste is required.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection or for collecting finger-prick or capillary whole blood is required. For testing, a timer is required. If multiple tests are being performed, more than one timer would be an advantage. If disposable pipettes are not provided, a calibrated micropipette and a micropipette stand are required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Required. The instructions for use submitted must correspond to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
73 | Material safety data sheet | Required. The technical document should provide detailed, comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazardous classification | Potential sources of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This information is usually included in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of in-country technical support | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification status | This assay in a RDT format is within the scope of WHO Prequalification as of March 2023. There is a WHO-prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 11
Antibodies to hepatitis C virus
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | GMDN code/ Term name | 48365/ Hepatitis C virus total antibody IVD, kit, enzyme immunoassay (EIA)
3 | EMDN code/ Term name | W0105020303/ HCV ANTIBODY (TOTAL)
4 | Generic name | Antibodies to hepatitis C virus (anti-HCV) immunoassay
5 | Specific type or variation (optional) | Qualitative assay
6 | Alternative name/s (optional) | Enzyme immunoassay (EIA) for antibodies to hepatitis C virus (anti-HCV)
7 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV antibody, hepatitis, microtitre plate enzyme immunoassay, HCV
Intended use
8 | Detection target | Hepatitis C virus antibodies
9 | Test purpose | To screen donated blood for HCV
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Assay format | Immunoassay (manual method)
12 | Specimen type | Plasma and serum
13 | Testing population | Adolescents and adults
14 | Intended user | Trained laboratory scientist, technologist, technician, specialist
15 | Level of the health-care system | For use in blood-screening laboratories
Performance characteristics
16 | Clinical sensitivity | > 99%
17 | Clinical specificity | > 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Limit of detection | Currently, there are no international units or international reference materials for measuring HCV antibodies, unlike for those of HBV. Two types of unit are used to express measurement results: signal-to-cut-off (s/co) ratios and the cut-off index (COI). The s/co ratio is obtained by dividing the signal strength for optical density of a sample by the signal strength for optical density of an internal cut off. As the s/co ratio is a ratio to the cut off and does not reflect the actual amount of antibody. In many test kits in which s/co ratios are used, the s/co is set between 0.75 and 1.1. The COI is obtained by converting the cut-off value (threshold value) obtained to 1.0 and indexing it. The COI standardizes measurements. As the COI is a ratio to the positive standard, it does not reflect the actual amount of antibody. In many test kits in which COI is used, it is set to 1.0.
20 | Rate of invalid, error and unreturnable results | < 5%
21 | Precision (repeatability/reproducibility) | Intra-assay (reproducibility): CV ≤ 13% Inter-assay (reproducibility at one site): CV ≤ 15%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Enzyme immunoassay with microwells or strips coated with recombinant HCV antigens. The test can detect antibodies to HCV genotypes 1–6 in human serum or plasma.
24 | Specimen(s) stability | At 2–8 °C for ≤ 72 h and at –20 °C for > 72h
25 | Specimen(s) volume | 10–100 µL
26 | Type of result | Qualitative (reactive or non-reactive)
27 | Time to result | 180–240 min
28 | End-point stability | < 60 min
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 40 specimens per day per operator
31 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
32 | Internal quality control | Positive, negative and blank controls are validated each time the kit is used to verify that the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement for controls or calibrators once they are opened should be clearly stated. The instructions should indicate the duration of stability of each control once opened.
38 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for ≥ 72 h and at –20 °C for > 72 h
39 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for ≥ 72 h and at –20 °C for > 72 h
40 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for ≥ 72 h and at –20 °C for > 72 h
41 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
42 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
43 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement for reagents once they are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
45 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument ohysical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time to charge battery | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Displayed parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability/availability of a test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Distilled or deionized water of adequate quality for enzyme immunoassay or enzyme-linked immunosorbent assay. See the instructions for use or package insert of the kit/reagent.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit components – testm reagents, consumables (if relevant) | Coated plate with wells, conjugate and positive and negative controls, wash buffer, substrate concentrate, stop solution and package insert or instructions for use.
66 | Reagent/kit size (number of tests) | ≥ 20 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents that the kits do not contain are required. See the package insert or instructions for use of the kit or reagent.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader/enzyme immunoassay spectrophotometer, an incubator or water bath, a microwell plate or strip washer/enzyme immunoassay washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with hood and a cover for calibrating the micropipettes that can weigh ≤ 0.001 g (1 µL), a micropipette stand, a timer, a plate cover, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | The instructions for use submitted must correspond to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. The certificate should be retained by the testing centre with the dates on which the lot was used.
73 | Material safety data sheet | The technical document should provide detailed, comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | Potential sources of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This information is usually included in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of in-country technical support | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification status | This assay is within the scope of WHO Prequalification as of March 2023. There is a WHO=prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 12 (Auto)
Antibodies to hepatitis C virus |  | This technical specification is for one specific test to be run on an automated analyser. As automated analysers are usually used to run a number of different tests, however, the selection of equipment should also depend on the other tests performed in the facility and not only the test described below.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | GMDN code/ Term name | 48365/ Hepatitis C virus total antibody IVD, kit, enzyme immunoassay (EIA)
3 | EMDN code/ Term name | W0105020303/ HCV ANTIBODY (TOTAL)
4 | Generic name | Antibodies to hepatitis C virus (anti-HCV) immunoassay
5 | Specific type or variation (optional) | Qualitative assay
6 | Alternative name/s (optional) | Chemiluminescent immunoassay for antibodies to hepatitis C virus (anti-HCV) Chemiluminescent enzyme immunoassay for antibodies to hepatitis C virus (anti-HCV)
7 | Keywords (optional) | Hepatitis C virus antibody, anti-HCV, hepatitis, chemiluminescent immunoassay, CLIA, chemiluminescent enzyme immunoassay, CLEIA, HCV
Intended use
8 | Detection target | Hepatitis C virus antibodies
9 | Test purpose | To screen donated blood for HCV
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Assay format | Immunoassay (automated method)
12 | Specimen type | Plasma and serum
13 | Testing population | Adolescents and adults
14 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
15 | Level of the health-care system | For use in blood-screening laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 99%
17 | Clinical specificity | ≥ 99%
18 | Analytical specificity | Assay does not cross react with other human diseases relevant to the local setting or with common interfering substances. If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Limit of detection | Currently, there are no international units or international reference materials for measuring HCV antibodies, unlike for those of HBV. Two types of unit are used to express measurement results: signal-to-cut-off (s/co) ratios and the cut-off index (COI). The s/co ratio is obtained by dividing the signal strength for optical density of a sample by the signal strength for optical density of an internal cut off. As the s/co ratio is a ratio to the cut off and does not reflect the actual amount of antibody. In many test kits in which s/co ratios are used, the s/co is set between 0.75 and 1.1. The COI is obtained by converting the cut-off value (threshold value) obtained to 1.0 and indexing it. The COI standardizes measurements. As the COI is a ratio to the positive standard, it does not reflect the actual amount of antibody. In many test kits in which COI is used, it is set to 1.0.
20 | Rate of invalid, error and unreturnable results | Not applicable
21 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV < 10%, inter-assay precision (reproducibility): CV < 15%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Chemiluminescent immunoassay or chemiluminescent enzyme immunoassay can detect HCV genotypes 1–6.
24 | Specimen(s) stability | At 2–8 °C for 48 h and at –20 °C for > 48 h
25 | Specimen(s) volume | ≤ 200 µL
26 | Type of result | Qualitative
27 | Time to result | ≤ 60 min
28 | End-point stability | Not applicable
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | Low sample volume laboratory Specimens per day per analyser: ≤ 500 (Sample throughput per hour per analyser: ≤ 62 ) Middle sample volume laboratory Specimens per day per analyser: 500–1000 (Sample throughput per hour: 63–125) High sample volume laboratory Specimens per day per analyser: ≥ 1000 (Sample throughput per hour: ≥ 126)
31 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
32 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify that the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected from all heat and light sources
36 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of the analyser's reagent unit or reagent storage
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
38 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 48 h, at –20 °C or colder for > 48 h
39 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 48 h, at –20 °C or colder for > 48 h
40 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 48 h, at –20 °C or colder for > 48 h
41 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
42 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
43 | On-board stability for controls/calibrators (temperature and humidity) | At the temperature of the analyser's measuring unit
44 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from all heat and light sources. Any specific requirement for controls or calibrators once they are opened should be clearly stated. The instructions should indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 9 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months’ shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Depends on the brand or model and whether bench-top or floor-standing
48 | Weight of the device (kg) | Depends on the brand or model and whether bench-top or floor-standing
49 | Power requirements | 100–240 VAC/50–60 Hz, or 200–240 VAC/50–60 Hz
50 | Time to charge battery | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Ideally, an uninterruptible power source
53 | Operating conditions (temperature, humidity and altitude) | Temperature range: 18–30 °C; humidity: < 80%; altitude: ≤ 2000 m
54 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle ware
55 | Displayed parameters | Digital display of test results, electronic and printed readout
56 | Display languages | English or local language
57 | Built-in memory storage capacity | Analyser for low sample volume laboratory: ≥ 800 tests Analyser for middle sample volume laboratory: ≥ 2000 tests Analyser for high sample volume laboratory: ≥ 10 000 tests
58 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programme database for shring information. Capability for data transmission, e.g. via USB, electronic health records and health programme databases.
59 | Open or closed system | Both open and closed systems
60 | Multi-disease testing capabilitt/availability of a test menu | Multi-disease testing
Infrastructure requirements
61 | Water requirements | For analysers with a wet system, deionized water of adequate quality. See the analyser's operational manual. instructions for use or package insert of kit/reagents. For analysers with s dry system, water is not required.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit components – test, reagents, consumables (if relevant) | In general, buffer/diluent, antigen and conjugate. If only cups, trays, cartridges or wells for the test are included in the test kit, other reagents specified by the manufacturer are required. Please see the package insert or instructions for use of the test kit.
66 | Reagent/kit size (number of tests) | ≥ 42 tests
67 | Consumables required but not provided in the test kit | For collecting venous whole blood, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection, alcohol swabs, tissues, permanent markers/pens, biohazard bags and disposable gloves are required. For testing, controls, calibrator and other reagents and consumables specified by manufacturer, disposable gloves, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste and biohazard bags are required. Please see the package insert or instructions for use of the test kit.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, analyser (if specified for the kit/reagent, that specified is required), a timer, micropipettes and a micropipette stand, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
69 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer
Documentation
70 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also include recommended procedures for disposal and any probable hazards to the environment and/or community.
71 | Instructions for use | Instructions for use submitted in the local language must correspond to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. The certificate should be retained by the testing centre with the dates on which the lot was used.
73 | Material safety data sheet | Required. The manufacturer should supply detailed,comprehensive documentation for each method on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | The manufacturer should clearly document potential sources of harm in kit inserts. Classification should be according to the Global Harmonized System, e.g. health, physical or environmental hazard. More detail is usually provided in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification, e.g. simple waste disposal, hazardous waste, incineration etc.
Training, installation and use
76 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range should be assessed.
77 | Requirements for installation and calibration | The manufacturer should be responsible for installing and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
78 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
79 | Warranty | The manufacturer must provide a minimum of 1 year’s warranty from the date of handover; the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during warranty must be specified.
80 | Preventive maintenance | The minimal maintenance expected to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either fixed or based on the number of tests the instrument processes, which will de dictated by the workload.
81 | Corrective maintenance | Specific requirements for calibration and test procedures must be described in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or supplier will describe in detail the maintenance support they will provide.
82 | Type of service contract (including leasing or reagent rental) | Must cover labour, repair, spare parts, shipping and logistics costs and training. Possibility of leasing instrument for smaller devices.
83 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
84 | Availability of in-country technical support | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
85 | Estimated lifespan | 5 years after the start of use/installation on condition that the regular maintenanceand inspection specified in the instruction manual, replacement of maintenance parts and necessary repairs or maintenance measures are performed under normal use and operating conditions.
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification status | This assay is within the scope of WHO Prequalification as of March 2023. There is a WHO-prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 13
Combined antibodies to HCV and HCV core antigen
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2023-03-29
iv | Date of publication | 2023-06-06
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category
2 | GMDN code/ Term name | 48381/ Hepatitis C virus antibody/antigen IVD, kit, enzyme immunoassay (EIA)
3 | EMDN code/ Term name | W0105020308/ HEPATITIS C VIRUS (HCV) ANTIGEN/ANTIBODY COMBO TEST
4 | Generic name | Combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg) immunoassay
5 | Specific type or variation (optional) | Qualitative assay
6 | Alternative name/s (optional) | Enzyme immunoassay (EIA) for combined antibodies to HCV (anti-HCV) and HCV core antigen (HCVcAg)
7 | Keywords (optional) | Hepatitis C virus antibody and core antigen, hepatitis, microtitre plate enzyme immunoassay, HCV
Intended use
8 | Detection target | Hepatitis C virus antibodies and core antigen
9 | Test purpose | To screen donated blood for HCV
10 | Specific disorder or condition | Hepatitis C virus (HCV) infection
11 | Assay format | Immunoassay (manual method)
12 | Specimen type | Plasma and serum
13 | Testing population | Adolescents and adults
14 | Intended user | Trained laboratory scientist, technologist, technician, specialist with expertise
15 | Level of the health-care system | For use in blood-screening laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 99%
17 | Clinical specificity | ≥ 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or with common interfering substances.If some cross-reactivity is present, the manufacturer must state this as a limitation of their product, and decision-makers should consider whether any interference observed will have a significant effect on the performance of the product.
19 | Limit of detection | Currently, there are no international units or international reference materials for measuring HCV antibodies, unlike for those of HBV. Two types of unit are used to express measurement results: signal-to-cut-off (s/co) ratios and the cut-off index (COI). The s/co ratio is obtained by dividing the signal strength for optical density of a sample by the signal strength for optical density of an internal cut off. As the s/co ratio is a ratio to the cut off and does not reflect the actual amount of antibody. In many test kits in which s/co ratios are used, the s/co is set between 0.75 and 1.1. The COI is obtained by converting the cut-off value (threshold value) obtained to 1.0 and indexing it. The COI standardizes measurements. As the COI is a ratio to the positive standard, it does not reflect the actual amount of antibody. In many test kits in which COI is used, it is set to 1.0. On HCV core antigen: < 1700 IU/mLThe COI is a ratio number to the positive standard and does not reflect the actual amount of antibody. HCV core antigen: No data. The limit of detection of the HCV core antigen automated method is ≤ 3 fmol/L.
20 | Rate of invalid, error and unreturnable results | < 5%
21 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): CV ≤ 7%, inter-assay precision (reproducibility): CV ≤ 15%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Enzyme immunoassay, microwells coated with recombinant HCV antigens and anti-HCVcAg monoclonal antibody. The test should be able to detect HCV genotypes 1–3 at least. Ideally, the test can detect HCV subtypes 1–6.
24 | Specimen(s) stability | At 2–8 °C for 72 h, at –20 °C beyond 72 h
25 | Specimen(s) volume | 50–100 µL
26 | Type of result | Qualitative (reactive or non-reactive)
27 | Time to result | 120–180 min
28 | End-point stability | 10–20 min
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | ≥ 40 specimens per day per operator
31 | Test limitations | The manufacturer should state any test limitations, including conditions that have not been validated with the test, e.g. testing population, sample types, users.
32 | Internal quality control | Positive, negative and blank controls should be validated each time the kit is used to verify that the performance of the assay is as qualified. Re-testing should be recommended in the case of invalid results. The package insert should include a description of how how positive, negative and invalid results are read and determined.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is compatible. Ideally, the proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detection of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated. The instructions should indicate the duration of stability of each reagent once opened.
38 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h, at –20 °C beyond 72 h
39 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h, at –20 °C beyond 72 h
40 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h, at –20 °C beyond 72 h
41 | Transport stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
42 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight
43 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected from humidity and direct sunlight. Any specific requirement for controls once opened should be clearly stated. The instructions should indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents upon manufacture (months) | ≥ 12 months for kits/reagents and 6 months for controls/calibrators from date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months shelf-life remaining when product arrives at point of use. Note: Please see p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577)
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time to charge battery | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Displayed parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability/availability of a test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Distilled or deionized water of adequate quality for enzyme immunoassay and enzyme-linked immunosorbent assay. See the instructions for use or the package insert of the kit/reagent.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit components – test, reagents, consumables (if relevant) | Coated plate with wells, conjugate and positive and negative controls, wash buffer, substrate concentrate, stop solution and package insert or instructions for use.
66 | Reagent/kit size (number of tests) | ≥ 96 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues, sodium hypochlorite (bleach) for sterilization before disposal of infectious waste, controls, pipette tips, 96-well dilution plates, controls, calibrator and other reagents not included in the kits are required. See the package insert or instructions for use of the kit or reagent.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. If serum or plasma is used, a centrifuge, a micropipette, a micropipette stand and a stand for containers for plasma or serum are required. As common equipment for venous whole blood collection and preparation of plasma and serum, a container for used needles is required. For testing, a microwell plate or strip reader/enzyme immunoassay spectrophotometer, an incubator or water bath, a microwell plate or strip washer/enzyme immunoassay washer, a vacuum pump, a vortex mixer, calibrated single and multi-channel micropipettes, an electronic scale with hood and a cover for calibrating the micropipettes that can weigh ≤ 0.001 g (1 µL), a micropipette stand, a timer, a plate cover, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | The instructions for use submitted must correspond to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product. This certificate should be retained by the testing center along with the dates that the lot number was in use.
73 | Material safety data sheet | The technical document should provide detailed, comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage or use; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazardous classification | Potential sources of harm. Classification according to the Global Harmonized System, e.g. health, physical or environmental hazard. This information is usually included in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g. simple rubbish or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of in-country technical support | Availability of a local authorized service provider or distributor is recommended if direct procurement from the manufacture was not used.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification status | This assay is within the scope of WHO Prequalification as of March 2023. There is a WHO-prequalified product, and prequalification is recommended. Please see WHO Public Reports for In Vitro Diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective actions | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## Acknowledgement and Purpose
Purpose and aim
The WHO Model List of Essential In-vitro Diagnostics (EDL) is a list of recommended types of tests to support diagnosis of infectious diseases and noncommunicable diseases at each level of the health-care system. The EDL serves as an evidence-based reference for countries to develop their own national lists for selecting and using IVDs. The EDL is also published in an electronic format (eEDL). Technical specifications for IVDs are criteria and baseline requirements to ensure good quality, safety, performance and efficacy. Detailed technical specifications are necessary for countries to select and procure appropriate tests. The EDL do not provide technical specifications necessary for selection and procurement of specific IVDs, however, these technical specifications for IVDs in the EDL have been developed to be published in an electronic format in the eEDL. The purpose of technical specifications for IVDs in the EDL is to advance universal health coverage, address health emergencies and promote healthier populations by helping countries to select and procure appropriate IVD products. Technical specifications for IVDs provide detailed technical information to support decisions on procurement of good-quality, safe IVDs to improve in-vitro diagnostic testing, particularly in low- and middle-income countries.
Acknowledgements
Technical contributions: Nikhil Moorchung (WHO consultant, India), Emma English (WHO consultant, United Kingdom), Jitendra Kumar (WHO consultant, India), Olga Perovic (WHO consultant, South Africa), Ehab Saad (WHO consultant, South Africa), Francis Gabriel Moussy (WHO headquarters), Ana Aceves Capri (WHO headquarters), Naofumi Hashimoto (WHO headquarters), Adriana Velazquez Berumen (WHO headquarters), Philippa Easterbrook (WHO headquarters), Niklas Luhmann (WHO headquarters) and Ute Ströher (WHO headquarters).
Conflicts of interest and declarations of interest
All the WHO consultants involved in preparing these technical specifications provided declarations of interests, and no conflicts were identified.
Funding
Funding and in-kind contributions to this project were provided by the following organizations: • The Bill & Melinda Gates Foundation • The Foreign, Commonwealth and Development Office, United Kingdom • The Economic Research Institute for the Association of Southeast Asian Nations and East Asia
Updating
The document will be updated regularly, every 3 years.

## Process
Process
Technical specifications define the minimum requirements for ensuring the quality, safety and efficacy of the products. Development of technical specifications consists of the six steps listed below.
•Step1: Review existing WHO guidance and specifications for the procurement, donation or lease of good-quality IVDs to develop a common template of generic specifications. The guidance documents analysed were: Guidance for procurement of in vitro diagnostics and related laboratory items and equipment; Guidelines for the diagnosis, prevention and management of cryptococcal disease in HIV-infected adults, adolescents and children: supplement to the 2016 consolidated guidelines on the use of antiretroviral drugs for treating and preventing HIV infection; Guidelines for managing advanced HIV disease and rapid initiation of antiretroviral therapy; and the WHO Model List of Essential In Vitro Diagnostics and WHO internal procurement templates.
•Step 2: WHO commissions a team of external consultants to develop technical specifications for the IVDs listed in the WHO-EDL. Development of specifications comprises: horizon scanning of existing tests and platforms; analysis of products on the market after approval by stringent regulatory authorities; analysis of international, regional and country standards; review of WHO guidance and guidelines for all the diseases included in the EDL; review of peer-reviewed publications, textbooks and published qualitative and quantitative comparative studies of tests and platforms; comparison of test specifications (i.e. accuracy, precision, positive and negative predictive value, likelihood ratios), comparison and review of the package inserts and user manuals of products approved by stringent regulatory authorities; comparison and review of documents and data on the safety and performance of approved products on the websites of stringent regulatory authorities; results of external quality assessments and technical expert opinion. Each template is drafted or reviewed by at least two qualified technical experts. All consultants provided declarations of Interest, and no conflicts were identified.
•Step 3: Internal review of all draft technical specifications by the EDL secretariat. •Step 4: Review of draft technical specifications by relevant disease programmes. •Step 5: 3-week public consultation on the draft technical specifications by posting them on the WHO website and sharing them on relevant networks and with non-State actors in official relations with WHO. •Step 6: Compile all feedback, and finalize the technical specifications before adding them to the electronic EDL.

## Evidence and Input
Evidence and Input
Sources of evidence:
1. Guidance for procurement of in vitro diagnostics and related laboratory items and equipment. Geneva: World Health Organization; 2017 (https://apps.who.int/iris/handle/10665/255577). 2. Guidelines on hepatitis B and C testing. Geneva: World Health Organization; 2017 (https://www.who.int/publications/i/item/978924154998). 3. Guidelines for the care and treatment of persons diagnosed with chronic hepatitis C virus infection. Geneva: World Health Organization; 2018 (https://apps.who.int/iris/handle/10665/273174). 4. Technical specifications series for submission to WHO prequalification-diagnostic assessment; TSS-8 – Immunoassays to detect hepatitis C antibody and/or antigen. Geneva: World Health Organization; 2019(https://apps.who.int/iris/bitstream/handle/10665/327936/9789241516518-eng.pdf?ua=1). 5. Technical specifications series for submission to WHO prequalification-diagnostic assessment; TSS-10 – In vitro diagnostic medical devices used for the qualitative and quantitative detection of hepatitis C RNA. Geneva: World Health Organization; 2023 (https://apps.who.int/iris/bitstream/handle/10665/366372/9789240057487-eng.pdf). 6.Technical specifications series for submission to WHO prequalification-diagnostic assessment; TSS 16 – Hepatitis C rapid diagnostic tests for professional use and/or self-testing, 2021 update. Geneva: World Health Organization; 2022 (https://apps.who.int/iris/bitstream/handle/10665/366073/9789240033757-eng.pdf?sequence=1&isAllowed=y).
7. Morris C, Prescott G, Hockley J. Collaborative study to evaluate the proposed 5th WHO IS for hepatitis C virus (HCV) RNA for nucleic acid amplification technique (NAT)-based assays. ‎Geneva: World Health Organization; 2015 (https://apps.who.int/iris/handle/10665/70782?locale-attribute=pt&). 8. Fryer F, Heath AB, Wilkinson DE, Minor PD, Collaborative Study Group. Collaborative study to evaluate the proposed 4th WHO international standards for hepatitis C virus (‎HCV)‎ for nucleic acid amplification technology (‎NAT)‎-based assays. ‎Geneva: World Health Organization 2011 (https://apps.who.int/iris/handle/10665/70782?locale-attribute=pt&). 9. Nick S, Volkers P, Ross S, Scheiblauer H, Collaborative Study Group. Collaborative study to establish the first WHO IS for detection of hepatitis C virus core antigen. ‎Geneva: World Health Organization; 2014 (https://www.who.int/publications/m/item/WHO-BS-2014.2247).
10. Hepatitis C market memo July 2022. Boston (MA): Clinton Health Access Initiative, Inc.; 2022 (https://chai19.wpenginepowered.com/wp-content/uploads/2022/07/HCV-2022-Market-Memo_vf.df). 11. HCV market intelligence report 2021 and preliminary HBV market insights. Boston (MA): Clinton Health Access Initiative, Inc.; 2021 (https://chai19.wpenginepowered.com/wp-content/uploads/2021/08/Hepatitis-C-Market-Report_2021-FINAL-1.pdf). 12. List of HIV diagnostic test kits and equipment classified according to the Global Fund quality assurance policy. Version 48, 29 March 2023. Geneva: The Global Fund to Fight AIDS, Malaria and Tuberculosis (https://www.theglobalfund.org/media/5878/psm_productshiv-who_list_en.pdf). 13. Hepatitis C diagostic technology landscape, May 2019. Geneva: Unitaid; 2019 (https://unitaid.org/assets/HepC-Dx-Tech-Landscape_May2019.pdf). 14. HCV diagnostic landscape 2015. Geneva: Unitaid; 2015 (https://unitaid.org/assets/UNITAID-HCV_Diagnostic_Landscape-1st_edition-1.pdf). 15. Putting HIV and HCV to the test, 3rd edition. Paris: Médecins sans Frontières; 2017 (https://msfaccess.org/putting-hiv-and-hcv-test-3rd-ed-2017). 16. Alter MJ, Kuhnert WL, Finelli L. Guidelines for laboratory testing and result reporting of antibody to hepatitis C virus. Mortal Morbid Wkly Rep. 2003;52(RR03);1–16 (https://www.cdc.gov/mmwr/preview/mmwrhtml/rr5203a1.htm). 17. Freiman JM, Tran TM, Schumacher SG. HCV core antigen testing for diagnosis of HCV infection: a systematic review and meta-analysis. Ann Intern Med. 2016;165(5):345–55. doi:10.7326/M16-0065.
18. Vetter BN, Reipold EI, Ongarello S,Audu R, Ige FA, Alkhazashvili M et al. Sensitivity and specificity of rapid diagnostic tests for hepatitis C virus with or without HIV coinfection: a multicentre laboratory evaluation study. J Infect Dis. 2022;226(3):420–30. doi:10.1093/infdis/jiaa389. 19. Warkad SD, Song KS, Pal D, Nimse SB. Developments in the HCV screening technologies based on the detection of antigens and antibodies. Sensors. 2019;19(19):4257. doi:10.3390/s19194257. 20. Tang W, Chen W, Amini A, Boeras D, Falconer J, Kelly H et al. Diagnostic accuracy of tests to detect hepatitis C antibody: a meta-analysis and review of the literature. BMC Infect Dis. 2017;17(Suppl_1):695. doi:10.1186/s12879-017-2773-2. 21. Odari EO, Budambula NL, Nitschko H. Evaluation of an antigen-antibody "combination" enzyme linked immunosorbent assay for diagnosis of hepatitis C virus infections. Ethiop J Health Sci. 2014;24(4):343–52. doi:10.4314/ejhs.v24i4.10. 22. Freiman JM, Tran TM, Schumacher SG, White LF, Ongarello S, Cohn J et al. Hepatitis C core antigen testing for diagnosis of hepatitis C virus infection: a systematic review and meta-analysis. Ann Intern Med. 2016;165(5):345–55. doi:10.7326/M16-0065. 23. Tillmann HL. Hepatitis C virus core antigen testing: role in diagnosis, disease monitoring and treatment. World J Gastroenterol. 2014;20(22):6701–6. doi:10.3748/wjg.v20.i22.6701. 24. Moretti M, Pieretti B, Masucci A,Sisti D, Rocchi M, Delprete E. Role of signal-to-cutoff ratios in hepatitis C virus antibody detection. Clin Vaccine Immunol. 2012;19(8):1329–31. doi:10.1128/CVI.00175-12.
25. Toyoda H, Kumada T, Kiriyama S, Sone Y, Tanikawa M, Hisanaga Y et al. Changes in hepatitis C virus (HCV) antibody status in patients with chronic hepatitis C after eradication of HCV infection by interferon therapy. Clin Infect Dis. 2005;40(6):e49–54. doi:10.1086/428128. 26. Nick S, Scheiblauer H. Sensitivities of CE-marked HIV, HCV, and HBsAg assays. J Med Virol. 2007;79:S59–64. doi:0.1002/jmv.20965. 27. Villar LM, Cruz HM, Barbosa JR, Bezerra CS, Portilho MM, de Paula Scalioni L. Update on hepatitis B and C virus diagnosis. World J Virol. 2015;4(4):323–42. doi:10.5501/wjv.v4.i4.323. 28. Lange B, Roberts T, Cohn J, Greenman J, Camp J, Ishizaki A et al. Diagnostic accuracy of detection and quantification of HBV-DNA and HCV-RNA using dried blood spot (DBS) samples – a systematic review and meta-analysis. BMC Infect Dis. 2017;17(Suppl_1):693. doi:10.1186/s12879-017-2776-z. 29. Nick S, Scheiblauer H. Sensitivities of CE-marked HIV, HCV, and HBsAg assays. J Med Virol. 2007;79:S59–64. doi:10.1002/jmv.20965. 30. Dabbs DJ, editor. Diagnostic immunohistochemistry: theranostic and genomic applications, 5th edition. Amsterdam: Elsevier; 2018 (https://www.eu.elsevierhealth.com/diagnostic-immunohistochemistry-9780323477321.html).
37. Fiedler BA, editor. Managing medical devices within a regulatory framework. Amsterdam: Elsevier Science; 2016 (https://www.elsevier.com/books/managing-medical-devices-within-a-regulatory-framework/fiedler/978-0-12-804179-6). 38. Tobin JJ, Walsh G. Medical product regulatory affairs: Pharmaceuticals, diagnostics, medical Devices. Hoboken (NJ): Wiley; 2011 (https://www.wiley.com/en-us/Medical+Product+Regulatory+Affairs:+Pharmaceuticals,+Diagnostics,+Medical+Devices-p-9783527318773). 39. Wong J, Tong RKY, editors. Handbook of medical device regulatory affairs in Asia (second edition). Singapore: Jenny Stanford Publishing; 2018 (https://www.jennystanford.com/9789814774314/handbook-of-medical-device-regulatory-affairs-in-asia-second-edition/). 40. Mantus D, Pisano DJ, editors. FDA Regulatory Affairs, third edition. Boca Raton (FL): CRC Press; 2014 (https://www.routledge.com/FDA-Regulatory-Affairs-Third-Edition/Mantus-Pisano/p/book/9781841849195).