## Speculum, vaginal
THERMAL ABLATION (HAEMOSTASIS) UNIT
i | Version No. | 1
ii | Date of initial version | December 2019
iii | Date of last modification
iv | Date of publication | February 2020
v | Completed / submitted by | WHO
NAME, CATEGORY AND CODING
1 | WHO Category / Code
2 | Generic name | Thermal ablation device
3 | Specific type or variation (optional) | Portable, handheld thermal ablation device Benchtop thermal ablation device
4 | Alternative name/s (optional) | Thermal coagulator; thermocoagulator; cold coagulator; Semm coagulator; thermoablation device, Haemostasis unit
5 | Keywords (optional) | Thermal ablation device; cervical precancer; destruction of abnormal human cervical tissue.
PURPOSE OF USE
6 | Clinical or other purpose | A thermal ablation device is a self-contained, electrically powered medical instrument with a probe tip heated to 100°C designed to destroy tissue of the uterine cervix.
7 | Level of use (if relevant) | Handheld device: health posts, clinics or outreach, hospitals Benchtop device: hospitals or clinics with mains electricity
8 | Clinical department/ward (if relevant) | Family medicine; gynaecology; outpatient clinic; outreach; oncology; obstetrics; surgery; nursing services
9 | Overview of functional requirements | • Probe tip temperature controlled to reach 100°C • Visual and/or audible cues to ensure working temperature reached • Simple and easy to use, appropriate for all levels of care. • Handheld device has added feature of portability with rechargeable batteries.
TECHNICAL CHARACTERISTICS
10 | Detailed requirements | Portable, handheld thermal ablation device OR Benchtop thermal ablation device: • Probe tip temperature controlled to reach 100°C • Visual and/or audible cues to ensure working temperature reached • Simple and easy to use, appropriate for all levels of care. • Handheld device has added feature of portability with rechargeable batteries. • Rated IPX1 (console) and IPX7 (therapy probe and or instrument cable on benchtop)
11 | Displayed parameters | User interface should provide audible and/or visual feedback: • Handheld: light/beep when temperature reached • Benchtop: display to indicate temperature.
12 | User adjustable settings | Benchtop: user can adjust temperature between 60°-120°C
PHYSICAL/CHEMICAL CHARACTERISTICS
13 | Components (if relevant) | • Minimum of 2 probe tips required as follows: o One probe must be flat. o The second probe can be either flat or can have a gentle nipple extrusion not exceeding 5mm (so as to anchor in centre of cervix but not to ablate endocervix). o Probes should not have any sharp edges. o Varying diameters, ranging from 8 mm to 25 mm o Biocompatible, material that will not adhere to cervix o Reusable and thus able to be decontaminated Approximate dimensions of the thermal ablation devices: - Handheld: 5 cm (W), 20 cm (H), 5 cm (D) and weighs <400g - Benchtop: 35 cm (W), 15 cm (H), 10 cm (D) and weighs 3.5 kg Both devices come with power adaptors Recommended extra battery for handheld unit
14 | Mobility, portability (if relevant) | Handheld device Can be easily carried in a case or backpack, with all accessories, all totalling <2kg Benchtop device: Not intended for portability
15 | Raw Materials (if relevant) | N/A
UTILITY REQUIREMENTS
16 | Electrical, water and/or gas supply (if relevant) | • The unit (either benchtop or handheld when charging) is suggested to be connected to a reliable power source • Handheld, battery operated device to have a minimum cumulative runtime of 1 hour on a single charge. • Electrical source requirements (based on country/setting of use): o Amperage: _____________; o Voltage: _______________; o Plug type:______________ .
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
17 | Accessories (if relevant) | • Speculum • Light source • Timer
18 | Sterilization process for accessories (if relevant) | • Probes should withstand repeated cycles of decontamination, done so according to manufacturers’ instructions. • Device handle, charging base, benchtop unit, power supply should be cleaned and decontaminated after each use according to manufacturer’s instructions.
19 | Consumables / reagents (if relevant) | N/A
20 | Spare parts (if relevant) | • Consider extra probes to account for decontamination cycles and patient scheduling. • A backup battery pack is recommended for handheld devices. Note: probes and spares are not interchangeable between devices of different brands and models and can vary in their design and lifetime.
21 | Other components (if relevant) | N/A
PACKAGING
22 | Sterility status on delivery (if relevant) | N/A
23 | Shelf life (if relevant) | N/A
24 | Transportation and storage (if relevant) | Storage area should be clean and dust-free, dry, cool, well-lit, venti- lated and vermin-proof. Store device in original packaging on a shelf or on in a storage cabinet. Devices should be able to withstand storage temperatures ranging from 15°C to 30°C, relative humidity ≤ 60% (non-condensing), and be protected from dripping water.
25 | Labelling (if relevant) | Labelling on the primary packaging should include the name and/ or trademark of the manufacturer and should adhere to the most current version of ISO 15223 – 1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements.
ENVIRONMENTAL REQUIREMENTS
26 | Context-dependent requirements | Environmental conditions vary globally and can be extreme; how- ever, the following are tenable: - Operating temperature: 15°C to 35°C - Operating relative humidity: ≤ 85%, non-condensing It is the responsibility of the procurement body to ensure that the manufacturer’s recommended operation and storage conditions are respected. If the device requires that the operating and/or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure.
TRAINING, INSTALLATION AND UTILISATION
27 | Pre-installation requirements(if relevant) | Handheld: no pre-installation requirements. Benchtop: if necessary, clear instructions/diagrams for assembly/ reassembly must be included.
28 | Requirements for commissioning (if relevant) | N/A
29 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Clinical staff training in cervical precancer lesion treatment guidelines and device use to be provided.
30 | User care(if relevant) | Prior to use: • Inspect for visible damage to device (handle, probes, any connections). • Make sure that no parts are missing or loose. • Make sure that connecting elements between instruments function properly.
WARRANTY AND MAINTENANCE
31 | Warranty | Minimum one year. Specific inclusions and exclusions to be listed. Contact details of manufacturer, supplier and local service agent to be provided.
32 | Maintenance tasks | Handheld device: no user-serviceable parts. Benchtop device: see manual or contact manufacturer for information
33 | Type of service contract | Costs and types of post-warranty service contract available shall be described (if applicable).
34 | Spare parts availability post-warranty | 8 years minimum, starting from installation/commissioning
35 | Software / Hardware upgrade availability | Costs and types of post-warranty service contract available shall be described (if applicable).
DOCUMENTATION
36 | Documentation requirements | • Instructions for use and service manuals to be provided (includ- ing procedures for decontamination) • User language preference prioritized, English is mandatory. • Contact details of manufacturer, supplier and local service agent. • Certificate of calibration and inspection to be provided (if appli- cable). • List to be provided of equipment and procedures required for local calibration and routine maintenance. • List to be provided of common spares and accessories, with part numbers.
DECOMMISSIONING
37 | Estimated Life Span | Handheld device: 7 years (can vary) Benchtop main unit: approximately 10 years if maintenance and service requirements are met.
SAFETY AND STANDARDS
38 | Risk Classification | US FDA: Device Class 2 EU: Class IIa
39 | Regulatory Approval / Certification | Compliance to (where applicable, but not limited to): - National Regulatory Agency/Authority (NRA) requirements compliance - Approval by regulatory body of country of manufacturer (if applica- ble) And at least one of: - FDA 510k clearance (US FDA) - CE mark (EU), with indication of Notifying Body (when applicable) - Other regulatory body in an IMDRF founding member country such as Australia, Canada, or Japan.
40 | International standards | Compliant with active version of the following standards (or equivalent): General manufacturing: • ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes • ISO 14971: Medical Devices - Application of Risk Management to Medical Devices • ISO 15223-1: Medical devices - Symbols to be used with medical device labels, labelling and information to be supplied - Part 1: General requirements Safety & product standards • IEC 60601-1 - Medical electrical equipment - Part 1: General requirements for basic safety and essential performance • IEC 60601-1-2: Medical electrical equipment - Part 1-2 General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests Probe-specific requirements: • ISO 10993-1: Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process • ISO 10993-5: Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity • ISO 10993-10: Biological evaluation of medical devices Part 10: Tests for irritation and skin sensitization Battery-operated only: • IEC 62133 - Secondary cells and batteries containing alkaline or other non-acid electrolytes – Safety requirements for portable sealed secondary cells o Part 1: Nickel o Part 2: Lithium
41 | Regional / Local Standards | Country-specific and regional standards may apply
42 | Regulations | US regulations: 21 CFR part 820 Quality System Regulation EU regulations: European Commission Regulation (EU) No. 2017/745 (replacing 93/42/EEC)