## Foglio1
Pulse oximeter - fingertip
Published on: 19 November 2020 - Version: 2
Version 1: September 2019 (WHO-UNICEF)
Technical Specification | Minimum requirements
General technical requirements | SpO2 and pulse rate monitor integrated into finger/toe clip. Configurations required to apply to adults and children, and all skin pigmentations. Suitable for spot check. SpO2 detection to include the range: 70–99%. SpO2 resolution: 1% or less. SpO2 accuracy (in the range at least 70–99%): within ± 3%. If equipment is capable of a wider SpO2 detection range, the accuracy over that wider range shall be stated. Pulse rate detection to include the range: 30–240 bpm. Pulse rate resolution: 1 bpm or less. Pulse rate accuracy within ± 3 bpm. Suitable for detection in low perfusion conditions (as per ISO 80601-2-61, test method must be described). Design must enable use in demanding environments, e.g. shock, vibration and free fall tests as per tests in ISO 80601-2-61. Available probe sizes must accommodate finger/toe thicknesses at least including the range 8–25 mm. Automatic correction for movement, ambient light artefacts (as per ISO 80601-2-61, test method must be described). Display shows % SpO2, pulse rate, signal quality, sensor error or disconnect and low battery status. Enclosure to have ingress protection level IPX2 or better. Any aspects of usability as per IEC 62366-1 must be described. Disinfectable with hospital grade detergents. Automatic power-off. Hours of continuous use, or number of tests, per battery set shall be stated. Batteries must allow at least 2500 spot checks calculated at 30 s per spot check, or at least 12 hours of operation, or better. Operated by internal battery. If rechargeable, batteries may be charged via USB connector or by external AC charger. Rechargeable batteries are preferred.
Displayed parameters | SpO2, pulse rate, battery and system status and preferably signal quality.
Alarms | Audible and visual alarms for sensor error or disconnected, system errors, low battery. Audible and visual alarms for low/high saturation and pulse rate.
Consumables | Rechargeable and/or non rechargeable batteries: 2 sets
Accessories | Battery charger (AC or USB if relevant): 1 per equipment. Replacement flexible cover for patient finger contact (if removable): 2 per equipment. Carry case and/or lanyard.
Documentation requirements (English language mandatory) | User: manuals, hard and soft copies, in English (mandatory) and other languages (preferable). Certificate of calibration and inspection (other means of assurance may be considered). Troubleshooting separate manual or as part of the user manual.
Primary packaging label | Name and/or trademark of the manufacturer. Electrical power input requirements (voltage, frequency and socket type) and safety use and storage (keep away from oil, grease and petroleum-based or flammable products as well as smoking or open flames). Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: IEC 60601-1 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-1 Medical electrical equipment – Part 1-1: General requirements for safety – Collateral standard: Safety requirements for medical electrical systems. IEC 60601-1-2 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. ISO 80601-2-61 Medical electrical equipment – Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment. ISO/IEEE 11073-10404 Health informatics – Personal health device communication – Part 10404: Device specialization – Pulse oximeter (if capacity for data connection to a computer is included). IEC 60068-2-31 Environmental testing – Part 2-31: Tests –Test Ec: Rough handling shocks, primarily for equipment-type specimens. IEC 62366-1 Medical devices – Part 1: Application of usability engineering to medical devices. IEC 62133 – Secondary cells and batteries containing alkaline or other non-acid electrolytes – Safety requirements for portable sealed secondary cells. Part 1: Nickel, Part 2: Lithium.
Warranty | 2 years with regards efficiency and quality of the product. Availability of accessories and spare parts for at least 2 years.