## Kangaroo, chair, mechanical
Chair, kangaroo care, mechanical
i | Version No.
ii | Date of initial version
iii | Date of last modification | October 2023
iv | Date of publication
v | Completed / submitted by | WHO working group
Name, category and coding
1 | WHO Category / Code | Under development
2 | Generic name | Recliner, Kangaroo
3 | Specific type or variation (optional) | This is not a medical device
Purpose of use
4 | Clinical or other purpose | Kangaroo, chair with accessories; mainly intended for immediate kangaroo care (iKMC) processes. Mechanical and manually adjustable comfortable beds/chairs for the mothers, to maintain an upright or semi-recumbent position for resting and sleeping with offering KMC
5 | Level of use (if relevant) | General outpatient and outreach services for primary care; First-level (district) hospital services; Second-level and third-level hospital services and specialized outpatient services
6 | Clinical department/ ward(if relevant)
7 | Overview of functional requirements | Mechanical manually adjustable. Mobile but majorly stationary in place.
Technical characteristics
10 | Detailed requirements | Ability to fully recline. Seat height at least 18", seat depth at least 20". Synchronized head-foot rest. At least with 75mm comfort foam. Antimicrobial vinyl material. Safe working load of at least 250Kg. Twin locking castors (at least 100mm) with 2 lockable brakes. Hoist clearance at least 130mm.
11 | Displayed Alarms | N/A
12 | User adjustable settings | N/A
Physical/chemical characteristics
13 | Components(if relevant) | N/A
14 | Mobility, portability(if relevant) | Able to move around designated areas.
15 | Raw Materials(if relevant) | High resistance to corrosion (tropical environment). Frame, side panels, base, has smooth finishing allowing for easy cleaning preferred. Caster wheel: single or double wheels, mold-on type, non-hooded (for easy maintenance)
Utility requirements
16 | Electrical, water and/or gas supply (if relevant) | N/A
Accessories, consumables, spare parts, other components
17 | Accessories (if relevant) | Basic IkMC (immediate Kangaroo Mother care) information kit. Set of Pillows.
18 | Sterilization process for accessories (if relevant) | Reusable medical furniture should be disinfected with apropiate solution according to manufacturer, cleaned in any health facilities as appropriate.
19 | Consumables / reagents (if relevant) | N/A
20 | Spare parts (if relevant) | N/A
21 | Other components (if relevant) | N/A
Packaging
22 | Sterility status on delivery (if relevant) | N/A
23 | Shelf life (if relevant) | N/A
24 | Transportation and storage (if relevant) | Capable of being transported/stored continuously in ambient temperature of -30˚C to +60˚C and relative humidity of 15 to 95%. Capable of use continuously in ambient temperature of 0˚C to +40˚C and relative humidity of 15 to 95%
25 | Labelling (if relevant) | Labeling on the primary packaging to include: Name and/or trademark of the manufacturer. Information for particular storage conditions (temperature, pressure, light, humidity). Manufacturer’s product reference Manufacture and expiry date, lot number Main characteristics (dimensions, capacity) Recommended storage conditions, Information for handling
Environmental requirements
26 | Context-dependent requirements | N/A
Training, installation and utilization
27 | Pre-installation requirements(if relevant) | Not required.
28 | Requirements for commissioning (if relevant) | N/A
29 | Training of user/s (if relevant) | World Health Organization‎, Kangaroo mother care: a practical guide (https://apps.who.int/iris/handle/10665/42587)
30 | User care(if relevant) | User manual for handling, instructions for disinfection and basic maintenance should be provided.
Warranty and maintenance
31 | Warranty | At least 2 years
32 | Maintenance tasks | Cleaning procedures. Stability control.
33 | Type of service contract | N/A
34 | Spare parts availability post-warranty | N/A
35 | Software / Hardware upgrade availability | N/A
Documentation
36 | Documentation requirements | The instruction manual should include: assembly instructions, standard operating procedure for using the product, day-to-day maintenance, a maintenance schedule, storage condition, cleaning, troubleshooting, safety details, warranty details with a scope, and postal, electronic, and telephone contacts of authorized dealer/country representative as appropriate. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
Decommissioning
37 | Estimated Life Span | N/A
Safety and standards
38 | Standards, for the manufacturer and the equipment | N/A
39 | Regulatory Approval / Certification | Regulatory approval/marketing license.
40 | International standards | N/A