## Foglio1
Ventilator for sub-acute care
Published on: 19 November 2020 - Version: 1
Technical Specification | Minimum requirements
General technical requirements | Medical air compressor or turbine in-built, with inlet filter. Possibility for using external low-pressure oxygen (approx. 20 psi) as source (preferable). If oxygen, high pressure input port (> 35 psi [2.4bar]). Oxygen-air mixture accuracy of 4%. Oxygen conserve feature (preferable). Internal function testing/leak testing. Event log for errors traceability (preferable). At least IP21 degree of protection to the harmful ingress of water (higher preferable). Capability to work with dual-limb breathing circuits. Capability to connect to an active humidifying system.
Ventilation modes | Non-invasive ventilation. It must include at least one mandatory and invasive ventilation mode. Pressure control ventilation (PCV). Volume control ventilation (VCV). Pressure support ventilation (PSV). Synchronized intermittent mandatory ventilation (SIMV) (preferable). Pressure regulated volume control (PRVC) (or similar preferable).
Monitored and controlled parameters (by user) | Air and externally supplied oxygen mixture ratios fully controllable. FiO2: 21–100%. Tidal volume: 50–1000 mL (preferable). Inspiratory pressure: 0–40 cmH2O. I:E ratio. RR: 10–60 breaths/min, minimum. PEEP: at least 0–20 cmH2O.
Displayed parameters (colour and graphics preferable) | Display with brightness and contrast control capabilities and/or any additional and equivalent technological solution to improve the readability in low ambient light and in sunlilght. Real-time scalar waveforms for flow, volume and pressure, at least two simultaneously. Status indicators for ventilator mode, battery status, patient data, alarm settings. Airway pressures (peak, mean and PEEP). Tidal volume (expired). Minute volume (expired). I:E ratio. Inspiration and expiration times. RR. FiO2. Occlusion pressure detection. Air and oxygen pressure. Spontaneous ventilation. Leak. Spontaneous minute volume (preferable).
Alarms, related to gas delivered (adjustable, visual and audible) | High/low FiO2. High/low minute volume. High/low inspiratory pressure. Breathing circuit disconnect. Low minute volume (preferable). Apnoea.
Alarms, related to equipment operation (visual and audible) | Gas supply failure. Power failure. Self-diagnostics failure alarm. Low battery.
Consumables, labelled “single use” (included and mentioned in a disaggregated list) | Single-limb breathing circuit for adult with standard connector with 22 mm outside diameter (at least 60). Single-limb breathing circuit for paediatric (if needed) with standard connector with 22 mm outside diameter (at least 40). Double-limb breathing circuits for adult with standard outlet/inlet connectors with 22 mm outside diameter (preferable) (at least 60). Double-limb breathing circuits for paediatric (if needed) with standard outlet/inlet connectors with 22 mm outside diameter (preferable) (at least 40). Non-invasive ventilation masks, for adult (at least 50 in different sizes). Non-invasive ventilation masks, for paediatric (if needed) (at least 12). Viral/bacterial filters 99.99% efficiency minimum; inspiratory and expiratory as applicable (at least 180). Exhaled gas filter, when applicable. Heat and moisture exchanger filters (HMEF) (at least 90).
Reusable accessories (included and mentioned in a disaggregated list) | Single-limb breathing circuit with standard connector with 22 mm outside diameter. Double-limb breathing circuits with standard outlet/inlet connectors with 22 mm outside diameter (preferable). Exhalation valve (if applicable). CO2 sensors (preferable). Standard hoses and connectors (i.e. DISS/NIST as applicable for the country) for oxygen wall outlets and cylinder, as required to operate. Compatible active humidifier provided.
Spare parts (included and mentioned in a disaggregated list) | 1 year’s spare parts kit as per preventive maintenance programme, including exhalation valves and also oxygen sensors (if applicable).
Portability | Preferable: portable equipment with mechanical strength to lever rough handling. Alternative option: mounted on a mobile cart with at least 2 castors fitted with breaks.
Power supply (voltage, frequency and plug vary across the countries) | Operates from AC power electric line: 100–240 V AC ± 10% / 50–60 Hz ± 10%. In-built rechargeable battery. Automatic switch from AC power electric-line mode to battery operating mode and vice versa. Continuous in battery operating mode with standard ventilation not less than 4 hours. Total re-charging time not greater than 6 hours.
Documentation (included) | Instruction for use; service manual and product information to be provided in English language, at least.
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent, (including the technical tests for safety and performance from accredited laboratory or third party), for (latest version recommended but compliance to previous standards versions could be accepted): IEC 60601-1:2005 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-2:2014 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. ISO 10651-5:2006 Lung ventilators for medical use – Particular requirements for basic safety and essential performance – Part 5: Gas-powered emergency resuscitators. ISO 80601-2-74:2017 Medical electrical equipment – Part 2-74: Particular requirements for basic safety and essential performance of respiratory humidifying equipment. ISO 18562-1:2017 Biocompatibility evaluation of breathing gas pathways in healthcare applications – Part 1: Evaluation and testing within a risk management process (if applicable). ISO 20789:2018 Anaesthetic and respiratory equipment – Passive humidifiers (if applicable).
Warranty | Minimum 2 years. Availability of accessories, consumables and spare parts for at least 2 years.