## Speculum, vaginal
CRYOSURGICAL UNIT, GYNAECOLOGICAL
i | Version No. | 1
ii | Date of initial version | December 2019
iii | Date of last modification
iv | Date of publication | February 2020
v | Completed / submitted by | WHO
NAME, CATEGORY AND CODING
1 | WHO Category / Code | Cryosurgery units, XD9LJ9
2 | Generic name | Cryosurgical Unit, Gynaecological
3 | Specific type or variation (optional) | N/A
4 | Alternative name/s (optional) | Cryotherapy Unit; Cryoprobes; Cryostats; Cryo Units; CSU; Probes,
5 | Keywords (optional) | surgical, gynaecology, dermatology, cryotherapy, cryocautery, cryosurgery, “cryo”
PURPOSE OF USE
6 | Clinical or other purpose | Cryosurgical unit with a gas (carbon dioxide,CO2, or nitrous oxide, N2O) cooled or electrically-cooled cryoprobe intended to destroy tissue of the uterine cervix with the application of extreme cold.
7 | Level of use (if relevant) | Health posts, health centres, specialty clinics, hospitals
8 | Clinical department/ward (if relevant) | Family medicine; gynaecology; outpatient clinic; outreach; oncolo- gy; obstetrics; surgery; nursing services
9 | Overview of functional requirements | The metal probe head (cryotip) makes contact with epithelium to destroy tissue.
TECHNICAL CHARACTERISTICS
10 | Detailed requirements | • Cryosystems include the hand-held unit with use-specific cryotips: • Temperature at probe edge shall be no greater than -20 °C (can be colder, NOT warmer) • If gas-cooled: - either carbon dioxide (CO2) or nitrous oxide (N2O) is used - must be a closed system, whereby the cryogen circulates through probe head, and then back through shaft. • If electrically-cooled, an ethanol-based solution and electricity are necessary.
11 | Displayed parameters | N/A
12 | User adjustable settings | Pressure gauge (for gas-cooled units), cryogun trigger
PHYSICAL/CHEMICAL CHARACTERISTICS
13 | Components (if relevant) | • Minimum of 2 cryotips are required: - One cryotip must be flat. - The second cryotip can be either flat or can have a gentle nipple extrusion not exceeding 5mm (so as to anchor in centre of cervix but not to ablate endocervix). - Cryotip diameters ranging from 17 mm to 23 mm - Biocompatible, material that will not adhere to cervix - Reusable and thus able to be decontaminated • The overall length of the cryoshaft and cryotip assembly should be between 170 and 200 mm. Gas tank, pressure gauge, handheld unit with probe, scavenging/ suction system if using N2O
14 | Mobility, portability (if relevant) | Can be portable or handheld.
15 | Raw Materials (if relevant) | N/A
UTILITY REQUIREMENTS
16 | Electrical, water and/or gas supply (if relevant) | For gas-based: N2O or CO2 The N2O units should only be used with scavenging ability. For electrically cooled: • The unit is suggested to be connected to a continuous, reliable power source • Electrical source requirements (based on country/setting of use): o Amperage: _____________ o Voltage: _______________ o Plug type: ______________
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
17 | Accessories (if relevant) | • Speculum • Light source • Timer
18 | Sterilization process for accessories (if relevant) | • Probes should withstand repeated cycles of decontamination, done so according to manufacturers’ instructions. • Device handle, charging base or gas tank and accessories (whichever applicable) to cleaned and decontaminated after each use according to manufacturer’s instructions.
19 | Consumables / reagents (if relevant) | Gas-based: N2O or CO2 gas Electrically cooled: ethanol-based solution
20 | Spare parts (if relevant) | Consider extra probes to account for decontamination cycles and patient scheduling. Note: probes and spares are not interchangeable between devices of different brands and models and can vary in their design and lifetime.
21 | Other components (if relevant) | N/A
PACKAGING
22 | Sterility status on delivery (if relevant) | N/A
23 | Shelf life (if relevant) | N/A
24 | Transportation and storage (if relevant) | Storage area should be clean and dust-free, dry, cool, well-lit, ventilated and vermin-proof. Store device in original packaging on a shelf or on in a storage cabinet. Devices should be able to withstand storage temperatures ranging from 15°C to 30°C, relative humidity ≤ 85% for gas-based systems or ≤ 60% for electrically cooled systems (both non-condensing), and be protected from dripping water. For compressed gases in cylinders (for gas-based system), both storage and transport can pose as a risk. Transport and storage shall adhere to local regulations; however, gases should never be stored at temperatures in excess of 30°C.
25 | Labelling (if relevant) | Labelling on the primary packaging should include the name and/ or trademark of the manufacturer and should adhere to the most current version of ISO 15223 – 1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements.
ENVIRONMENTAL REQUIREMENTS
26 | Context-dependent requirements | Environmental conditions vary globally and can be extreme; howev- er, the following are tenable: - Operating temperature: 15°C to 35°C - Operating relative humidity: ≤ 85%, non-condensing It is the responsibility of the procurement body to ensure that the manufacturer’s recommended operation and storage conditions are respected. If the device requires that the operating and/or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure.
TRAINING, INSTALLATION AND UTILISATION
27 | Pre-installation requirements(if relevant) | N/A
28 | Requirements for commissioning (if relevant) | Supply of either a gas (N2O or CO2) or ethanol-based solution a requisite for operations
29 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Clinical staff training in cervical precancer lesion treatment guidelines and device use to be provided.
30 | User care(if relevant) | • The cryosurgical unit shall be decontaminated between patients. • Probes should be inspected regularly for mechanical integrity.
WARRANTY AND MAINTENANCE
31 | Warranty | Minimum one year. Specific inclusions and exclusions to be listed. Contact details of manufacturer, supplier and local service agent to be provided.
32 | Maintenance tasks | • Routine check of probe and device mechanical integrity recommended (refer to user manual). • List shall be provided of equipment and procedures required for routine inspection. • Advanced maintenance tasks required are not recommended to be performed in the hospital due to safety concern.
33 | Type of service contract | Costs and types of post-warranty service contract available shall be described (if applicable).
34 | Spare parts availability post-warranty | 8 years minimum, starting from installation/commissioning
35 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
36 | Documentation requirements | • Instructions for use and service manuals to be provided (including procedures for decontamination) • User language preference prioritized, English is mandatory. • Contact details of manufacturer, supplier and local service agent. • Certificate of calibration and inspection to be provided (if applicable). • List to be provided of equipment and procedures required for local calibration and routine maintenance. • List to be provided of common spares and accessories, with part numbers.
DECOMMISSIONING
37 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
38 | Risk Classification | US FDA: Device Class 2 EU: Class IIa
39 | Regulatory Approval / Certification | Compliance to (where applicable, but not limited to): - National Regulatory Agency/Authority (NRA) requirements compliance - Approval by regulatory body of country of manufacturer (if applicable) And at least one of: - FDA 510k clearance (US FDA) - CE mark (EU), with indication of Notifying Body (when applicable) - Other regulatory body in an IMDRF founding member country such as Australia, Canada, or Japan.
40 | International standards | Compliant with active version of the following standards (or equivalent): General manufacturing: • ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes • ISO 14971: Medical Devices - Application of Risk Management to Medical Devices • ISO 15223-1: Medical devices - Symbols to be used with medical device labels, labelling and information to be supplied - Part 1: General requirements Safety and product standards for electrically-cooled systems: • IEC 60601-1 - Medical electrical equipment - Part 1: General requirements for basic safety and essential performance • IEC 60601-1-2: Medical electrical equipment - Part 1-2 General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests Safety and product standards for gas-based systems: • ISO 21969 High-pressure flexible connections for use with medical gas systems Probe-specific requirements: • ISO 10993-1: Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process • ISO 10993-5: Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity • ISO 10993-10: Biological evaluation of medical devices Part 10: Tests for irritation and skin sensitization
41 | Regional / Local Standards | Country-specific and regional standards may apply
42 | Regulations | US regulations: 21 CFR part 820 Quality System Regulation EU regulations: European Commission Regulation (EU) No. 2017/745 (replacing 93/42/EEC)