## Foglio1
Oxygen therapy flowmeter, dial-type (The specification detailed is for general reference only and it refers to the "Flowmeter, Thorpe tube for oxygen 0-15 L/min")
Published on: 19 November 2020 - Version: 2
Version 1: September 2019 (WHO-UNICEF)
Technical Specification | Minimum requirements
General technical requirements | Thorpe tube flowmeter, to measure and regulate the flow of medical gas is composed of inlet and outlet ports, a regulator, a valve and a clear tapered measuring tube. It is suitable for connection with various medical gas sources, such as a centralized system, cylinders or compressors. DISS style inlet and outlet, or if required by end-user other international standard fittings, e.g. 1/8 inch FNPT female, 3/8 inch BSP female, UNI EN 737, DIN, DISS, AFNOR, Ohmeda, Chemtron, Puritan Bennet, Schrader. Transparent, clearly readable and graduated (metric system) column, shatter-resistant polymer certified for medical use. Clearly visible graduation, 270 or more degrees of visibility. Needle valve and body constructed of brass or aluminium. Calibrated at 345–380 kPa (3.4–3.8 bar, 50–55 psi) inlet gauge pressure. Inlet gauge pressure (nominal) > 380–413 kPa (3.8–4.1 bar, 55–60 psi), peak gauge inlet pressure 690 kPa (6.9 bar, 100 psi). Pressure-compensated design to give specified accuracy for whole range of input pressures. Minimum flow rate to be zero, i.e. fully closed. Maximum flow rate when fully open to be stated. Anti-slip knob. Available in international ISO and ANSI colour-coding systems for oxygen and medical air. Disinfectable with hospital grade detergents. Availability of various outlet adapters (tubing nipples/Christmas trees), with ISO, ANSI and generic colour-coding and suitable for all international standard outlet fittings, including (but not limited to) threaded, non-threaded, 6 mm barbed and 9/16 inch UNF female thread for oxygen and medical air. Brass/steel/aluminium/polymers/hard plastic body and valve, certified for medical use. Material: polypropylene, polycarbonate, acrylic or transparent equivalent biocompatible plastic/polymer certified for medical use, unbreakable or shatter resistant, for the column.
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity). Gas type, calibration temperature and pressure should be specified on the label.
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: Colour coding ISO or ANSI for medical gases. Conforms to ISO, NFPA, and/or CGA standards, and/or UL or CSA approved. ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen. ISO 15002 Flow-metering devices for connection to terminal units of medical gas pipeline systems. ISO 18562 Biocompatibility evaluation of breathing gas pathways in healthcare applications. ISO 10524 Pressure regulators for use with medical gases. ISO 18082 Anaesthetic and respiratory equipment – Dimensions of non-interchangeable screw-threaded (NIST) low-pressure connectors for medical gases. ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements. ISO 5359 Low-pressure hose assemblies for use with medical gases. ISO 32 Colour coding for medical gases.
Warranty | 2 years with regards efficiency and quality of the product. Availability of accessories and spare parts for at least 2 years.