## Foglio1
Tubing, medical gases
Published on: 19 November 2020 - Version: 2
Version 1: September 2019 (WHO-UNICEF)
Technical Specification | Minimum requirements
General technical requirements | Medical grade tubing (interior diameter 5 mm) used for the circulation of medical gases. Translucent, flexible and soft plastic tubing of uniform diameter. Anti-kink tubing, non-permanent deformation if kinked or bent too tight. Oxygen and air/oxygen mixture compatibility, as per ISO 15001. Round shape section, internal diameter range 3–5 mm (1/8–3/16 inch), compatible with standard 6 mm barbed fitting. Length: 25 m (preferable). Connectors/adapters for oxygen therapy/respiratory support circuits (preferable).
Material | Rubber or soft plastic tubing, semi-rigid, PVC or other material, FDA Title 21/USP VI compliant and certified for medical use, hardness > 60 Shore A (ASTM D-2240).
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: Conforms to ISO, NFPA and/or CGA standards, and/or UL or CSA approved. ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen. ISO 18562 Biocompatibility evaluation of breathing gas pathways in healthcare applications. ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements. ISO 18190 Anaesthetic and respiratory equipment – General requirements for airways and related equipment. ISO 18562-1 Biocompatibility evaluation of breathing gas pathways in healthcare applications – Part 1: Evaluation and testing within a risk management process. ISO/DIS 17256 Anaesthetic and respiratory equipment — Respiratory therapy tubing and connectors. ISO 5359:2014 Anaesthetic and respiratory equipment — Low-pressure hose assemblies for use with medical gases.