## Foglio1
Ventilator medical, child/newborn, with Continuous Positive Airway Pressure (CPAP) and accessories (The specification detailed below is for general reference only and it refers to adult/paediatric device NOT dedicated for newborn )
Published on: 19 November 2020 - Version: 3
Version 2: 18/06/2014 (modification of the initial version)
Version 1: 13/06/2012 (initial version)
Technical Specification | Minimum requirements
General technical requirements | Maintains continuous positive pressure in airway. Easy to operate user interface, numbers and displays to be clearly visible. Leakage compensation capability (preferable). In-built air compressor or turbine. Oxygen inlet. Capability to connect to an active humidifier system (preferable). Noise level < 35 dB at mid pressure range. Expiratory relief features that reduce the pressure slightly at the end of each breath to make it easier for the patient to exhale (preferable). Pressure ramp option that starts pressure at low level and slowly increases over a period (preferable). All parts withstand high disinfection procedures. Inspiration trigger for auto start. Class I or Class II or internally powered. Protection IP21 required (IP22 preferable).
Ventilation modes | Non-invasive CPAP
Monitored and controlled parameters (by clinical user) | FiO2: 21–100% (preferable). Pressure: 4–20 cmH2O.
Displayed parameters (colour and graphics preferable) | Display easily readable in low ambient light and sunlight. Pressure: cmH2O. FiO2 (%) (preferable). Flow (preferable). Air leak (%) (preferable). RR (preferable).
Alarms, related to gas delivered (visual and audible) | High/low pressure and/or minute ventilation. High/low oxygen (preferable). Breathing circuit disconnection.
Alarms, related to equipment operation (visual and/or audible) | Lack of water (preferable). System failure. Air filter to be replaced. Power failure (preferable). Low battery (preferable).
Consumables, labelled “single use” (included and mentioned in a disaggregated list) | Inlet bacteria filters, if applicable (at least 180). Expiratory filters, high efficiency. Full face mask with tubing (for paediatric and universal fit for adult), alternative oral-nasal mask for adult and paediatric with tubing (at least 12 adult in different sizes). Helmet for adult and paediatric with tubing (preferable).
Accessories, reusable (included and mentioned in a disaggregated list) | Full face mask with tubing (for paediatric and universal fit for adult), alternative oral-nasal mask for adult and paediatric with tubing (at least 2 adult and 2 paediatric); withstands high-level disinfection and sterilization. Helmet for adult and paediatric with tubing (preferable); withstands high-level disinfection and sterilization (preferable). Humidifier accessory if not integrated. Standard hoses and connectors (i.e. DISS/NIST, barb, as applicable for the country) for oxygen wall outlets and cylinder, as required to operate. Mains power cable ≥ 2 m, as required to operate.
Spare parts (included and mentioned in a disaggregated list) | 1 year’s spare parts kit as per preventive maintenance programme (or as part of a service level agreement).
Portability | Portable equipment with mechanical strength to lever rough handling.
Power supply (voltage, frequency and plug vary across the countries) | Operates from AC power electric line: 100–240 V AC ± 10% / 50–60 Hz ± 10%. In-built rechargeable battery (preferable); if the equipment does not have an internal battery, an external battery or uninterruptible power supply should be included to provide battery back up in the case of AC power failure. Automatic switch from AC power electric-line mode to battery operating mode and vice versa, if applicable.
Documentation (included) | Instruction for user and clinical guidance; service manual and product information to be provided in English language, at least.
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent, (including the technical tests for safety and performance from accredited laboratory or third party), for (latest version recommended but compliance to previous standards versions could be accepted): IEC 60601-1 Medical electric equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-2:2014 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. ISO 80601-2-70 Medical electrical equipment – Part 2-70: Particular requirements for basic safety and essential performance of sleep apnoea breathing therapy equipment. ISO 80601-2-80 Medical electrical equipment – Part 2-80: Particular requirements for basic safety and essential performance of ventilatory support equipment for ventilator insufficiency. ISO 60601-1-1-8 Medical electrical equipment – Part 1-8: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems. If applicable, for the accessories and consumables: ISO 18562-1:2017 Biocompatibility evaluation of breathing gas pathways in healthcare applications – Part 1: Evaluation and testing within a risk management process. ISO 80601-2-74:2017 Medical electrical equipment – Part 2-74: Particular requirements for basic safety and essential performance of respiratory humidifying equipment. ISO 17510:2015 Medical devices – Sleep apnoea breathing therapy – Masks and application accessories.
Warranty | Minimum 2 years. Availability of accessories, consumables and spare parts for at least 2 years.