## Foglio1
Laryngoscope, newborn, set (The specification detailed below is for general reference only and it refers to the "ILaryngoscope, FO, diameter 28 mm, with blades" NOT specifc for NEWBORN)
i | Version No. | 2
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2020-07-22
iv | Date of publication | 2020-11-19
v | Completed / submitted by | WHO working group
Name, category and coding
1 | WHO Category / Code | Under development
2 | Generic name | Laryngoscope (rigid)
3 | Specific type or variation (optional) | rigid, curved/straight
4 | GMDN name | Rigid intubation laryngoscope, reusable
5 | GMDN code | 15076
6 | GMDN category | 02 Anaesthetic and respiratory devices, 09 Reusable devices
7 | UMDNS name | Laryngoscopes, Rigid
8 | UMDNS code | 15-076
9 | UNSPS code (optional) | N/A
10 | Alternative name/s (optional) | Intubation laryngoscope; Laryngoscope, rigid
11 | Alternative code/s (optional) | N/A
12 | Keywords (optional) | anaesthesia, anaesthesia, endotracheal, resuscitation, intubation
13 | GMDN/UMDNS definition (optional) | A handheld device intended to be used by anaesthesia/emergency service personnel to manipulate the tongue, preventing it from obstructing the oropharynx and enabling a clear view of the trachea for the insertion of an endotracheal (ET) tube prior to the delivery of inhalation anaesthesia and/or ventilation. It has a handle containing batteries to power its light (a small built-in light bulb or fibre optic light) for airway illumination, and a curved or straight blade of various designs and lengths that can be hinged/interchanged or integral. Some types can be magnetic resonance imaging (MRI) compatible. This is a reusable device to improve respiratory status of a patient, and to help in the treatment evaluation of patients suffering from chronic respiratory disorders (e.g. asthma, emphysema).
14 | EMDN code | Z12021003
15 | EMDN nomenclature | LARYNGOSCOPES
Purpose of use
16 | Clinical or other purpose | To manipulate the tongue and enable a clear view of the trachea for surgical/mechanical ventilation/ intubation procedures.
17 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital
18 | Clinical department/ ward(if relevant) | Outpatients, ear, nose and throat (ENT), operating theatre, emergency room (ER), intensive care unit (ICU)
19 | Overview of functional requirements | A light source on or via the blade illuminates the larynx to allow viewing and tube passage. The unit is handheld with internal batteries and has interchangeable, rigid blades of different sizes.
Technical characteristics
20 | Detailed requirements | LED technology light bulb. Fiber Optic (FO) direct transmission of the light. Equipment provided with compatible stainless steel autoclavable reusable set of blades with illumination included.
21 | Displayed parameters | N/A
22 | User adjustable settings | User shall be able to vary intensity of light in addition to powering on/off.
Physical/chemical characteristics
23 | Components(if relevant) | The system/equipment provided is composed at least by the following components/parts: handheld unit, single piece (with external diameter preferable not less than 25 mm); on/off switch to be robust and easy to use; external material to be non-ferrous; compatible set of blades; protective case for clean storage and safe transport. Unit case should be hard and splashproof. There must be no sharp edges on the unit.
24 | Mobility, portability(if relevant) | Easy and safe transport to be possible by hand.
25 | Raw Materials(if relevant) | Blades to be made of surgical grade stainless steel material or any equivalent or better quality material.
Utility requirements
26 | Electrical, water and/or gas supply (if relevant) | Powered by internal, rechargeable, replaceable battery. Battery cover to be secure but simple to open and battery compartment to be sealed against liquid ingress. Battery charger to operate from input supply 110–220 V, 60–50 Hz, ± 10% (battery charger built-in, preferably, or external) with protection against over-voltage and over-current line conditions.
Accessories, consumables, spare parts, other components
27 | Accessories (if relevant) | Protective case for clean storage and safe transport (with dedicated space to transport at least: the handle, set of batteries and 3 blades). At least the following 4 autoclavable blades-sets for adult and pediatric applications: Macintosh, size 2 (paediatric); Macintosh, size 4 (adult); Macintosh size 3 (preferable); Miller, size 2 (paediatric); Miller, size 3 (adult); Miller size 1 (preferable).
28 | Sterilization process for accessories (if relevant) | Supplier to describe any sterilisation process required for accessories, if any.
29 | Consumables / reagents (if relevant) | Tongue depressors, single use. Supplier to describe any additional necessary consumables, detailing shelf life and number of uses.
30 | Spare parts (if relevant) | Replacement battery pack, supplied empty of charge. Spare light bulb compatible with the device provided.
31 | Other components (if relevant)
Packaging
32 | Sterility status on delivery (if relevant) | N/A
33 | Shelf life (if relevant) | N/A
34 | Transportation and storage (if relevant) | Unit shall be supplied protectively packed for safe transportation and delivery.
35 | Labelling (if relevant) | N/A
Environmental requirements
36 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0–50 °C and relative humidity of 15–85%, preferably 90%. Capable of operating continuously in ambient temperature of 10–40 °C and relative humidity of 15–85%, preferably 90%.
Training, installation and utilization
37 | Pre-installation requirements(if relevant) | Not required.
38 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation. Supplier to perform installation, safety and operation checks before handover.
39 | Training of user/s (if relevant) | Training of users in operation and technicians in basic maintenance shall be provided.
40 | User care(if relevant) | The whole unit is to be cleanable with standard disinfectants usually used in hospitals.
Warranty and maintenance
41 | Warranty | Not less than 2 years with specific inclusions and exclusions to be clearly listed. Contact details of manufacturer, supplier and local service agent to be provided.
42 | Maintenance tasks | List shall be provided of equipment and procedures required for local calibration and routine maintenance.
43 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
44 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
45 | Software / Hardware upgrade availability | N/A
Documentation
46 | Documentation requirements | User/technical manual to be supplied in English (provision of versions in other UN languages, if available, will be an asset). Certificate of calibration and inspection to be provided, when applicable. List to be provided of equipment and procedures required for local calibration, if necessary, and routine maintenance. Battery disposal according local laws.
Decommissioning
47 | Estimated Life Span | Not less than 5 years.
Safety and standards
48 | Standards, for the manufacturer and the equipment | Certified quality management system for medical devices (e.g. ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory purposes). General quality management (e.g. ISO 9001:2015 Quality management systems – Requirements). Application of risk management to medical devices (e.g. ISO 14971:2019 Medical devices – Application of risk management to medical devices).
49 | Regulatory Approval / Certification | Free sales certificate (FSC). Certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
50 | International standards | Laryngoscope Compliance to the following international standards, when applicable, or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party). Reference to the last available version is recommended but compliance to previous standards versions could be considered. IEC 60601-1:2012 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-2:2007 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility - Requirements and tests. ISO 7376-1:1994 Laryngoscopic fittings – Part 1: Conventional hook-on type handle-blade fittings IEC 60601-2-18 Ed. 3.0:2009 (b) Medical electrical equipment – Part 2-18: Particular requirements for the basic safety and essential performance of endoscopic equipment. ISO 8600-1:2013 Endoscopes – Medical endoscopes and endotherapy devices – Part 1: General requirements. ISO 8600-4:1997 Optics and optical instruments – Medical endoscopes and certain accessories – Part 4: Determination of maximum width of insertion portion. ISO 8600-7:2012 Endoscopes – Medical endoscopes and endotherapy devices – Part 7: Basic requirements for medical endoscopes of water-resistant type.