## Foglio1
Ultrasound imaging system specific for TRUS
i | Version No. | 1
ii | Date of initial version | 10 October 2019
iii | Date of last modification
iv | Date of publication
v | Completed / submitted by | Working group (WHO/IAEA)
Name, category and coding
1 | WHO Category / Code | (under development)
2 | Generic name | Ultrasound imaging system
3 | Specific type or variation (optional) | Not applicable
4 | GMDN name | General purpose ultrasound imaging system
5 | GMDN code | 40761
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Scanners, Ultrasonic, General-Purpose
8 | UMDNS code | 15976
9 | UNSPS code (optional) | 42201700
10 | Alternative name/s (optional) | Not applicable
11 | Alternative code/s (optional) | Not applicable
12 | Keywords (optional) | Brachytherapy
13 | GMDN/UMDNS definition (optional) | A stationary or mobile (e.g. on wheels) assembly of devices designed to collect, display, and analyse ultrasound images during a variety of extracorporeal and/or intracorporeal (endosonography or endoscopic) ultrasound imaging procedures (e.g. cardiac, OB/GYN, endoscopy, breast, prostate, vascular, and intra-surgical imaging). It consists of a mains (AC-powered) data processing unit with integrated software and a monitor. It is typically presented as a mobile assembly which may support a wide variety of transducers and related application software packages; an ultrasound transducer(s) may be included.
14 | EMDN code | Z110401
15 | EMDN nomenclature | ULTRASOUND SCANNERS
Purpose of use
16 | Clinical or other purpose | Imaging to guide applicator placement for brachytherapy.
17 | Level of use (if relevant) | Hospital
18 | Clinical department/ ward(if relevant) | Radiation Oncology Department
19 | Overview of functional requirements | A mobile ultrasound unit for imaging of the treatment applicator during placement for brachytherapy. The unit shall include trans-abdominal transducers, detection system, display, and image capture and storage to allow visualization of various applicators.
Technical characteristics
20 | Detailed requirements | • A mobile clinical ultrasound imaging unit capable of operating in the following modes: 2D mode and M-mode (motion mode). • Two active transducer ports. • Two high-frequency convex transducers for trans-abdominal gynaecological examinations covering a frequency range of 2.0–5.0 MHz. • (Package 2) Two biplane transrectal transducers for sagittal and transverse imaging covering a frequency range of 3.0–12.0 MHz. • A single high-resolution monitor, at least 18-inch, for image display and image processing. The displayed image shall contain not less than 256 grey levels in 2D mode. • Keyboard with one button control. • A universal serial bus (USB) port, video output, onboard thermal printer and A4 laser printer. • Image processing measurements: length, height, width, angle, area and volume calculations. • Image storage in a clinically useful format and resolution on a hard drive and with capability to transfer images to external media with the Digital Imaging and Communications in Medicine (DICOM) standard. • A supply of water soluble and hypoallergenic ultrasound gel.
21 | Displayed parameters | Comprehensive parameters at control console.
22 | User adjustable settings | At control console
Physical/chemical characteristics
23 | Components(if relevant) | Major components of the ultrasound unit mobile frame, display, console, and transducers.
24 | Mobility, portability(if relevant) | Mobile
25 | Raw Materials(if relevant) | Not applicable
Utility requirements
26 | Electrical, water and/or gas supply (if relevant) | Single-phase electrical power and air-conditioning.
Accessories, consumables, spare parts, other components
27 | Accessories (if relevant) | Not applicable
28 | Sterilization process for accessories (if relevant) | Low temperature sterilization for probes.
29 | Consumables / reagents (if relevant) | Water-based gel
30 | Spare parts (if relevant) | Access to spare parts is required.
31 | Other components (if relevant) | Not applicable
Packaging
32 | Sterility status on delivery (if relevant) | Not applicable
33 | Shelf life (if relevant) | Not applicable
34 | Transportation and storage (if relevant) | If storage is required prior to installation, the storage area should be secure, clean and dust-free, dry, cool, well-lit, ventilated and vermin-proof.
35 | Labelling (if relevant) | Symbols used according to ISO-15233 standards.
Environmental requirements
36 | Context-dependent requirements | It is the responsibility of the hospital to ensure that the manufacturer’s recommended operation and storage conditions, as given in their site planning guide, are respected. If the device requires that the operating and/or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure.
Training, installation and utilization
37 | Pre-installation requirements(if relevant) | Not applicable
38 | Requirements for commissioning (if relevant) | Acceptance testing, commissioning of transducers, baselines for quality control.
39 | Training of user/s (if relevant) | Applications training in configuration, clinical operation, safety features and service mode.
40 | User care(if relevant) | The equipment shall only be used in accordance with departmental policies, procedures and work instructions, and consistent with manufacturer’s use instructions, to ensure safety and protection.
Warranty and maintenance
41 | Warranty | At least 12 months
42 | Maintenance tasks | Approximately 4 service days per year. Tasks as per manufacturer’s preventative maintenance schedule.
43 | Type of service contract | Full-service contract, including parts, service and repair or in-house maintenance engineering service.
44 | Spare parts availability post-warranty | 9 years minimum
45 | Software / Hardware upgrade availability | To be specified by manufacturer.
Documentation
46 | Documentation requirements | Comprehensive user manual.
Decommissioning
47 | Estimated Life Span | 5–10 years
Safety and standards
48 | Risk classification | US FDA: Device Class 2
49 | Regulatory approval/certification
50 | International standards | IEC, Medical Electrical Equipment, Part 1: General Requirements for Safety, Rep. IEC 60601-1:2005+AMD1:2012. IEC, Medical electrical equipment - Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment, IEC 60601-2-37:2007+AMD1:2015 CSV. IEC, Ultrasonics - Pulse-echo scanners - Part 1: Techniques for calibrating spatial measurement systems and measurement of point-spread function response, IEC 61391-1:2006+AMD1:2017 CSV. IEC, Ultrasonics - Pulse-echo scanners - Part 2: Measurement of maximum depth of penetration and local dynamic range, IEC 61391-2:2010. IEC, Ultrasonics - Field characterization - Test methods for the determination of thermal and mechanical indices related to medical diagnostic ultrasonic fields, IEC 62359:2010+AMD1:2017 CSV.
51 | Regional/local standards | Country-specific and regional standards may apply.
52 | Regulations | USA: 21CFR 892.1560 Ultrasonic pulsed echo imaging system.