## Glucometer
Blood Glucose Meter, handheld
i | Version No. | Two (2)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2022-09-12
iv | Date of publication
v | Prepared by | World Health Organization
NAME, CATEGORY, AND CODING
1 | WHO Category / MTE Code
2 | Generic name | Glucose meter, handheld
3 | EMDN code / name | W0201060102 / Blood Glucose Meters
4 | Specific type or variation (optional) | Enzymatic with variable detection methods
5 | Alternative name/s (optional) | Glucometer; POCT glucose; Blood Glucose Self-Monitoring (SMBG); Haemoglucotest (HGT); Blood Glucose Meters, portable; Blood Glucose Monitors, portable; blood sugar meters, hypoglycemia detection units, portable glucose analyzers.
6 | Keywords (optional) | Glucose, hypoglycemia, hyperglycemia, SMBG, self monitoring, blood glucose, POCT, self-testing, diabetes
INTENDED USE
7 | Detection target | Blood glucose.
8 | Test purpose | To support diagnosis and/or monitoring type 1 and type 2 diabetes mellitus, and/or to support diagnosis and screening for diabetes, and/or to support diagnosis of hypoglycaemia and screening for intermediate hyperglycaemia, and/or to support diagnosis of impaired fasting glucose/impaired glucose tolerance, and/or to support screening for type 2 diabetes mellitus and impaired fasting glucose/impaired glucose tolerance.
9 | Specific disorder / condition | Diabetes mellitus; hypoglycemia; hyperglycemia
10 | Test format | Enzymatic method, commonly either glucose oxidase or glucose dehydrogenase, with photometric, amperometric, elecrochemical or biosensor detection (the method offered/available shall be specified)
11 | Specimen type | At least capillary whole blood and, preferably also venous whole blood.
12 | Testing population | Adult and paediatric
13 | Intended user | Equipment designed for single-patient self-testing and/or multiple-patients POC testing. Hand held portable devices, ideally training would be provided, however shall be usable without specialist training. Must be stated in instructions for use.
14 | Level of the healthcare system | All levels of healthcare setting (health centres, district hospitals, specialized hospitals, community based settings, health posts, home-care, etc.).
PERFORMANCE CHARACTERISTICS
15 | Accuracy | Analytical performance shall remain stable throughout the measurement of the analyte. The assay shall be linear across the entire measuring interval available. External Quality Assurance (EQA) at regular intervals is advisable for devices in healthcare settings (EQA watery controls only should not be preferred).
16 | Range of Measurement and related Accuracy | Measurement range not less than from 30 to 400 mg/dL (1.7 to 22.2 mmol/L), preferably from 20 up to 500 mg/dL (1.1 to 27,8 mmol/L). Accuracy must meet ISO-15197 standard, in particular: a) 95% of blood glucose results must be within 15% for values equal to or greater than 100 mg/dL, and 15 mg/dL for values below 100 mg/dL. b) 99% of results to fall within zones A or B of the Consensus Error Grid (Parkes error grid)
17 | Analytical specificity | Analytical specificity (interferences): method shall be evaluated and demonstrate no interference from ascorbate, lipid, protein, high/low hematocrit. At least the interference list of ISO 15197:2013 shall be evaluated and demonstrated.
18 | Limit of detection | Not higher than 30 mg/dL or 1.7 mmol/L (better not higher than 20 mg/dL or 1.1 mmol/L)
19 | Invalid/error/unreturnable rate | Test failure, low battery, test strips problems (i.e. expired or inserted wrong side) and malfunction messages should also be available. Error messages linked to the specific linearity range (when available).
20 | Precision (repeatibility/reproducibility) | Repeatability (within-run variability) CV <5.0 %. Must be stated in instructions for use.
21 | Trueness of measurement: bias | Typically <15% (better 10%) from target value or less than 0.83 mmol/L for lower glucose ranges, hematocrit levels will significantly affect result (bias) (see "Test Limitations" requirement for details).
TECHNICAL AND OPERATIONAL CHARACTERISTICS
22 | Principle of the assay | Enzymatic method (or equivalent; equivalence shall be demonstrated and clearly reported based on the standards required), commonly either glucose oxidase or glucose dehydrogenase, with photometric, amperometric, elecrochemical or biosensor detection
23 | Specimen(s) stability | Specimen shall be stable for at least 15 minutes (better for at least 30 minutes) from the time to the exposure to enviroment conditions. In case of venous blood samples are used, any specific need of anti-coagulants use should be clearly stated and detailed by the manufacturer.
24 | Specimen(s) volume | Sample volume at least < 15µl (in case of self-monitoring/single patient device: better if < 10µl). Manufacturer instructions for tubes and/or specific analytical instruments shall state additional requirements, if any.
25 | Type of result | In case of POCT/multiple patients device for professionals: quantitative values shall be clearly presented in display and in mmol/L or mg/dL (or both, when available, depending on manufacturer settings). In case of self-monitoring/single patient device: it should not be possible to switch between different units
26 | Time to result | In case of POCT/multiple patients device for professionals: results available in less than 1 minute (preferably less than 30 seconds). In case of self-monitoring/single patient device: results available in less than 30 seconds (preferable less than 10 seconds).
27 | End-point stability | Not applicable
28 | Ease of use for POC tests only: number of steps that require precision | Single step, easy to use devices.
29 | Specimen throughput per operator, per hour, per 8-hour working day or per batched run | In case of POCT/multiple patients device for professionals: not less than 50 specimens per 8-hour working day, depending on healthcare setting.
30 | Test limitations | Haematocrit acceptable range at least from 30% to 55% (better from 15% up to 65%). Device able to compensate for high haematocrit value. The manufacturer shall state ALL additional limitations, if any (i.e. blood concentrations of ascorbic acid or galactose, etc.)
31 | Internal quality control | Quality control features available. Quality Control material shall be run as defined by the manufacturer. In case of self-monitoring/single patient device: automatic meter calibration check (no "coding"). In case of POCT/multiple patients device for professionals: preferably automatic meter calibration check (no "coding") and/or calibration checks frequency and related activities clearly stated by the manufacturer (at least twice per year would be recommended).
32 | Compatability with external quality control material | External Quality Assessment (EQA) material shall be run from a National EQA provider (when available).
33 | Transport stability of kit/reagents (temperature and humidity) | At least in the range from 5°C to 35°C (better if up to at least 40°C), protected against high humidity and direct sunlight. Manufacturers instructions shall state specific detail and limitations, if any.
34 | Storage stability of kit/reagents (temperature and humidity) | At least in the range from 5°C to 35°C (better if up to at least 40°C), protected against high humidity and direct sunlight. Manufacturers instructions shall state specific detail and limitations, if any.
35 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
36 | In-use stability of kit/reagents (temperature and humidity) | Not applicable
37 | Transport stability for specimen collection media (temperature and humidity) | Not applicable
38 | Storage stability for specimen collection media (temperature and humidity) | Not applicable
39 | In-use stability for specimen collection media (temperature and humidity) | Not applicable
40 | Transport stability for controls/calibrators (temperature and humidity) | Temperature and humidity stability ranges shall be clearly stated by the manufacturer, when necessary
41 | Storage stability for controls/calibrators (temperature and humidity) | Temperature and humidity stability ranges shall be clearly stated by the manufacturer, when necessary
42 | On-board stability for controls/calibrators (temperature and humidity) | Not applicable
43 | In-use stability for controls/calibrators (temperature and humidity) | Based on manufacturer instructions
44 | Shelf-life of kit/reagents upon manufacture (months) | All strips should have at least 12 months expiry date (better 18 months) from the date of production (better: from the date of supply). Shelf life for reagents shall be anyway clearly stated by the manufacturer.
45 | Remaining shelf-life of kit/reagents upon delivery (months) | Strips should have a minimum of 2 (better 3 or more) months shelf life after opening the strip vial.
INSTRUMENT PHYSICAL AND TECHNICAL CHARACTERISTICS
46 | Size of device (Height x Width x Depth) | Handheld device (approx. average size as reference: 6 cm x 12 cm x 3 cm)
47 | Weight of the device (kg) | Not greater than 0.5 kg (preferably including batteries)
48 | Power requirements and characteristics | Operated by internal battery with not any memory loss if batteries are removed. Batteries may be single use, or rechargeable (preferred) with external AC battery charger, or by USB connection. Power requirement 100-240 V (± 10%)/ 50-60Hz (country dependant) for rechargeable devices. Charger, if used, must have protection against over-voltage and over-current line conditions. If rechargeable device is provided, operation should be preferably possible while charging. Battery should be replaceable without using any tools. Preferable: due to trasnport restrictions, the compatible batteries provided should be without Lithium
49 | Time to battery charge | Not higher than 3 hours
50 | Battery duration | In case of POCT/multiple patients device for professionals: not less than up to 30 days at an average of 20 tests per day and/or 600 tests in total (better: avereage of 30 tests per day and/or 900 tests in total). In case of self-monitoring/single patient device: not less than 100 tests.
51 | Alternative charging options | Solar
52 | Operating and storage conditions (temperature and humidity) | Glucometer: operating temperature range at least: 10°C - 40°C (better at least: 5°C - 45°C), humidity: 10-80% (better up to 90%) without condensation, and atmospheric pressure operating range: at least from 800 up to 1060hPa; additional altitude related operating limitations, up to at least 3000m, should be clearly stated by the manufacturer. Storage conditions range: at least -5°C - 50°C (better at least: -10°C - 55°C) and humidity 20 % – 80% RH. Test Strips packages: for high humidity (> 80%) or temperature conditions (> 40°C), the manufacturer shall guarantee that the container for strips, when provided, can be opened frequently without influencing the performance of strips/glucosemeter.
53 | User interface | LCD display, simple test menu. In case of POCT/multiple patients device for professionals: option for connection to laboratory computer and middleware should be also available. In case of self-monitoring/single patient device: technical solutions to improve the readability for private users with limited visual capacity should be preferably available.
54 | Displayed parameters | Digital display of test results. Electronic output. Audible and/or visual alert/alarm indicator at least for "low battery", "malfunction" and/or "strips error". Preferable: "expired strip" messages/alarm when a strip used is expired and additional alarm/warning for elevated or too low glucose levels
55 | Display languages | At least 20 languages should be available including Arabic, Chinese, English, French, Russian and Spanish. Local language of the Country of sales would be also preferred, when available.
56 | Built-in memory storage capacity | In case of POCT/multiple patients device for professionals: storage for at least 300 test results and 10 control results (better up to at least 500 test results and 20 control results). In case of self-monitoring/single patient device: storage for at least 100 test results (better up to at least 250 test-results for long term-storage) and, when available, 5 control results.
57 | Diagnostic connectivity | Capabilities for data transmission and storage e.g. via USB, bluetooth, cable. In case of self-monitoring/single patient device: electronic linkage to manufacturer for data updates and/or instrument monitoring should be also preferably available. In case of POCT/multiple patients device for professionals: Availability of a digital connection to the Electronic Patient Record AND/OR data connectivity compatible with laboratory information system (LIS) and programme database to enable information sharing. Electronic linkage to manufacturer for data updates and instrument monitoring should be also preferably available.
58 | Open or Closed system | Preferably open (able to use compatible consumables of different brands)
59 | Multidisease testing capabilities / Test menu availability | Dedicated blood glucose analyser
INFRASTRUCTURE REQUIREMENTS
60 | Water requirements | Not applicable
61 | Refrigeration or cold chain for storage of kit/reagents | Consumables shall be stored according to the Manufacturer’s instructions.
62 | Refrigeration or cold chain for reconstituted reagents and controls | Consumables shall be stored according to the Manufacturer’s instructions.
ACCESSORIES, CONSUMABLES, SPARE PARTS AND OTHER COMPONENTS
63 | Kit component – test / reagents / consumables (if relevant) | All reagents, accessories and consumables to have a "ready to start" system shall be provided by manufacturer. The components provided shall be clearly listed. At least the following components (consumables and accessories) shall be provided (at least strips but also the other consumables/accessories should be available in the local market.): • Test Strips, cartridges or cuvettes, depending on the manufaturer instructions for use • Single-use lancets to be used without a lancing device, with at least the following characteristics: pre loaded pressure operated (preferable); easy handling, two step operation and button operated design (preferable); single use, sterile and unable to be reused; needle is fully shielded before and after use; automatically retracts the needle into casing after sampling; edged, high quality ultra sharp needle, which can rapidly penetrate skin to reduce patient’s pain; protect the user from risk of blood borne infection. • If needed or when specifically required: specific lancing device and the necessary sets of compatible disposable lancets. Depending on the specific needs, safety lancets would be recommended. • Hand Case
In addition, in case of POCT/multiple patients device for professionals: • Set of control solutions (at least: low, medium and high), when necessary for calibration. In alternative and based on manufacturer instructions, the calibration materials shall be provided. • Data download cable (when necessary) In addition, in case of self-monitoring/single patient device: • When necessary: data download cable • When necessary and stated by manufacturer: quality control materials
64 | Reagent Kit size (number of tests) | Kit/package size of test strips (or other compatible consumables such as cuvettes, etc.) will depend on the health-care setting and the workload. In case of self-monitoring/single patient device: single packages containing the complete material for not more than 50 tests each should be available (better if not more than 25 per each package).
65 | Consumables required but not provided in the test kit | Specimen collection receptacles. List of reagents, consumables and accessories, including available compatible commercial products if any, shall be provided by the bidder, the distributor and/or the manufacturer.
66 | Other auxiliary laboratory equipment required but not provided | In case of POCT/multiple patients device for professionals: when needed, pipettes (autopipettes and glass), gloves, eye defenders, laboratory coat and other materials could be requested. In case of self-monitoring/single patient device: not applicable
67 | Spare parts (if relevant) | Set of spare compatible batteries (preferably available in the local market). List of spare parts, including available compatible commercial products if any, shall be provided by the bidder, the distributor and/or the manufacturer.
DOCUMENTATION
68 | Instrument operator manual | Operating and service manuals available in at least 20 languages (including Arabic, Chinese, English, French, Russian and Spanish) must include lists of important spares and accessories - with their part numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community. Preferably: manufacturers should provide the translation of the operating manual in the local lanugage of the country of sales
69 | Instructions for use | Required: IFU (Instructions For Use) submitted must relate to regulatory version registered for sale and use in country of supply. Manufacturers should provide the translation of the IFU in the local lanugage of the country of sales. In case of self-monitoring/single patient device: solutions to improve the readability for private users with limited visual capacity should be preferably available.
70 | Certificate of analysis | Required. Apart from the specific product test pass report, a Certificate of Analysis must be provided by the manufacturer and must comply with the local and country requirements for the product. They shall contain the actual results obtained from final quality control for lot release performed for each lot of a product.
71 | Material safety data sheet (MSDS) | Required. The manufacturer shall supply detailed documentation for each method providing detailed and comprehensive information of a product related to (i) the health effects of exposure to the product; (ii) hazard evaluation related to the handling of the product, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
ENVIRONMENTAL AND SAFETY REQUIREMENTS
72 | Hazardous classification | The manufacturer shall clearly document potential source of harm in kit inserts. Classification shall be according to the Global Harmonized System (GHS) e.g. health hazard, physical hazard or environmental hazard. More specific details will usually be contained in the Material Safety Data Sheet (MSDS)
73 | Disposal requirements | Description of disposal techniques available and/or required, according to the hazard classification e.g. simple waste disposal, hazardous waste, incineration etc.
TRAINING, INSTALLATION AND UTILISATION
74 | Pre-installation requirements (if relevant) | Not applicable
75 | Requirements for installation and calibration | Power requirement 100-240 V (± 10%)/ 50-60Hz (country dependant) for rechargeable devices. Some devices will use single use batteries. The device is fully automatic and it shall be provided already calibrated. Specific calibration recurrent activities (if necessary) shall be clearly stated by the manufacturer in the documentation provided. Automatic shutoff function.
76 | Training of user/s (if relevant) | Manufacturers shall provide training details related to use and maintenance of the device. Preferable: video tutorial for equipment use and maintenance shall be available on line and recorded in a CD, DVD or USB delivered with the equipment (in case of self-monitoring/single patient device: special focus should be preferably put on illiterate private users and on users with limited visual capacity). In addition, in case of POCT/multiple patients device for professionals: User training topics shall include at least: 1) Normal operating mode, 2) Daily and periodic cleaning and maintenance, 3) Basic troubleshooting, 4) Risks in the use of the equipment. Maintenance training topics shall include at least: 1) Normal operating mode, 2) Daily and periodic cleaning and maintenance, 3) Advanced troubleshooting, 4) Accessories and spare parts description and codes, 5) Warranty conditions, 6) Supplier and Manufacturer contact references.
WARRANTY AND MAINTENANCE
77 | Warranty | Manufacturer must provide a minimum of two (2) years warranty from the date of puchase. The warranty shall cover all delivered equipment and accessories. The date of commencement, duration of warranty period, specific exclusions/inclusions and other conditions such as maintenance support during warranty must be specified. Distrbution and after-sale services networks to be specified by the manufacturer or by the bidder at the time of the purchase. In case that the "Supplier" propose a manufacturer warranty extension, at least the following requirements shall be included: a. The warranty covers all equipment and accessories delivered. b. Specified duration including corrective maintenance. c. It includes corrective maintenance: labour and spare parts with no call limitations with the sole exclusion of damage from proven misuse. d. Corrective maintenance during warranty can be carried out at the beneficiary premise, or by mailing the broken equipment at supplier or manufacturer premise and mailing back the repaired equipment or the subtituting one to the Beneficiary premise. In the case of transportation all the costs and risks will be covered by the Supplier. e. The time elapsed between the communication of the faulty equipment and the response of the Supplier will be, within the warranty period, no more than 2 business days. f. The corrective maintenance technical service must be performed by a company based in the country of destination. g. Supplier will notify the end user in case of any equipment Recall / Field safety Notice / Alerts as issued by the manufacturer/NRA or any other relevant authorities and organizations.
78 | Preventive maintenance | The device has the capability to be cleaned and disinfected easily and to hold up under frequent disinfection.The minimal maintenance expected to be conducted by user and supplier shall be stated. Frequency of maintenance shall be based either on a fixed time period or based on the number of tests the instrument processes; this will de dictated by workload. Generally will include external cleaning of the device and change of batteries.
79 | Corrective maintenance | Not applicable
80 | Type of service contract (including leasing or reagent rental) | Possibility of instrument loan should be preferably available. In case of POCT/multiple patients device for professionals: in addition to the warranty obligations, any service contract must cover labour, repair, spare parts, shipping and logistics costs and training.
81 | Spare parts availability post-warranty | Availability of spare parts (but also accessories and consumables) shall be at least 5 years
82 | In-country technical support availability | In addition to the warranty obligations, the manufacturer or agreed distributor should have in-country technical support available to provide adequate maintenance/repair support, when needed.
DECOMMISIONING
83 | Estimated lifespan | The analyser shall have a minimum life span of 5 years; true life span will depend on throughput of the analyser. (see 'How to Plan and Budget for your Healthcare Technology' http://www.who.int/management/plan_budget_healthcare.pdf)
QUALITY AND REGISTRATION
84 | Global regulatory approvals | For product to be considered, it shall have marketing authorization for supply within the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (see https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y).
85 | Regulatory approval / certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, for example but not limited to the following product’s risk classifications: EUROPE/CE 2017/745 conformity or equipment with CE mark according with 93/42 EEC; USA/FDA certification (Food and Drug Administration) for internal US market; CANADA/SOR/98-282 conformity certification; AUSTRALIA/TGA Conformity certification; JAPAN/ PMDA pre-market approval; etc.
86 | Standards for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). Application of risk management to medical devices (e.g. ISO 14971).
87 | WHO prequalification status | Prequalification is recommended If the test category is within the scope of WHO PQ. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr If the test category is not within the scope of WHO PQ then not applicable.
88 | Post-market surveillance | A summary of current post-market information by an applicant should be provided, when available, as an indicator of product performance. Contractual agreements shall require supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system. Whether the supplier is a manufacturer, an authorised representative, or a distributor, each has an important role in ensuring timely feedback of complaints, the identification of non-conformity, and subsequent recall, corrections or corrective actions.
89 | Field Safety Collective Actions | It is recommended that manufacturers/providers have communication channels in place to notify any corrective actions to final users, such communication channels shall be declared in the contract.
90 | Standards for product performance | Compliance to the following international standards or to regional or national equivalent, (including the technical tests for safety and performance from accredited laboratory or third party). Latest version recommended but compliance to previous standards versions could be accepted: • IEC 60601-1-2 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests • IEC 60601-1-6 Medical electrical equipment - Part 1-6: General requirements for basic safety and essential performance - Collateral standard: Usability • ISO 15197 In vitro diagnostic test systems — Requirements for blood-glucose monitoring systems for self-testing in managing diabetes mellitus • ISO 22870:2016: Point-of-care testing (POCT) — Requirements for quality and competence.

## Acknow&Process
Technical Specifications Development and Revision process
Period (N. of weeks - Quarter / year) | Process phase: Description | Stakeholders involved
10 weeks - 2nd quarter/2022 | TS definition: preparation, discussion and agreement of the first draft done with staff and consultants of MD Dept of WHO HQ | MD department - WHO HQ
6 weeks - 2nd and 3rd quarter/2022 | Revision Phase 1: Revision of the first draft done by other WHO stakeholders (staff and consultants) | WHO HQ - WHO EURO - PAHO - WHO SEARO - WHO AFRO - WHO WPRO
6 weeks - 3rd quarter/2022 | Revision Phase 2: Revision of the first draft done by UN agencies and WHO Collaborating Centres | UNOPS - UNDP - UNICEF - Fondazione IRCCS Policlinico San Matteo di Pavia - Vermont University - Saudi FDA - CECMED Cuba
6 weeks - 3rd quarter/2022 | Revision Phase 3: Revision of the first draft done by Non-State-Actors and other Non-WHO stakeholders | ICRCCC; FIND; UNIGE; GMTA; IMFBE; HUMATEM; DITTA; INTERNATIONAL DIABETES FEDERATION; LIFE FOR A CHILD; IAPO; IFBLS.
2 weeks - 3rd quarter/2022 | Final Draft: based on the results, feedbacks and comments coming from the three previous phases of revision, preparation of the TS Final Draft | MD department - WHO HQ
2 weeks - 3rd quarter/2022 | Final Revision: Revision of the Final Draft with ALL the stakeholders involved in the three previous phases | ALL stakeholders involved in the previous phases
1 week - 4th quarter/2022 | TS Final document: based on the results, feedbacks and comments coming from the Final Revision phase, finalization of the TS Final document | MD department - WHO HQ
Acknowledgments
Name |  | Organization
Velazquez Berumen Adriana |  | WHO, Staff
Moussy Francis Gabriel |  | WHO, Staff
Aceves Capri Ana |  | WHO, Staff
Hashimoto Naofumi |  | WHO, Staff
Humbert Tifenn Lucile Marie |  | WHO, Staff
Coulibaly Sheick Oumar |  | WHO, Staff
Ying Lin Ling |  | WHO, Staff
Roglic Gojka |  | WHO, Staff
Stroher Ute |  | WHO, Staff
Lemgruber Alexandre |  | WHO, Staff
Tusabe Martha |  | WHO, Staff
Wehbi Mohamad |  | WHO, Staff
Himley Stephen |  | WHO, Staff
Iyengar Swathi Bhargav |  | WHO, Staff
Colbert Allison |  | WHO, Staff
Taguembou Diana |  | WHO, Staff
Thangavelu Sasikala |  | WHO, Staff
Shae, Kylie |  | WHO, Staff
Shin Jinho |  | WHO, Staff
Tebutt Emma |  | WHO, Staff
Kijo Agnes |  | WHO, Staff
Sougou Aissatou Sarassa |  | WHO, Staff
Mulerwa Martha |  | WHO, Staff
Annan Edith Andrews |  | WHO, Staff
Tay Siang Kiu |  | WHO, Staff
Francesco Ribolzi |  | WHO, Consultant
Claudio Meirovich |  | WHO, Consultant
Albert Einstien |  | WHO, Consultant
Rosales Alfonso |  | WHO, Consultant
Enriquez Rodriguez Nilda |  | WHO, Consultant
Legaspi Cheryl Valerie |  | WHO, Consultant
Vitale Umberto |  | UNOPS
Wandani Sebonego |  | UNICEF
Jimenez Gabriela |  | UNDP
Lago Paolo |  | S.Matteo Hospital (Pavia, Italy)
Clark Tobey |  | VERMONT UNIVERSITY
Asally Razan |  | SAUDI FDA
Martinez Dulce |  | CECMED
Aebischer Perone Sigiriya |  | ICRC
Vetter Beatrice |  | FIND
Safary Elvis |  | FIND
Martinez Guillermo |  | FIND
Beran David |  | UNIGE
Miller Ashley |  | GMTA
Sachse Tatjana |  | GMTA
Judd Tom |  | IFMBE
Yadin David |  | IFMBE
Blanc-Gonnet Cathy |  | HUMATEM
Seebohm Anabel |  | DITTA
Graham Ogle |  | LIFE FOR A CHILD
Bilińska Jolanta |  | IAPO
Brady Sarah |  | IFBLS
Culliton Marie |  | IFBLS