## Foglio1
Ventilator medical, child/newborn, transport, with accessories (The specification detailed below is for general reference only and it refers to a "Ventilator for Transport, for adult and paediatric with accessories")
Published on: 19 November 2020 - Version: 3
Version 2: 18/06/2014 (modification of the initial version)
Version 1: 13/06/2012 (initial version)
Technical Specification | Minimum requirements
General technical requirements | Medical air compressor integral to unit, with inlet filter, or high performance turbines, or Venturi systems. External low-flow oxygen (preferable). If oxygen high-pressure input port (> 35 psi). Oxygen-air mixture accuracy of not higher than 10%. Oxygen conserve feature (preferable). Internal function testing/leak testing. Event log for errors traceability (preferable). All parts withstand high disinfection procedures. At least IP21 degree of protection to the harmful ingress of water. Polyvinyl chloride (PVC) materials must be avoided in the patient gas pathway.
Ventilation modes | Pressure control ventilation (PCV). Volume control ventilation (VCV). Synchronized intermittent mandatory ventilation (SIMV) (preferable). Pressure support ventilation (PSV) (preferable). Pressure regulated volume control (PRVC) (or similar preferable). Non-invasive ventilation (CPAP/ BiPAP).
Monitored and controlled parameters (by user) | Air and externally supplied oxygen mixture ratios fully controllable. FiO2: at least from 40% (or less) to at least 97% (or more). Tidal volume: 20–1000 mL (preferable). Inspiratory pressure: 0–40 cmH2O. I:E ratio. RR: 10–60 breaths/min, minimum.
Displayed parameters (colour and graphics preferable) | Display easily readable in low ambient light and sunlight. Real-time scalar waveforms for flow, volume and pressure, at least two simultaneously (preferable). Status indicators for ventilator mode, battery status, patient data, alarm settings. Airway pressures (peak and PEEP). Tidal volume (expired). Minute volume (expired). I:E ratio. Inspiration time. RR. FiO2. Occlusion pressure detection or combination of apnoea and high-pressure alarms. Air and/or oxygen pressure. Spontaneous ventilation. Leak. Gas availability/time remaining (preferable).
Alarms, related to gas delivered (visual and audible) | High/low FiO2 (preferable). High/low minute volume. High/low inspiratory pressure. Breathing circuit disconnect. Apnoea.
Alarms, related to equipment operation (visual and/or audible) | Gas supply failure. Power failure. Low battery.
Consumables, labelled “single use” (included and mentioned in a disaggregated list) | Breathing circuits for adult: single or double-limb (prefered) with standard outlet/inlet connectors with 22 mm outside diameter (at least 60). Breathing circuits for paediatric (if needed): single or double-limb (preferred) with standard outlet/inlet connectors with 22 mm outside diameter (at least 40). Exhaled gas filter (when applicable). Heat moisture exchanger filters (HMEF) (at least 90). Non-invasive ventilation masks, for adult (at least 50 of different sizes). Non-invasive ventilation masks, for paediatric (if needed) (at least 12). Viral/bacterial filters 99.99% efficiency minimum; inspiratory and expiratory, as applicable (at least 180).
Accessories, reusable (included and mentioned in a disaggregated list) | Breathing circuits for adult: double-limb with standard outlet/inlet connectors with 22 mm outside diameter (at least 2). Breathing circuits for paediatric (if needed): double-limb with standard outlet/inlet connectors with 22 mm outside diameter (at least 2). Exhalation valve if applicable. CO2 sensors (preferable). Standard hoses and connectors (i.e. DISS/NIST as applicable for the country) for oxygen wall outlets and cylinder, as required to operate.
Spare parts (included and mentioned in a disaggregated list) | 1 year’s spare parts kit as per preventive maintenance programme (preferable).
Portability | Portable equipment with mechanical strength to lever rough handling.
Power supply (voltage, frequency and plug vary across the countries) | Operates from AC power electric line: 100–240 V AC ± 10% / 50–60 Hz ± 10%. In-built rechargeable battery. Automatic switch from AC power electric-line mode to battery operating mode and vice versa. Continuous in battery operating mode with standard ventilation not less than 4 hours. Total re-charging time not greater than 6 hours.
Documentation (included) | Instruction for use; service manual and product information to be provided in English language, at least.
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for (latest version recommended but compliance to previous standards versions could be accepted): IEC 60601-1:2005 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-2:2014 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. ISO 10651-3:1997 Lung ventilators for medical use - Part 3: Particular requirements for emergency and transport ventilators. ISO 80601-2-12:2011 - Particular requirements for basic safety and essential performance of critical care ventilators. IEC 60601-1-2:2014 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. ISO 18562-1:2017 Biocompatibility evaluation of breathing gas pathways in health-care applications – Part 1: Evaluation and testing within a risk management process (if applicable). ISO 20789:2018 Anaesthetic and respiratory equipment – Passive humidifiers (if applicable). N 13718-1:2015 Medical vehicles and their equipment - Air ambulances - Part 1: Requirements for medical devices used in air ambulances (preferable). RTCA DO-160G Environmental conditions and test procedures for airborne equipment (preferable). DIN EN 794-3 Lung ventilators - Part 3: Particular requirements for emergency and transport ventilators (preferable). IEC 60601-1-8:2006 Medical electrical equipment - Part 1-8: General requirements for basic safety and essential performance - Collateral standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems (preferable).
Warranty | Minimum 2 years. Availability of accessories, consumables and spare parts for at least 2 years.