## Foglio1
Tube, endotracheal, with cuff (Several sizes)
Published on: 19 November 2020 - Version: 1
WHO Item name: | Required technical specifications: | Dimensions / Sizes / Presentation: | Material: | Certificates for product safety and quality assurance, as requested or equivalent
Tube, endotracheal, No. 4, with cuff, sterile, single use | Endotracheal tube designed to be introduced into the trachea via the mouth or nose, ensuring an unobstructed upper airway to convey gases and vapours to and from the lungs during anaesthesia, resuscitation and other situations where the patient is not properly ventilated. It consists in a thin, flexible, transparent and single hollow cylinder, with an anatomical curvature Magill- type of 37.5°, black and legibly depth markings and graduation in cm, with radio-opaque continuous line mark, with cuff and pilot balloon, with a standard connector in the proximal end. The proximal end of the tube safely fits the connector from which size is selected according to the tube size. The connector is straight and double-ended, with the proximal end being an outer, standard 15 mm internal diameter, conical tip that allows the tube to be connected to the ventilation system (breathing circuit or manual resuscitator). The distal end of the tube is open and bevelled (obliquely cut), atraumatic, with Murphy’s eye. The cuff, situated at the distal end of the tube, provides an airtight seal between the tube and the tracheal wall. It seals the lungs against the liquid secretions sloshing around in the upper airway. And, it ensures that the environment below the cuff can be pressurised and ventilated with a carefully controlled gas mixture. The cuff has a low pressure in order to avoid inadequate pressure on the tracheal mucous membrane to prevent damage or even necrosis. It typically has a capacity of 10 mL. The cuff is inflated via a small-bore inflation tube welded to the outside of the tracheal tube or built into its wall. The pilot balloon indicates the cuff distension. One end is connected to the cuff through a thin inflation tube located close to the proximal end. The other end has a spring-loaded, one-way valve that maintains a pre-set pressure in the circuit, and has Luer tip connector for syringes. The endotracheal tubes are standard in all aspects: dimension, markings and connectors. Initial sterilization method: ethylene oxide gas or Gamma radiation. Individually peel-packed in paper and/or plastic. | Tube: Size: 4. Internal diameter: 4.0 mm. External diameter: 6.7 mm. Length: 210 mm, minimum. Connector: Proximal end: standard internal diameter: 15 mm. Distal end (fits the tube): length: 11 mm, minimum. Cuff: Diameter: 10.5 mm. | Tube: polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: polyvinyl chloride (PVC); medical grade; transparent or blue. Cuff: siliconized rubber, or siliconized polyvinyl chloride (PVC). Pilot balloon: siliconized latex, or siliconized polyvinyl chloride (PVC). | ISO 5361:2016: Anaesthetic and respiratory equipment – Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products – Ethylene oxide – Requirements for the development, validation and routine control of a sterilization process for medical devices.
Tube, endotracheal, No. 5, with cuff, sterile, single use | Endotracheal tube designed to be introduced into the trachea via the mouth or nose, ensuring an unobstructed upper airway to convey gases and vapours to and from the lungs during anaesthesia, resuscitation and other situations where the patient is not properly ventilated. It consists in a thin, flexible, transparent and single hollow cylinder, with an anatomical curvature Magill- type of 37.5°, black and legibly depth markings and graduation in cm, with radio-opaque continuous line mark, with cuff and pilot balloon, with a standard connector in the proximal end. The proximal end of the tube safely fits the connector from which size is selected according to the tube size. The connector is straight and double-ended, with the proximal end being an outer, standard 15 mm internal diameter, conical tip that allows the tube to be connected to the ventilation system (breathing circuit or manual resuscitator). The distal end of the tube is open and bevelled (obliquely cut), atraumatic, with Murphy’s eye. The cuff, situated at the distal end of the tube, provides an airtight seal between the tube and the tracheal wall. It seals the lungs against the liquid secretions sloshing around in the upper airway. And, it ensures that the environment below the cuff can be pressurised and ventilated with a carefully controlled gas mixture. The cuff has a low pressure in order to avoid inadequate pressure on the tracheal mucous membrane to prevent damage or even necrosis. It typically has a capacity of 10 mL. The cuff is inflated via a small-bore inflation tube welded to the outside of the tracheal tube or built into its wall. The pilot balloon indicates the cuff distension. One end is connected to the cuff through a thin inflation tube located close to the proximal end. The other end has a spring-loaded, one-way valve that maintains a pre-set pressure in the circuit, and has Luer tip connector for syringes. The endotracheal tubes are standard in all aspects: dimension, markings and connectors. Initial sterilization method: ethylene oxide gas or Gamma radiation. Individually peel-packed in paper and/or plastic. | Tube: Size: 5. Internal diameter: 5.0 mm. External diameter: 6.7 mm. Length: 250 mm, minimum. Connector: Proximal end: standard internal diameter: 15 mm. Distal end (fits the tube): length: 12 mm, minimum. Cuff: Diameter: 13 mm. | Tube: polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: polyvinyl chloride (PVC); medical grade; transparent or blue. Cuff: siliconized rubber, or siliconized polyvinyl chloride (PVC). Pilot balloon: siliconized latex, or siliconized polyvinyl chloride (PVC). | ISO 5361:2016: Anaesthetic and respiratory equipment – Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products – Ethylene oxide – Requirements for the development, validation and routine control of a sterilization process for medical devices.
Tube, endotracheal, No. 6, with cuff, sterile, single use | Endotracheal tube designed to be introduced into the trachea via the mouth or nose, ensuring an unobstructed upper airway to convey gases and vapours to and from the lungs during anaesthesia, resuscitation and other situations where the patient is not properly ventilated. It consists in a thin, flexible, transparent and single hollow cylinder, with an anatomical curvature Magill- type of 37.5°, black and legibly depth markings and graduation in cm, with radio-opaque continuous line mark, with cuff and pilot balloon, with a standard connector in the proximal end. The proximal end of the tube safely fits the connector from which size is selected according to the tube size. The connector is straight and double-ended, with the proximal end being an outer, standard 15 mm internal diameter, conical tip that allows the tube to be connected to the ventilation system (breathing circuit or manual resuscitator). The distal end of the tube is open and bevelled (obliquely cut), atraumatic, with Murphy’s eye. The cuff, situated at the distal end of the tube, provides an airtight seal between the tube and the tracheal wall. It seals the lungs against the liquid secretions sloshing around in the upper airway. And, it ensures that the environment below the cuff can be pressurised and ventilated with a carefully controlled gas mixture. The cuff has a low pressure in order to avoid inadequate pressure on the tracheal mucous membrane to prevent damage or even necrosis. It typically has a capacity of 10 mL. The cuff is inflated via a small-bore inflation tube welded to the outside of the tracheal tube or built into its wall. The pilot balloon indicates the cuff distension. One end is connected to the cuff through a thin inflation tube located close to the proximal end. The other end has a spring-loaded, one-way valve that maintains a pre-set pressure in the circuit, and has Luer tip connector for syringes. The endotracheal tubes are standard in all aspects: dimension, markings and connectors. Initial sterilization method: ethylene oxide gas or Gamma radiation. Individually peel-packed in paper and/or plastic. | Tube: Size: 6. Internal diameter: 6.0 mm. External diameter: 8.0 mm. Length: 290 mm, minimum. Connector: Proximal end: standard internal diameter: 15 mm. Distal end (fits the tube): length: 13 mm, minimum. Cuff: Diameter: 18.5 mm. | Tube: polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: polyvinyl chloride (PVC); medical grade; transparent or blue. Cuff: siliconized rubber, or siliconized polyvinyl chloride (PVC). Pilot balloon: siliconized latex, or siliconized polyvinyl chloride (PVC). | ISO 5361:2016: Anaesthetic and respiratory equipment – Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products – Ethylene oxide – Requirements for the development, validation and routine control of a sterilization process for medical devices.
Tube, endotracheal, No. 7, with cuff, sterile, single use | Endotracheal tube designed to be introduced into the trachea via the mouth or nose, ensuring an unobstructed upper airway to convey gases and vapours to and from the lungs during anaesthesia, resuscitation and other situations where the patient is not properly ventilated. It consists in a thin, flexible, transparent and single hollow cylinder, with an anatomical curvature Magill- type of 37.5°, black and legibly depth markings and graduation in cm, with radio-opaque continuous line mark, with cuff and pilot balloon, with a standard connector in the proximal end. The proximal end of the tube safely fits the connector from which size is selected according to the tube size. The connector is straight and double-ended, with the proximal end being an outer, standard 15 mm internal diameter, conical tip that allows the tube to be connected to the ventilation system (breathing circuit or manual resuscitator). The distal end of the tube is open and bevelled (obliquely cut), atraumatic, with Murphy’s eye. The cuff, situated at the distal end of the tube, provides an airtight seal between the tube and the tracheal wall. It seals the lungs against the liquid secretions sloshing around in the upper airway. And, it ensures that the environment below the cuff can be pressurised and ventilated with a carefully controlled gas mixture. The cuff has a low pressure in order to avoid inadequate pressure on the tracheal mucous membrane to prevent damage or even necrosis. It typically has a capacity of 10 mL. The cuff is inflated via a small-bore inflation tube welded to the outside of the tracheal tube or built into its wall. The pilot balloon indicates the cuff distension. One end is connected to the cuff through a thin inflation tube located close to the proximal end. The other end has a spring-loaded, one-way valve that maintains a pre-set pressure in the circuit, and has Luer tip connector for syringes. The endotracheal tubes are standard in all aspects: dimension, markings and connectors. Initial sterilization method: ethylene oxide gas or Gamma radiation. Individually peel-packed in paper and/or plastic. | Tube: Size: 7. Internal diameter: 7.0 mm. External diameter: 9.3 mm. Length: 320 mm, minimum. Connector: Proximal end: standard internal diameter: 15 mm. Distal end (fits the tube): length: 16 mm, minimum. Cuff: Diameter: 24 mm. | Tube: polyvinyl chloride (PVC), or polyurethane; medical grade; transparent. Connector: polyvinyl chloride (PVC); medical grade; transparent or blue. Cuff: siliconized rubber, or siliconized polyvinyl chloride (PVC). Pilot balloon: siliconized latex, or siliconized polyvinyl chloride (PVC). | ISO 5361:2016: Anaesthetic and respiratory equipment – Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products – Ethylene oxide – Requirements for the development, validation and routine control of a sterilization process for medical devices.
Tube, endotracheal, No. 8, with cuff, sterile, single use | Endotracheal tube designed to be introduced into the trachea via the mouth or nose, ensuring an unobstructed upper airway to convey gases and vapours to and from the lungs during anaesthesia, resuscitation and other situations where the patient is not properly ventilated. It consists in a thin, flexible, transparent and single hollow cylinder, with an anatomical curvature Magill- type of 37.5°, black and legibly depth markings and graduation in cm, with radio-opaque continuous line mark, with cuff and pilot balloon, with a standard connector in the proximal end. The proximal end of the tube safely fits the connector from which size is selected according to the tube size. The connector is straight and double-ended, with the proximal end being an outer, standard 15 mm internal diameter, conical tip that allows the tube to be connected to the ventilation system (breathing circuit or manual resuscitator). The distal end of the tube is open and bevelled (obliquely cut), atraumatic, with Murphy’s eye. The cuff, situated at the distal end of the tube, provides an airtight seal between the tube and the tracheal wall. It seals the lungs against the liquid secretions sloshing around in the upper airway. And, it ensures that the environment below the cuff can be pressurised and ventilated with a carefully controlled gas mixture. The cuff has a low pressure in order to avoid inadequate pressure on the tracheal mucous membrane to prevent damage or even necrosis. It typically has a capacity of 10 mL. The cuff is inflated via a small-bore inflation tube welded to the outside of the tracheal tube or built into its wall. The pilot balloon indicates the cuff distension. One end is connected to the cuff through a thin inflation tube located close to the proximal end. The other end has a spring-loaded, one-way valve that maintains a pre-set pressure in the circuit, and has Luer tip connector for syringes. The endotracheal tubes are standard in all aspects: dimension, markings and connectors. Initial sterilization method: ethylene oxide gas or Gamma radiation. Individually peel-packed in paper and/or plastic. | Tube: Size: 8. Internal diameter: 8.0 mm. External diameter: 10.7 mm. Length: 340 mm, minimum. Connector: Proximal end: standard internal diameter: 15 mm. Distal end (fits the tube): length: 16 mm, minimum. Cuff: Diameter: 26 mm. | Tube: polyvinyl chloride (PVC), or polyurethane; medical grade; transparent. Connector: polyvinyl chloride (PVC); medical grade; transparent or blue. Cuff: siliconized rubber, or siliconized polyvinyl chloride (PVC). Pilot balloon: siliconized latex, or siliconized polyvinyl chloride (PVC). | ISO 5361:2016: Anaesthetic and respiratory equipment – Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process.
Tube, endotracheal, No. 9, with cuff, sterile, single use | Endotracheal tube designed to be introduced into the trachea via the mouth or nose, ensuring an unobstructed upper airway to convey gases and vapours to and from the lungs during anaesthesia, resuscitation and other situations where the patient is not properly ventilated. It consists in a thin, flexible, transparent and single hollow cylinder, with an anatomical curvature Magill- type of 37.5°, black and legibly depth markings and graduation in cm, with radio-opaque continuous line mark, with cuff and pilot balloon, with a standard connector in the proximal end. The proximal end of the tube safely fits the connector from which size is selected according to the tube size. The connector is straight and double-ended, with the proximal end being an outer, standard 15 mm internal diameter, conical tip that allows the tube to be connected to the ventilation system (breathing circuit or manual resuscitator). The distal end of the tube is open and bevelled (obliquely cut), atraumatic, with Murphy’s eye. The cuff, situated at the distal end of the tube, provides an airtight seal between the tube and the tracheal wall. It seals the lungs against the liquid secretions sloshing around in the upper airway. And, it ensures that the environment below the cuff can be pressurised and ventilated with a carefully controlled gas mixture. The cuff has a low pressure in order to avoid inadequate pressure on the tracheal mucous membrane to prevent damage or even necrosis. It typically has a capacity of 10 mL. The cuff is inflated via a small-bore inflation tube welded to the outside of the tracheal tube or built into its wall. The pilot balloon indicates the cuff distension. One end is connected to the cuff through a thin inflation tube located close to the proximal end. The other end has a spring-loaded, one-way valve that maintains a pre-set pressure in the circuit, and has Luer tip connector for syringes. The endotracheal tubes are standard in all aspects: dimension, markings and connectors. Initial sterilization method: ethylene oxide gas or Gamma radiation. Individually peel-packed in paper and/or plastic. | Tube: Size: 9. Internal diameter: 9.0 mm. External diameter: 12.0 mm. Length: 350 mm, minimum. Connector: Proximal end: standard internal diameter: 15 mm. Distal end (fits the tube): length: 16 mm, minimum. Cuff: Diameter: 28 mm. | Tube: polyvinyl chloride (PVC) or polyurethane; medical grade; transparent. Connector: polyvinyl chloride (PVC); medical grade; transparent or blue. Cuff: siliconized rubber, or siliconized polyvinyl chloride (PVC). Pilot balloon: siliconized latex, or siliconized polyvinyl chloride (PVC). | ISO 5361:2016: Anaesthetic and respiratory equipment – Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process. ISO 11135:2014: Sterilization of health-care products – Ethylene oxide – Requirements for the development, validation and routine control of a sterilization process for medical devices.