## Foglio1
Ultrasound scanner (The specification detailed is for general reference only and it refers to a "Ultrasound, portable, with accessories")
Published on: 19 November 2020 - Version: 1
Alternative naming: Cart-based point-of-care ultrasound scanners, also called portable ultrasound scanners (both touchscreen and/or standard laptop equipment and cart-based)
Technical Specification | Minimum requirements
General technical requirements | Capable of generating imaging procedures involving lungs, heart, abdomen, pelvis, blood vessels, musculoskeletal and soft tissue. Console: laptop style console design, optional touchscreen combined with conventional user- control panel. Total weight of the equipment 5–10 kg (preferable). Dimensions, approximately: 35–45 cm (L); 35–45 cm (H); 5–10 cm (D). Battery duration: minimum 2 hours under normal use conditions. Clear protective control panel cover for infection control. Imaging focusing: adjustable focal depth, either manually or automatically synchronized to the selected scanning depth. Zooming capability with automated image optimization. Depth range selection: capable of multiple depth range selection, synchronized with automatic focal zone selection. Field of view: imaging depth at least 20 cm. Image orientation: capable of lateral and vertical inversion (in B-mode). Image modes at least: • 2-D imaging • M-mode • B/M mode • dual 2-D/colour image mode with cine loop • pulsed wave Doppler, colour Doppler imaging (CDI), power Doppler imaging (PDI), duplex, continuous wave Doppler, triplex mode (optional). Needle enhancement ability. Software applications that include at least: • obstetrics/gynaecological measurements and calculations, including gestational sac mean diameter, femur length, crown-rump length, biparietal diameter and abdominal circumference, enabling estimation of gestational age and foetal weight (preferable); • small parts/soft tissues; • lung; • vascular/basic cardiac quantification; • easy selection of callipers; • measurements capabilities (distance, area and circumference by ellipse and trace method); • capability to be upgraded with additional software applications. Probe-dependant applications with factory-default pre-sets at least: cardiac, peripheral vascular, abdominal adult, abdominal paediatric, small parts, lung, MSK-general, MSK-superficial, obstetrics/gynaecological, trans-vaginal. Equipment with zoom functionality available. Screen annotations capture patient data, date and time, scanning protocols, probes. Text annotations and body markers and image orientation indicator. Transducer ports: at least two active transducer ports permanently available; capability of electronic (preferable) or manually switch between probes.
Monitor and display | Screen monitor: high-definition (HD) digital black and white and colour liquid-crystal display (LCD) monitor of at least 25 cm diagonal (across), equivalent to 10 inches (at least 15 inches preferable), with reflection filter or adequate alternative. Screen monitor protection available. Laptop monitor fold-down and lock mechanism of the screen for safe and easy transportation (if applicable). User-friendly control panel: easy to use, logical and orderly control panel: for quick and easy location of most common functions. Back lighting or highlighting of application knobs/buttons/controls.
Communication and storage | Data communication, storage and transfer interface: USB minimum, high-definition multimedia interface (HDMI) preferable. DICOM 3.0 conformity. Digital image storage capacity: at least 64 GB of cine memory for static images and dynamic clips. Cine loop: freeze and cine-loop functions. Image grey scale: 256 shades of grey and video output of 625 lines/frame 150 dB full time dynamic range. Capability for database of patient images and information.
Consumables | Ultrasound transmission gel for 1000 patients’ exams. Cleaning and disinfection products for 1000 patients’ exams. Compatible printing paper for 3 months’ operation (when printer available).
Accessories (included) | Transducers: • Phased-array 2–5 MHz for basic cardiac and lung studies. • Phased-array up to 8 MHz for paediatric patients. • Broadband curvilinear at least 2–5 MHz for general abdominal, obstetrics/gynaecological and lung ultrasound applications. This should have colour, power and spectral Doppler capabilities. M-mode is desirable for obstetrics. • Linear-array high frequency broadband at least 5–10 MHz, with colour, power and spectral Doppler capabilities for vascular and small parts. • Capability to connect endocavitary transducers. Dedicated probes covers should be provided. If applicable, matching trolley (compact and lightweight) or compatible wheeled cart, with gel bottle holders, drawer, or dedicated space for accessories, as well as a place for scanner positioning and easy orientation. Cables and other connection accessories. Storage security lock/chain and key Wheeled cart or dedicated trolley (if applicable) with gel bottle holders, drawer or dedicated space for accessories, place for scanner positioning and easy orientation. Equipment provided with necessary quality assurance and control testing tools.
Power supply (voltage, frequency and plug vary across countries) | Equipment must be connected to a reliable and continuous source of energy. Operates from AC power electric line: 100–240 V ~, 50/60 Hz. Built-in rechargeable battery included. Automatic switch from AC power electric line mode to battery operating mode and vice versa. Power supply: power supply may vary according to country. Working time in battery mode and standard operations not less than 2 hours. Battery recharging time not more than 4 hours.
Documentation (included, minimum in English language) | Hard and soft copies in English (provision of versions in other UN languages, if available, will be an asset). Certificate of calibration and inspection. User manual with specific protocols for cleaning, disinfection, troubleshooting. Service manual with calibration and routine maintenance. Contact details for after sales service. Contact details of manufacturer, supplier and local service agent.
Primary packaging label | Labelling on the primary packaging to include: • Name and/or trademark of the manufacturer. • Model or product reference. • Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer and the equipment | Certified quality management system for medical devices (e.g. ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory purposes). General quality management (e.g. ISO 9001:2015 Quality management systems – Requirements). Application of risk management to medical devices (e.g. ISO 14971:2019 Medical devices – Application of risk management to medical devices).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party). Reference to the last available version is recommended but compliance to previous standards versions could be considered (as/when applicable). • ISO 29821:2018 Condition monitoring and diagnostics of machines – Ultrasound – General guidelines, procedures and validation. • IEC 60601-1:2005+AMD1:2012 Consolidated version – Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. • IEC 60601-1-2:2014 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. • IEC 60601-2-37:2007 Medical electrical equipment – Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment. • IEC 61157:2007+AMD1:2013 Consolidated version – Standard means for the reporting of the acoustic output of medical diagnostic ultrasonic equipment. • IEC 60601-1-4 Medical electrical equipment – Part 1-4: General requirements for safety – Collateral standard: Programmable electrical medical systems.
Warranty | The system shall be covered by a warranty of at least 1 year including parts and labour, including the probes, starting as of the date of successful on-site acceptance, as per testing and acceptance. Spare parts availability for the equipment lifespan (not less than 7 years). Warranty shall include all necessary spare parts, shipment to site, cost of replacement work, personnel, disposal of faulty parts, and software (patches, upgrades, and updates).
ADDITIONAL SERVICES REQUIRED
Testing and acceptance | The system, prior to shipment, shall be tested by designated agents for conformity with manufacturer’s performance specifications and the minimum requirements specified. The results of the testing of the system shall be documented by the contractor in an acceptance protocol that shall be signed by the end user.
Training and manuals | User care instructions and protocols (manuals) to be provided, including guidance for replacement of accessories and consumables and safe decontamination of reusable parts (if applicable), indicating if they are generic or brand related. Technical maintenance protocols and manuals to be provided. Trainings for users and technical maintenance teams (provided also in online format, if available). Trainings shall include topics related to the use of quality assurance and control tools provided for testing.