## Speculum, vaginal
SPECULUM, VAGINAL
i | Version No. | 1
ii | Date of initial version | December 2019
iii | Date of last modification
iv | Date of publication | February 2020
v | Completed / submitted by | WHO
NAME, CATEGORY AND CODING
1 | WHO Category / Code | Speculum, vaginal: XD9478 Speculum, vaginal, reusable: XD1KF1
2 | Generic name | Speculum, vaginal
3 | Specific type or variation (optional) | Vaginal speculum, bivalved
4 | Keywords (optional) | surgical, gynaecology, vaginal exam
PURPOSE OF USE
5 | Clinical or other purpose | A vaginal speculum is a device intended to open the vaginal canal to enable a healthcare provider to visually inspect the cervix and collect vaginal or cervical specimens and/or perform surgical operations in a woman’s lower genital tract.
6 | Level of use (if relevant) | Hospital, clinic, or health post
7 | Clinical department/ward (if relevant) | Family medicine; gynaecology; obstetrics; outpatient clinic; outreach.
8 | Overview of functional requirements | Bivalved vaginal specula have two blades that are self-retaining to facilitate visualization of the cervix for observation, testing, or to carry out a procedure.
TECHNICAL CHARACTERISTICS
9 | Detailed requirements | • Specify: reusable or single-use • Bivalved and self-retaining to maintain an open vaginal canal • To be available in a variety of sizes (large, medium, small, AND narrow e.g. Pedersen) • Type to be specified based on user preference or availability Example types and sizes: Collins: o Large – blade length: 110mm (+/- 5%), blade width: 40mm (+/- 5%). o Medium – blade length: 100mm(+/- 5%), blade width: 35mm (+/- 5%). o Small – blade length: 85mm(+/- 5%), blade width: 30mm (+/- 5%). Cusco: o Large – blade length: 11.5cm (+/- 5%), blade width: 3.5cm (+/- 5%). o Medium – blade length: 9.5cm (+/- 5%), blade width: 3.5cm (+/- 5%). o Small – blade length: 7.5cm (+/- 5%), blade width: 2cm (+/- 5%). Graves: o Large – blade length: 115mm (+/- 5%), blade width: 35mm (+/- 5%). o Medium – blade length: 95mm(+/- 5%), blade width: 35mm (+/- 5%). o Small – blade length: 75mm(+/- 5%), blade width: 20mm (+/- 5%). Pederson (narrower version of graves): o Large – blade length: 115mm (+/- 5%), blade width: 25mm (+/- 5%). o Medium – blade length: 95mm(+/- 5%), blade width: 22mm (+/- 5%). o Small – blade length: 75mm(+/- 5%), blade width: 13mm (+/- 5%)
10 | Displayed parameters | N/A
11 | User adjustable settings | N/A
PHYSICAL/CHEMICAL CHARACTERISTICS
12 | Components (if relevant) | N/A
13 | Mobility, portability (if relevant) | N/A
14 | Raw Materials (if relevant) | Reusable: metal alloys (typically non-quenched, non-magnetic, austenitic stainless steel) and shall be autoclavable. Single-use: highstrength plastic (for example acrylics) and are supplied as sterile.
UTILITY REQUIREMENTS
15 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
16 | Accessories (if relevant) | N/A
17 | Sterilization process for accessories (if relevant) | N/A
18 | Consumables / reagents (if relevant) | N/A
19 | Spare parts (if relevant) | N/A
20 | Other components (if relevant) | N/A
PACKAGING
21 | Sterility status on delivery (if relevant) | Reusable: N/A Single-use: sterile
22 | Shelf life (if relevant) | N/A
23 | Transportation and storage (if relevant) | Storage area should be clean and dust-free, dry, cool, well-lit, ventilated and vermin-proof. Store device in original packaging on a shelf or on in a storage cabinet. Devices should be able to withstand storage temperatures ranging from 15°C to 30°C, relative humidity ≤ 85% (non-condensing), and be protected from dripping water.
24 | Labelling (if relevant) | Labelling on the primary packaging should include the name and/or trademark of the manufacturer and should adhere to the most current version of ISO 15223 – 1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements.
ENVIRONMENTAL REQUIREMENTS
25 | Context-dependent requirements | Environmental conditions vary globally and can be extreme; however, the following are tenable: - Operating temperature: 15°C to 35°C - Operating relative humidity: ≤ 85%, non-condensing It is the responsibility of the procurement body to ensure that the manufacturer’s recommended operation and storage conditions are respected. If the device requires that the operating and/ or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure.
TRAINING, INSTALLATION AND UTILISATION
26 | Pre-installation requirements(if relevant) | N/A
27 | Requirements for commissioning (if relevant) | N/A
28 | Training of user/s (if relevant) | Clinical staff training in vaginal examinations and gynaecological procedures.
29 | User care(if relevant) | N/A
WARRANTY AND MAINTENANCE
30 | Warranty | One year
31 | Maintenance tasks | N/A
32 | Type of service contract | N/A
33 | Spare parts availability post-warranty | N/A
34 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
35 | Documentation requirements | • Instructions for use and service manuals to be provided (including procedures for decontamination). • User language preference prioritized, English is mandatory. • Contact details of manufacturer, supplier and local agent.
DECOMMISSIONING
36 | Estimated Life Span | N/A
SAFETY AND STANDARDS
37 | Risk Classification | US FDA: Device Class 1 for metal speculum. EU: Class I US FDA: Device Class 2 for non-metal speculum. EU: Class I
38 | Regulatory Approval / Certification | Compliance to (where applicable, but not limited to): • National Regulatory Agency/Authority (NRA) requirements compliance • Approval by regulatory body of country of manufacturer (if applicable) And at least one of: • FDA 510k clearance (US FDA) • CE mark (EU), with indication of Notifying Body (when applicable) • Other regulatory body in an IMDRF founding member country such as Australia, Canada, or Japan.
39 | International standards | Compliant with active version of the following standards (or equivalent): For reusable products: • ISO 7153-1: Surgical instruments –Materials -- Part 1: Stainless steel • ISO 13402: Surgical and dental hand instruments -- Determination of resistance against autoclaving, corrosion and thermal exposure. For singe-use products, supplied as sterile: • ISO 17664: Processing of health care products – Information to be provided by the medical manufacturer for the processing of medical devices • ISO 11135: Sterilization of health-care products - Ethylene oxide • ISO 11137: Sterilization of health care products - Radiation • ISO 11607: Packaging for terminally sterilized medical devices
40 | Regional / Local Standards | Country-specific and regional standards may apply
41 | Regulations | US regulations: 21 CFR part 820 Quality System Regulation 21 CFR part 878.1800 - Speculum and accessories. EU regulations: European Commission Regulation (EU) No. 2017/745 (replacing 93/42/EEC) "