## Foglio1
Nasal oxygen cannula with prongs
Published on: 19 November 2020 - Version: 2
Version 1: September 2019 (WHO-UNICEF)
Technical Specification | Minimum requirements
General technical requirements | Cannula with nasal prongs designed for easy administration of medicinal oxygen through patient nostrils; single use. Low-resistance tubing, round shape section, designed for low-flow procedures, typically 0–15 L/min, where the delivered gas does not meet all the inspiratory demand and entrains ambient air. The twin prongs nasal tips are soft and smoothly finished to ensure equal oxygen flow to both nostrils. They are connected to a lip support and harness (one tube right/left side). The harness is fully adjustable (over the patient’s ear) with a double tubing (right and left side), interlinked through a moulded Y-connector to the oxygen supply line. All tubing is soft and flexible, kink resistant, with star lumen, and with proximal end with a universal, funnel-shaped connector to oxygen source. Tubing compatibility with standard oxygen connection tubing, 3–5 mm internal diameter and 7–8 mm external diameter, and 15/22 mm diameter ventilation tubing; available with “standard” and “universal” hose end connector. Individually packed in a sealed plastic envelope. Non-sterile. Box of 50 or 100 units.
Sizes | Adult: outer diameter of the prong: 6 mm; tube length: 1.5–2 m. Paediatric: outer diameter of the prong: 3.7 mm; tube length: 1.5–2 m.
Material | Rubber or soft plastic tubing and prongs, semi-rigid and allowing freedom of movement, PVC or other material, FDA Title 21/USP VI compliant and certified for medical use, hardness > 60 Shore A (ASTM D-2240).
Primary packaging label | Single use. Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: ISO 11712:2009 Anaesthetic and respiratory equipment – Supralaryngeal airways and connectors. ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen. ISO 18562 Biocompatibility evaluation of breathing gas pathways in healthcare applications. ISO 18190 Anaesthetic and respiratory equipment – General requirements for airways and related equipment. ISO 18562-1 Biocompatibility evaluation of breathing gas pathways in healthcare applications Part 1: Evaluation and testing within a risk management process. ISO/DIS 23368 Anaesthetic and respiratory equipment – Low flow nasal cannula for oxygen therapy. ISO/DIS 17256 Anaesthetic and respiratory equipment – Respiratory therapy tubing and connectors. ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements.