## BED wards
Bed, hospital (the specification detailed below refers to a hospital bed, mechanical/hydraulic )
i | Version No. | 1
ii | Date of initial version | 2021-10-06
iii | Date of last modification
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Hospital bed
3 | Specific type or variation (optional) | mechanical, hydraulic
4 | UMDNS name | Beds; Beds, Hydraulic; Beds, Mechanical
5 | UMDNS code | 10342; 10353; 10357
6 | UNSPS code (optional)
7 | Alternative name/s (optional) | Patient bed, Ward bed, Manual bed, Mechanical bed, Hydraulic bed
8 | Alternative code/s (optional)
9 | Keywords (optional) | mechanical, hydraulic, ward, manual
10 | EMDN code | V0806; V080602
11 | EMDN nomenclature | MEDICAL BEDS; MANUAL MEDICAL BEDS
PURPOSE OF USE
12 | Clinical or other purpose | Hospital bed for general and specialized wards, used during patient hospitalization.
13 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital.
14 | Clinical department/ward(if relevant) | Wards
15 | Overview of functional requirements | Hospital bed for general and specialized wards, mechanically and/or hydraulically controlled from foot end side, fixed height (option for variable height), 3 sections.
TECHNICAL CHARACTERISTICS
16 | Detailed requirements | Bed frame with at least the following characteristics: • Made in robust combination of tubular steel tube preferably epoxy coated • The top of the bed shall be composed by at least 3 sections • Equipped with plastic bumpers and receptacles for patient-lifting and IV poles • Two support handles shall be positioned on either side of the rising backrest • Made of steel tube with at least 1.5mm of thickness • Head and foot end easily removable (for resuscitation) • Head and foot end made of Stainless Steel with plastic-covered bends and grips or Laminated headboard Mattress, mattress platform and pillow with at least the following characteristics: • The platform base consisting of round/oval steel tubes crossed of rectangular cross-pieces, joints, and hinges made of cold-drawn steel. • Galvanized Grid platform with mattress retainers. • Mattress manufactured by laminating a centerpiece of chip foam, at least 120 mm thick, and of density approximately 80kg/m3. • Mattress enclosed by a waterproof cover, hinged where two pieces are joined. • The mattress size should appropriate as per bed size. • Pillow made of anti-bacteria, anti-static, toxic-free, high-density foam core, provided with removable and washable hard cotton cover (preferable)
17 | Displayed parameters | N.A.
18 | User adjustable settings | Controls and movements with at least the following characteristics: • Back-rest and Knee-brake controlled/adjusted by screw-gear based retractable crank handles preferably positioned at the foot end. • Calf section with drop-prevention device • Backrest Elevation: at least 50º. • Knee-brake Elevation: at least 20º. • Vertical movement control/adjustment via mechanically/hydraulically operated actuator with preferably dual-sided pedal control (OPTIONAL).
PHYSICAL/CHEMICAL CHARACTERISTICS
19 | Components(if relevant) | Dimensions of the bed plane, at least: 220 cm (lenght) x 90 cm (wide), approx. Workload capacity up to not less than 220Kg.
20 | Mobility, portability(if relevant) | The bed shall be mounted on suitable castors, two of which shall be free running and two braked castors (diameter at least 120 mm)
21 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
22 | Electrical, water and/or gas supply (if relevant) | N/A
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
23 | Accessories (if relevant) | At least the following accessories should be provided: • Telescopic IV-pole with 4 hooks. • Chart Holder. • Urine/Liquids sack holder. • Bed size pair of stainless steel folding side-rails (OPTIONAL) • Blanket Support (OPTIONAL)
24 | Sterilization process for accessories (if relevant) | N/A
25 | Consumables / reagents (if relevant) | N/A
26 | Spare parts (if relevant) | N/A
27 | Other components (if relevant) | N/A
PACKAGING
28 | Sterility status on delivery (if relevant) | N/A
29 | Shelf life (if relevant) | N/A
30 | Transportation and storage (if relevant) | N/A
31 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
32 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%.Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%.
TRAINING, INSTALLATION AND UTILISATION
33 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover. Local clinical staff to affirm completion of installation
34 | Requirements for commissioning (if relevant) | N/A
35 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
36 | User care(if relevant) | Table layout to enable easy cleaning and sterilization of all surfaces, with no fluid traps.
WARRANTY AND MAINTENANCE
37 | Warranty | At least 24 months
38 | Maintenance tasks | Replacement oil or grease sufficient for two years’ maintenance.
39 | Type of service contract | N/A
40 | Spare parts availability post-warranty | At least 10 years
41 | Software / Hardware upgrade availability | N/A
DOCUMENTATION
42 | Documentation requirements | All supporting documentation, operation, service and user manuals must be presented at least in the official language of the country in which the equipment will be used.
DECOMMISSIONING
43 | Estimated Life Span | 15 years
SAFETY AND STANDARDS
44 | Regulatory Approval / Certification | Free Sales Certificate (FSC) provided by any of the following countries: Australia, Canada, Japan, USA and/or European Community (e.g. FDA and/or CE certificate given by a third certified party for the specific medical devices proposed (no only declaration of conformity).
45 | Standards for the manufacturer | NMPA certificate provided (Chinese Suppliers only). Certified Quality Management System for medical devices (e.g. ISO 13485, or Good Manufacturing Practice - GMP) or General quality management (e.g. ISO 9001: Quality management systems – Requirements). Application of risk management to medical devices (e.g. ISO 14971).
46 | Standards for the product performance | Compliance to the following international standards or to regional or national equivalent, (including the technical tests for safety and performance from accredited laboratory or third party). Latest version recommended but compliance to previous standards versions could be accepted: - EN-ISO 2409 Paints and varnishes - Cross-cut test. - EN 288: Specification and approval of welding procedures for metallic materials. Part 3: welding procedures tests for the arc welding of steels. Or EN 25817: Arc-welded joints in steel. Guidance on quality levels for imperfections. - IEC:60601-2-52: Medical Electrical Equipment - Part 2-52: Particular Requirements For The Basic Safety And Essential Performance Of Medical Beds (only in case the . Alternative national equivalent tests can be considered.