## Foglio2
Radiographic mobile digital equipment
i | Version No. | 3
ii | Date of initial version | Version 1: 27/08/2012 (initial version); Version 2: 18/06/2014 (modification of the initial version)
iii | Date of last modification
iv | Date of publication | 19/11/2020
v | Completed / submitted by | WHO and IAEA working groups
Name, category and coding
1 | WHO Category / Code | (under development)
2 | Generic name | Mobile X-ray system
3 | Specific type or variation (optional) | Mobile, portable, digital
4 | UNSPS code (optional) | N.A.
5 | EMDN name | Mobile Radiographic Units
6 | EMDN code | Z11039016
7 | Alternative name/s (optional) | X-ray mobile equipment, mobile; general radiographic X-ray mobile equipment
8 | Alternative code/s (optional) | N.A
9 | Keywords (optional) | portable, imaging, radiography, mobile, radiology, digital radiography
Purpose of use
10 | Clinical or other purpose | General-purpose mobile diagnostic x-ray system used in a variety of x-ray imaging applications and to support out of Hospital infrastructure screening/diagnostic/monitoring interventions and applications. With regard to the COVID-19 pandemic and, in particular, COVID-19 related lung damage and connected complications, mobile radiographic digital technology offers benefits in chest imaging and assessing levels of COVID-19 probability and disease severity (1), anyway this equipment will serve for orthopedic and other general purposes too.
11 | Level of use (if relevant) | The mobile radiographic digital equipment is primarily used in the emergency unit and critical care unit/ICU of district general hospitals, specialized hospitals and health care centers. However, it can also be used in outreach operations (i.e. mobile clinics).
User of the equipment: radiographers/radiological technologists and medical professionals trained on safety, radiation protection and image management
12 | Clinical department/ward(if relevant) | Not relevant
13 | Overview of general functional requirements | The mobile radiographic digital equipment described here requires a direct digital radiography (DDR) technology with a flat-panel detector. The equipment consists of a wheeled cart that transports an X-ray generator, an X-ray tube, a tube telescopic arm support, collimators, on-board computer and a flat-panel digital detector. The equipment is a battery-powered unit, with a battery and charging system and a motor drive that is powered by the battery. The main AC line voltage is used to charge the batteries and, in some models, to power the equipment during standard operations (this dual power supply reduces the unit’s downtime by permitting the equipment to operate while recharging the batteries).
Technical characteristics and minimum requirements
14 | General technical requirements | Equipment easy to assemble/install/set up and be fully functional and operational as a complete standalone solution for acquisition, review, presentation, display, storage and transfer of radiographic images in a resource-limited setting.
Equipment: mobile, motorized, battery and AC power operated (AC mandatory for charging and, preferably, for standard working operations too).
Digital Imaging and Communications in Medicine (DICOM) 3.0 compatible image storage and transfer required.
The system should be capable of storing at least 2000 images with comprehensive post-processing options (Dose Area Product [DAP] meter to be integrated).
Capacity for removable media storage, to transfer data through different options (CD, DVD and/or USB), to send images through existing network port, and, preferably, to have wireless transfer of images through hospital wireless network (wireless and cable connections shall be provided).
Integrated Ethernet connectivity required.
At least 20 anatomical programmes shall be available. Software for COVID-19 detection should be included (preferable) (if available).
An integrated image review monitor to be included in the configuration. Anti-scatter grid or software for scatter correction (preferable) (if available). Equipment provided with DAP device/capability to record the patient dose.
15 | Detailed technical requirements | kVp range at least 40–120 kVp, digitally displayed. mAs range at least 0.5–200 mAs or more.
Minimum exposure time is 8.0 msec or less. Maximum exposure time 4 sec. Automatic exposure control facility (preferable) (if available).
Tube power rating at least 20 kW (measured at 100 kVp).
Rotating anode with dual focal spots and the maximum focal spot not higher than 1.3 mm (equivalent output/technology could be considered).
Heat storage capacity of the anode at least 120 000 HU. Cooling rate not less than 14 000 HU/min.
Total filtration at least of 2.5 mm aluminium equivalent.
An integrated/paired digital radiography flat-panel detector: wireless and/or with cable (wireless flat detector preferable); sub mm pixel size, active detector area not less than 35 x 43 cm.
Capability for alphanumeric annotation of images. Equipment total weight in the range 100–500 kg.
16 | Digital detector | Image quality: spatial resolution better than 3 lp/mm. Pixel pitch: < 150 x 150 µm.
Grayscale: at least 4096 (12-bit).
Preview image access time: less than 10 sec after X-ray exposure.
17 | Displayed and user-adjustable parameters and settings | Image to be displayed immediately after exposure.
Digital display of mAs, kV, KAP/DAP and an electronic timer. Low battery indicator/alarm.
Exposure status lights on main control and/or collimator (standby, ready up, exposure). Image display to be contrast- and brightness-adjustable, at least 18 inches diagonal size.
Exposures by remote control should also be possible, with operating distance higher than approximately 10 m.
The exposure release switch n the imaging suite should be detachable, with a cord of at least 5 m long.
18 | System components and other physical characteristics | An X-ray tube support with telescopic arm.
The tube stand must be fully counter-balanced for rotation in all directions. Articulated arm for imaging with any patient position.
Source to image receptor distance (SID) range not less than 100–200 cm. Frame with column/arm rotation range not less than ± 180 degrees.
Adjustable multi-leaf collimator, rotatable ± 90 degrees, with patient centring light. All cables shall be concealed in the arm system.
Collimation light to confirm the radiation field size.
Unit base wheels must be easily accessible for cleaning.
19 | Mobility and portability | When motor or battery is non-functional, free movement by pushing must be possible. Equipment speed capacity not less than approximately 1.5 km/hr.
Motorized movement capable of ascending slope of up to at least 7 degrees from horizontal.
The unit must have an effective system for parking, transport and emergency braking. Unit base wheels must be easily accessible for cleaning.
Utility requirements
20 | Power supply (voltage, frequency and plug vary across countries) | AC power input to be 120 and/or 220 V ± 10%, 50–60 Hz, single phase, fitted with compatible mains plug. X-ray exposures without power supply (battery mode exposure) preferable (if available).
Motor battery to be sealed lead-acid type, recharged by main unit power connection and recharging time not higher than 8 hours.
Battery total energy capacity up to at least 20 000 mAs.
Resettable overcurrent breaker to be fitted on both live and neutral supply lines.
Voltage corrector/stabilizer to allow safe and stable operations at ± 20% of local rated voltage (if necessary).
Accessories, consumables, spare parts, other components
21 | Accessories and spare parts | Must be supplied with protective dust cover at least for control panel.
To be supplied with at least 1 adult-size protective lead apron and 1 thyroid shield. Portable radiation hazard warning signs to be supplied with unit.
List of important spares and accessories to be provided with their part numbers and cost. Spare parts availability for the equipment lifespan (not less than 7 years).
Equipment provided with necessary quality assurance phantoms to check the image quality and calibration of the mobile radiography device. Phantom holder should be also provided as any quality assurance/control necessary tool.
Packaging
22 | Packaging, transport and environmental requirements | The system, for the shipment by air to end users, shall be packed in accordance with international standards that are applicable for the shipment by air of this kind of equipment.
Capable of being stored continuously in ambient temperature of 0–50 °C and relative humidity (RH) of 15–90%.
Capable of operating continuously in ambient temperature of 10–40 °C and RH 15–90%. Labelling on the primary packaging to include:
Name and/or trademark of the manufacturer.
Production year.
Model or product reference.
Information for particular storage conditions (temperature, pressure, light, humidity).
Safety, standards and warranty
23 | Standards, for the manufacturer and the equipment | Certified quality management system for medical devices (e.g. ISO 13485:2016 – Medical devices – Quality management systems – Requirements for regulatory purposes).
General quality management (e.g. ISO 9001:2015 – Quality management systems – Requirements). Application of risk management to medical devices (e.g. ISO 14971:2019 – Medical devices – Application of risk management to medical devices).
24 | Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country.
Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
25 | Standards, for product performance | Compliance to the following international standards or to regional or national equivalent, as applicable (including the technical tests for safety and performance from accredited laboratory or third party). Reference to the last available version is recommended but compliance to previous standards versions could be considered (as/when applicable).
IEC 60601-1:2005+AMD1:2012 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance.
IEC 60601-1-2:2014 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests.
IEC 60336:2005 X-ray tube assemblies for medical diagnosis – Characteristics of focal spots.
IEC 60601-1-3:2008+AMD1:2013 Part 1-3: General requirements for basic safety and essential performance
– Collateral standard: Radiation protection in diagnostic X-ray equipment.
IEC 60601-2-28:2017 Part 2-28: Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis.
IEC 60601-2-54:2009+AMD1:2015+AMD2:2018 Part 2-54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and fluoroscopy.
26 | Warranty and post- warranty | The system shall be covered by a warranty of at least 1 year, including all spare parts and labour, starting as of the date of successful on-site acceptance, as per testing and acceptance.
Warranty shall include all necessary spare parts, shipment to site, cost of replacement work, personnel, disposal of faulty parts, and software (patches, upgrades and updates).
Specific warranty inclusions and exclusions shall be clearly listed. At least two service visits per year shall be made under warranty.
List to be provided of important spares and accessories, with their part numbers and cost. Spare parts availability for the equipment lifespan (not less than 7 years).
List of equipment and procedures required for local calibration and routine maintenance shall be provided. Advanced maintenance tasks required shall be documented.
Costs and types of post-warranty service contract available shall be described.
The guaranteed time-period of support and software upgrade (if relevant) availability post-warranty shall be described.
Any variation to be indicated in the offer.
Additional services required Testing and acceptance • Factory acceptance test (FAT): the system, prior to shipment, shall be tested for conformity of the system with the manufacturer’s performance specifications and the minimum requirements specified. The FAT report shall be provided by the supplier. • Supplier shall perform, when applicable: installation, calibration, safety and operation checks before handover. • On-site acceptance test (SAT): the system, after delivery, shall be tested by the contractor together with the user to demonstrate that the performance meets the manufacturer’s performance specifications and the minimum requirements specified as determined by WHO, IAEA and users. The results of the SAT shall be documented in an acceptance test protocol that shall be signed by the end user (after consultation with the hospital medical physicist) and the manufacturer. • The results of the testing of the system shall be documented by the contractor in an acceptance protocol that shall be signed and dated by the user. Training and manuals • User, technical and maintenance manuals, hard and soft copies, to be supplied in English language (provision of versions in other UN languages, if available, will be an asset). Supplier has to describe any materials contained in the device that are classified as hazardous under local regulations. • Specific cleaning and disinfection instructions for IPC included (preferable). • Training of users in operation (including basic radiation protection topics and the use of quality assurance/control tools provided for testing), and training to technicians for basic maintenance shall be provided (provided also in an online format, if available). • Contact details of manufacturer, supplier and lists of local service agents to be provided together with the documentation.