## Foglio1
Bubble humidifier, non-heated
Published on: 19 November 2020 - Version: 2
Version 1: September 2019 (WHO-UNICEF publication)
Technical Specification | Minimum requirements
General technical requirements | Non-heated bubble humidifier. Graduated, transparent humidification bottle. Graduation should show minimum and maximum water level. Detachable metal or rigid durable polymer cap with gas. DISS inlet connector. 6 mm barbed (or specify alternate style) outlet. Humidification chamber working volume available between 150–500 mL. Graduation options available in metric, imperial and both units. Flow rate capacity up to 15 L/min. Pressure relief safety valve ≥ 14 kPa (0.14 bar, 2 psi). All components to be capable of disinfection including: bottle, diffuser, tubing, O-ring/seals, inlet and outlet connectors, cover lid in between patients. Supplier must define decontamination procedure. Bottle, diffuser and tubes made of polypropylene, polycarbonate or equivalent biocompatible plastic/ polymer certified for medical use, unbreakable or shatter resistant. Cap and connectors made of brass/steel/other biocompatible metal or polymer certified for medical use. Pressure valve made of brass chromium plated or equivalent metal certified for medical use.
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: Conforms to ISO, NFPA and/or CGA standards, and/or UL or CSA approved. ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen. ISO 18562 Biocompatibility evaluation of breathing gas pathways in healthcare applications. ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements. ISO 18190 Anaesthetic and respiratory equipment – General requirements for airways and related equipment. ISO 18562-1 Biocompatibility evaluation of breathing gas pathways in healthcare applications – Part 1: Evaluation and testing within a risk management process. ISO 8185 Respiratory tract humidifiers for medical use – Particular requirements for respiratory humidification systems.