## Foglio1
Mobile diagnostic X-ray digital system (C-arm)
i | Version No. | 1
ii | Date of initial version | 10 October 2019
iii | Date of last modification
iv | Date of publication
v | Completed / submitted by | Working group (WHO/IAEA)
Name, category and coding
1 | WHO Category / Code | (under development)
2 | Generic name | C-arm fluoroscopic X-ray unit
3 | Specific type or variation (optional) | Not applicable
4 | GMDN name | Mobile general-purpose fluoroscopic X-ray system, digital
5 | GMDN code | 37646
6 | GMDN category | 12 Diagnostic and therapeutic radiation devices
7 | UMDNS name | Radiographic Units, Mobile
8 | UMDNS code | 13272
9 | UNSPS code (optional) | 42201800
10 | Alternative name/s (optional) | Not applicable
11 | Alternative code/s (optional) | Not applicable
12 | Keywords (optional) | Brachytherapy
13 | GMDN/UMDNS definition (optional) | A mobile (within an imaging facility) general-purpose diagnostic fluoroscopic X-ray system that uses a C-arm and digital techniques for image capture, display and manipulation and is designed to be used in a variety of general-purpose applications requiring real-time fluoroscopic imaging capabilities. It includes spot-film capabilities in addition to the fluoroscopic features and is intended to optimize the capability of users to visually and quantitatively evaluate the anatomy and physiological function of various targeted body areas in real-time. It is frequently used in conjunction with an ingested or injected X-ray contrast medium. Images can be both real-time and delayed formats.
14 | EMDN code | Z11030704
15 | EMDN nomenclature | DIGITAL C-ARM REMOTE CONTROLLED RADIOGRAPHIC/FLUOROSCOPIC STATIONS
Purpose of use
16 | Clinical or other purpose | X-ray imaging of patients for radiotherapy treatment planning.
17 | Level of use (if relevant) | Hospital
18 | Clinical department/ ward(if relevant) | Radiation Oncology Department
19 | Overview of functional requirements | A unit providing planar X-ray imaging of patients for brachytherapy treatment planning. The mobile unit shall include stand, control console, display, C-arm, X-ray source and detector to allow imaging, including fluoroscopic imaging, of brachytherapy applicators. The unit shall be Digital Imaging and Communications in Medicine (DICOM)-compatible for transfer of images to a treatment planning system (TPS).
Technical characteristics
20 | Detailed requirements | • A high-frequency X-ray generator with potential in the range of 40–110 kV and X-ray tube with dual focal spots. • A fluoroscopic mode, including foot switch and last image hold capability as well as a radiographic mode. • A flat panel imaging detector with dimensions at least 21 cm x 21 cm. • Dual high-resolution flat display screens of at least 18-inch diameter with last image hold capability. • A C-arm gantry with orbital movement of at least 120° (– 30° to + 90°). • Motor-driven vertical travel greater than or equal to 40 cm. • Horizontal travel greater than or equal to 20 cm. • Focus-image receptor distance greater than or equal to 90 cm. • Free space greater than or equal to 76 cm. • A braking system for gantry control. • A DICOM 3.0 interface for digital networking of images (transfer of digital images to a brachytherapy TPS or hospital picture archiving and communication system [PACS]). • Hard disk image storage capability as well as external storage capability.
21 | Displayed parameters | Comprehensive parameters at control console.
22 | User adjustable settings | At control console
Physical/chemical characteristics
23 | Components(if relevant) | Major components of the C-arm fluoroscopic X-ray unit include X-ray tube, housing and generator, X-ray detector system, C-arm, display and control console.
24 | Mobility, portability(if relevant) | Mobile
25 | Raw Materials(if relevant) | Not applicable
Utility requirements
26 | Electrical, water and/or gas supply (if relevant) | Single-phase electrical power and air-conditioning.
Accessories, consumables, spare parts, other components
27 | Accessories (if relevant) | Not applicable
28 | Sterilization process for accessories (if relevant) | Not applicable
29 | Consumables / reagents (if relevant) | Not applicable
30 | Spare parts (if relevant) | Onsite spare parts kit is recommended.
31 | Other components (if relevant) | Not applicable
Packaging
32 | Sterility status on delivery (if relevant) | Not applicable
33 | Shelf life (if relevant) | Not applicable
34 | Transportation and storage (if relevant) | If storage is required prior to installation, the storage area should be secure, clean and dust-free, dry, cool, well-lit, ventilated and vermin-proof.
35 | Labelling (if relevant) | Symbols used according to ISO-15233 standards.
Environmental requirements
36 | Context-dependent requirements | C-arm fluoroscopic X-ray units are operated in brachytherapy rooms that provide radiation protection for staff and members of the public. It is the responsibility of the hospital to ensure that the manufacturer’s recommended operation and storage conditions, as given in their site planning guide, are respected. If the device requires that the operating and/or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure.
Training, installation and utilization
37 | Pre-installation requirements(if relevant) | It is the responsibility of the hospital to comply with requirements in the manufacturer’s site planning guide and ensure services are appropriate for operation of the equipment.
38 | Requirements for commissioning (if relevant) | Acceptance testing, commissioning of imaging system, baselines for quality control, comprehensive radiation survey.
39 | Training of user/s (if relevant) | Applications training in clinical operation, safety features and service mode.
40 | User care(if relevant) | The equipment shall only be used in accordance with departmental policies, procedures and work instructions, and consistent with manufacturer’s use instructions, to ensure safety and protection.
Warranty and maintenance
41 | Warranty | At least 12 months
42 | Maintenance tasks | Approximately 4 service days per year. Tasks as per manufacturer’s preventative maintenance schedule.
43 | Type of service contract | Full-service contract, including parts, service and repair or in-house maintenance engineering service.
44 | Spare parts availability post-warranty | 10 years minimum
45 | Software / Hardware upgrade availability | To be specified by manufacturer.
Documentation
46 | Documentation requirements | Comprehensive user manual, including emergency procedures.
Decommissioning
47 | Estimated Life Span | 10 years
Safety and standards
48 | Risk classification | US FDA: Device Class 2
49 | Regulatory approval/certification | Radiation regulatory approval for equipment possession, use and premises.
50 | International standards | IEC, Medical Electrical Equipment, Part 1: General Requirements for Safety, Rep. IEC 60601-1:2005+AMD1:2012. IAEA, Radiation Protection and Safety of Radiation Sources, Basic Safety Standards, No. GSR Part 3, 2014. IEC, Medical electrical equipment – Part 1–3: General requirements for basic safety and essential performance – Collateral Standard: Radiation protection in diagnostic X-ray equipment, IEC 60601-1-3:2008+AMD1:2013 CSV. IEC, Medical electrical equipment – Part 2–28: Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis, IEC 60601-2-28:2017. IEC, Medical electrical equipment – Part 2–54: Particular requirements for the basic safety and essential performance of X-ray equipment for radiography and radioscopy, IEC 60601-2-54:2009+AMD1:2015+AMD2:2018 CSV.
51 | Regional/local standards | Country-specific and regional standards may apply.
52 | Regulations | USA: 21CFR892.1650 Image-intensified fluoroscopic X-ray system Local radiation safety regulations may apply.