## Foglio1
Oxygen concentrator
Published on: 19 November 2020 - Version: 4
Version 3: September 2019 (WHO-UNICEF)
Version 2: 2015 (modification of the initial version)
Version 1: 2012 (initial version)
Technical Specification | Minimum requirements
General technical requirements | Provides a continuous flow of concentrated oxygen (> 82%) (preferably > 90%) from room air through one oxygen outlet. Continuous flow up to 5 L/min or 8 L/min or 10 L/min. Contains oxygen monitor to verify concentration. Requires continuous AC power source to operate. Power efficiency ≤ 70 W/L/min (preferable). User interface to be easy to operate; numbers and displays clearly visible and easily readable in low ambient light and sunlight. Digital or analogue meter that displays cumulative hours of device operation. Oxygen outlet(s) with 6 mm (1/4 inch) barbed fitting or equivalent. Oxygen outlet to be securely mounted and sheltered to reduce risk of being broken or bent. Flowmeter minimum flow rate of 0.5 L/min or less. Flowmeter adjustable, within minimum gradation intervals of 0.5 L/min for 5 L/min models, and 1 L/min for larger models. Noise level < 60 dB(A). Capable to be disinfected with hospital grade detergents. At least IP11 degree of protection to the harmful ingress of water (fluid spill resistance), preferable up to IP21. Mechanical shock resistance, mechanical vibration, electromagnetic compatibility and electrical safety testing. Capable of supplying the specified oxygen concentration continuously in ambient temperature from 10–40 °C, relative humidity from 15–85% (preferably up to 95%), and elevation from 0 to at least 2000 m. For operation at elevations higher than 2000 m, environmental requirements are less stringent; performance characteristics at such altitudes must be stated.
Displayed parameters | Oxygen flow rate (on flowmeter). Cumulative hours of operation.
User adjustable settings | Oxygen flow rate.
Alarms | Audible and/or visual alarms for: • Low oxygen concentration (< 82%). • Power supply failure. • High temperature. • Low battery (preferable). • Low high/no flow (preferable). • Low/high output pressure.
Accessories (included and mentioned in a disaggregated list) | • DISS and 6 mm barbed adaptor for each outlet (interchangeable between devices of different brands and models) (if applicable): 1 package of 20 per equipment. • Humidifier included, bubble and non-heated, single use is preferred (3 months’ supply required). Reusable may be acceptable with appropriate disinfection protocols.
Spare parts (included and mentioned in a disaggregated list) | 1-year spare parts kit as per preventive maintenance programme. Including: • Internal and external mounted filters for cleaning the air intake. • Spare battery set for alarm system (if applicable). • Spare mains power cable length ≥ 2.5 m (if applicable). • Replacement sets of spare fuses (if non-resettable fuses are used). • Sieve beds. Bidder must give a complete list of the specific spare parts included in their bid. Other spares that may be needed: circuit breaker, printed circuit board, sieve beds, compressor service kit, valves, wheels, motor capacitor, flowmeters and fan.
Mobility, portability | Whole unit to be movable with wheels on at least two castors. Unit weight to be < 27 kg.
Power supply, voltage, frequency and plug vary across countries | Equipment must be connected to a reliable and continuous source of energy. Operates from AC power electric line: 100–240 V/50–60 Hz. Main power cable and plug adapted for various countries. Mains power cable length ≥ 2.5 m. Electrical protection by resettable circuit breakers or replaceable fuses, fitted in both neutral and live lines. Single fuse in live line may be considered, but is less preferable.
Documentation requirements (English language mandatory) | Set: user and maintenance manuals, hard and soft copies, in English (mandatory) and other languages (preferable). Certificate of calibration and inspection. Troubleshooting, calibration and routine maintenance. List of all spare parts and accessories, with part numbers and contact details for parts supply. Document with contact details of manufacturer, supplier and local service agent.
Primary packaging label | Name and/or trademark of the manufacturer. Electrical power input requirements (voltage, frequency and socket type) and safety use and storage (keep away from oil, grease and petroleum-based or flammable products as well as smoking or open flames). Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: ISO 80601-2-69:2014 Medical electrical equipment – Part 2-69: Particular requirements for basic safety and essential performance of oxygen concentrator equipment. IEC 60601-1:2012 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-2:2014 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic disturbances – Requirements and tests. IEC 60601-1-6:2013 Medical electrical equipment – Part 1-6: General requirements for basic safety and essential performance – Collateral standard: Usability. IEC 60601-1-8:2012 Medical electrical equipment – Part 1-8: General requirements for basic safety and essential performance – Collateral standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems. IEC 60601-1-9:2013 Medical electrical equipment – Part 1-9: General requirements for basic safety and essential performance – Collateral standard: Requirements for environmentally conscious design. IEC 60601-1-11:2010 Medical electrical equipment – Part 1-11: General requirements for basic safety and essential performance – Collateral standard: Requirements for medical electrical equipment and medical electrical systems used in the home health-care environment. Compliance with ISO 8359 may be considered.
Warranty | 2 years with regards efficiency and quality of the product. Availability of accessories and spare parts for at least 2 years.