## Foglio1
Heat and moisture exchanger filter (HMEF)
Published on: 19 November 2020 - Version: 1
Technical Specification | Minimum requirements
General technical requirements | Heat and moisture exchanger filters (HMEF) for respiratory breathing systems (i.e. endotracheal tube); single use. 99.9% efficiency; humidification level > 30 mg H2O/L; dead space < 40 mL.
Sizes | Adult: 22 Fr/15 mm; transparent cover; maximum tidal volume (breathing filter): adapted for adult patient. Paediatric: connector 22 Fr/15 mm; transparent cover; maximum tidal volume (breathing filter): adapted for paediatric patient.
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity). If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word “LOT” (or equivalent harmonized symbol). Shelf life: 3 years.
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: ISO 9360-2:2001 Anaesthetic and respiratory equipment – Heat and moisture exchangers (HMEs) for humidifying respired gases in humans – Part 2: HMEs for use with tracheostomized patients having minimum tidal volumes of 250 ml. ISO 5367:2014 Anaesthetic and respiratory equipment – Breathing sets and connectors.