## Foglio1
Endotracheal tube introducer, bougie, sterile, single use
Published on: 19 November 2020 - Version: 1
WHO Item name: | Required technical specifications: | Dimensions / Sizes / Presentation: | Material: | Certificates for product safety and quality assurance, as requested or equivalent
Endotracheal tube introducer, bougie, sterile, single use, different sizes. | To assist with endotracheal intubations is used to guide the tube properly into the airway. Blue or yellow tube with graduated marking. Curved tip with distal rounded smooth tip. Initial sterilization method: ethylene oxide gas or Gamma radiation or equivalent as appropriate and applicable. Individually peel-packed in paper and/or plastic. At least the following sizes included in the set provided: 10 and 15 Fr. | Standard size: 10 Fr and 15 Fr, approximately 60–70 cm long | Flexible, medical grade, radiopaque, braided polyester base with a resin coating. | ISO 5361:2016: Anaesthetic and respiratory equipment – Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process.