## Foglio1
Endotracheal tube introducer, Stylet, sterile, single use
Published on: 19 November 2020 - Version: 1
WHO Item name: | Required technical specifications: | Dimensions / Sizes / Presentation: | Material: | Certificates for product safety and quality assurance, as requested or equivalent
Endotracheal tube introducer, stylet, sterile, single use, 10 Fr | Flexible and malleable guide (stylet) to be inserted into the endotracheal tube to guide it properly during the intubation. It has a soft and round end-tip. It can be shaped as needed. It has graduated marking. Manufacturer name and tube size are indicated on the tube. Initial sterilization method: ethylene oxide gas or Gamma radiation or equivalent as appropriate and applicable. Individually peel-packed in paper and/or plastic | Size 1: For endotracheal tube with internal diameter of 3.5–4.5 mm. Length: 30–45 cm, approximately. Diameter: 10Fr (3.3 mm). | Stylet: malleable metal alloy covered by polyvinyl chloride (PVC) coated, latex free; white; disinfectant resistant; withstand steam sterilization. | ISO 5361:2016: Anaesthetic and respiratory equipment – Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process.
Endotracheal tube introducer, stylet, sterile, single use, 14 Fr | Flexible and malleable guide (stylet) to be inserted into the endotracheal tube to guide it properly during the intubation. It has a soft and round end-tip. It can be shaped as needed. It has graduated marking. Manufacturer name and tube size are indicated on the tube. Initial sterilization method: ethylene oxide gas or Gamma radiation or equivalent as appropriate and applicable. Individually peel-packed in paper and/or plastic. | Size 2: For endotracheal tube with internal diameter of 4.5–6.0 mm. Length: 30–45 cm, approximately. Diameter: 14 Fr (4.6 mm). | Stylet: malleable metal alloy covered by polyvinyl chloride (PVC) coated, latex free; white; disinfectant resistant; withstand steam sterilization. | ISO 5361:2016: Anaesthetic and respiratory equipment – Tracheal tubes and connectors. ISO 10993-1:2018: Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process.