## Foglio1
Self-inflating neonatal resuscitation bag with masks
i | Version No. | 1
ii | Date of initial version | 2013-06-25
iii | Date of last modification | 2015-12-11
iv | Date of publication | 2015-12-30
v | Completed / submitted by | Working group (WHO/UNFPA/UNICEF)
Name, category and coding
1 | WHO Category / Code | (under development)
2 | Generic name | Resuscitator, hand operated neonatal, reusable.
3 | Specific type or variation (optional) | Multi-use self-inflating neonatal resuscitation bag with masks for preterm and full-term babies and low-birth-weight infants (less than 5 kg).
4 | GMDN name | Manual pulmonary resuscitator, reusable and resuscitator facemask, reusable.
5 | GMDN code | 17591 and 17170
6 | GMDN category | 02 Anaesthetic and respiratory devices, 09 Reusable devices
7 | UMDNS name | Resuscitators, pulmonary, manual [13-367]; resuscitators, pulmonary, manual, reusable [17-591].
8 | UMDNS code | 13-367 17-591
9 | UNSPS code (optional) | 42272303
10 | Alternative name/s (optional) | Resuscitation kit, resuscitation cabinets; CPR emergency kit; emergency kit, cardiopulmonary resuscitation, resuscitation bag and masks.
11 | Alternative code/s (optional) | S 37507; S 16161
12 | Keywords (optional) | Resuscitation bag, mask, resuscitator, reusable devices, valves
13 | GMDN/UMDNS definition (optional) | “Manual pulmonary resuscitator, reusable [17591] A hand-operated device designed to provide or assist ventilation in patients who are apnoeic or exhibit inadequate respiration. It typically employs entrained ambient air and includes a large flexible chamber that is hand ventilated, a gas reservoir, tubing, and a connector for attachment to a mask or endotracheal (ET) tube; oxygen (O2) from an O2 source may also be connected when necessary. It is used by emergency medical services (EMS) in ambulances, intensive care units (ICU), during internal patient transfer, accident and emergency (A&E), mass casualty incidents (MCI), and is generally placed strategically throughout a hospital. This is a reusable device.” “Resuscitator facemask, reusable [17170] A flexible, form-shaped device that is placed over a patient’s nose and mouth to direct ambient air, or medical oxygen (O2) and air, from a resuscitator to the upper airway and lungs. It is typically made of non-conductive sterilisable materials (e.g. silicone) that will create a gas-tight seal against the face. It will typically include a 15 mm and/or 22 mm connector and is available in a range of sizes (baby to adult). It will be directly attached to the resuscitator and held in place on the patient’s face by the operator. This device is intended for use with a breathing resuscitator but may be used for the delivery of anaesthesia gases. This is a reusable device."
Purpose of use
14 | Clinical or other purpose | Provide positive air pressure ventilation to newborn babies with asphyxia, babies who experience respiratory arrest, apnoea or respiratory distress requiring assisted ventilation, and babies who require assisted ventilation during procedures.
15 | Level of use (if relevant) | Health centre/district hospital/provincial hospital/specialized hospital/other health facilities that include maternity services.
16 | Clinical department/ ward(if relevant) | Nursing services, surgery, paediatrics, emergency medicine, obstetrics, intensive care unit, labour and delivery room.
17 | Overview of functional requirements | The resuscitator is used to ventilate newborns with a body weight less than 5 kg. The resuscitator can be used to efficiently maintain ventilation, or as resuscitation in other critical situations.
Technical characteristics
18 | Detailed requirements | A resuscitator is used to ventilate a neonate with a body weight of less than 5 kg. It is operated by hand and ventilation can be done with ambient air or with oxygen. A resuscitator can be totally disassembled, and is easy to clean and disinfect. All parts are manufactured from high-strength, long-life materials that require no special maintenance or storage conditions. A resuscitator is supplied as a complete set with: • non-rebreathing patient valve with a pressure limiting valve so that airway pressure does not exceed 4.5 kPa (45 cmH2O) and can transmit an airway pressure of at least 3 kPa (=30 cmH2O); • masks, translucent, in two different sizes: Size 0 (preterm and low-birth-weight baby), round type, outer diameter 35–50 mm; Size 1 (term baby), round type, outer diameter 50–65 mm silicone rubber or any material fulfilling at least the standards ISO 10993-1:2009; ISO 10993-5:2009; ISO 10993-10:2010, or equivalent; or classified as USP Class V; • compressible self-refilling ventilation bag: silicone rubber or any other material fulfilling ISO 10651-4; • bag size: 200–320 mL; intake valve with an optional nipple for O2 tubing: polycarbonate/polysulfone or any other material fulfilling the ISO 10651-4 or any other equivalent; • bag made of silicone and valve made of polycarbonate or polysulfone or any other sterilizable material complying with ISO10651-4 or equivalent • material: polycarbonate/polysulfone or any other sterilizable material fulfilling at least ISO 10651-4.
19 | Displayed parameters | N.A.
20 | User adjustable settings
Physical/chemical characteristics
21 | Components(if relevant) | Self-inflating neonatal resuscitation bag with masks for preterm and term babies. Patient valves with pressure relief valves o 45 cmH2O
22 | Mobility, portability(if relevant) | Portable and mobile.
23 | Raw Materials(if relevant) | Recommended material is silicone rubber for the bag and mask and polycarbonate/polysulfone for the valves. Any material fulfilling ISO 10993-1:2009; ISO 10993-5:2009; ISO 10993-10:2010 or equivalent or USP Class V is also recommended.
Utility requirements
24 | Electrical, water and/or gas supply (if relevant) | N.A.
Accessories, consumables, spare parts, other components
25 | Accessories (if relevant) | Masks for preterm, low-birth-weight babies (Size 0) and term babies (Size 1) (see Section 18 of this table for details).
26 | Sterilization process for accessories (if relevant) | The resuscitator and mask must be cleaned and disinfected after each use; following the manufacturer's instruction is recommended.
27 | Consumables / reagents (if relevant) | N.A.
28 | Spare parts (if relevant) | N.A.
29 | Other components (if relevant) | N.A.
Packaging
30 | Sterility status on delivery (if relevant) | Cleaned and disinfected before use.
31 | Shelf life (if relevant) | N.A.
32 | Transportation and storage (if relevant) | Primary packaging: unit of use. One resuscitator set in a plastic bag + box with the manufacturer's instruction for use. Labelling on the primary packaging: name and/or trademark of the manufacturer. Manufacturer's product reference. Type of product and main characteristics. If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word "LOT" (or equivalent harmonized symbol) (if applicable). Information for particular storage conditions (temperature, pressure, light, humidity, etc.) as appropriate (or equivalent harmonized symbol). Information for handling, if applicable (or equivalent harmonized symbol).
33 | Labelling (if relevant) | Symbols used according to ISO 15223 standards.
Environmental requirements
34 | Context-dependent requirements | N.A.
Training, installation and utilization
35 | Pre-installation requirements(if relevant) | N.A.
36 | Requirements for commissioning (if relevant) | N.A.
37 | Training of user/s (if relevant) | The resuscitator should only be operated someone who has received adequate basic training in resuscitation technique.
38 | User care(if relevant) | After each dismantling and cleaning, the resuscitator must be reassembled and tested to make sure that it works correctly. In view of its use, the item is considered an "emergency resuscitation item", which means that it must always be readily available and in a good working condition. Following the manufacturer's instructions manual is recommended.
Warranty and maintenance
39 | Warranty | Minimum 1 year.
40 | Maintenance tasks | N.A.
41 | Type of service contract | N.A.
42 | Spare parts availability post-warranty | N.A.
43 | Software / Hardware upgrade availability | N.A.
Documentation
44 | Documentation requirements | The resuscitator is supplied in a box as a complete set, with clear instructions/diagrams for use and assembly in the language of the user and a list of accessories/parts. The resuscitator is supplied with instructions and diagrams covering its functionality, how to use it, how to dismantle and assemble it, and how to clean, disinfect and sterilize it. It is supplied with an instruction manual covering item descriptions and functionality, including usage, maintenance and a list of spare parts.
Decommissioning
45 | Estimated Life Span | 5 years.
Safety and standards
46 | Standards, for the manufacturer and the equipment | Class B (GHTF Rule 6), Class II a (EU, Australia), Class II (Canada, Japan).
47 | Regulatory Approval / Certification | Optional: • registered in country of import, if applicable; • approval from the national regulatory agency of the manufacturer’s country; • FDA registration (United States); or • CE Mark (EU) for class IIa, with notified body number, or approval from other regulatory body in an IMDRF founding member country.
48 | International standards (for reference) | ISO 13485:2012 Medical devices - Quality management systems – Requirements for regulatory purposes: ISO 10651-4:2002 Lung ventilators -- Part 4: Particular requirements for operator-powered resuscitators (EN 13544-2 implied). For devices not intended to be used with oxygen, oxygen related clauses are optional For facemask (if not made of silicone): ISO 10993-1:2009 Biological evaluation of medical devices – Part 1: Evaluation and testing for mask ISO 10993-5:2009 Biological evaluation of medical devices – Part 5: Tests for in vitro cytotoxicity ISO 10993-10:2009 Biological evaluation of medical devices – Part 10: Tests for irritation and delayed-type hypersensitivityor classified as USP class V Optional: ISO 14971:2012 Medical devices – Application of risk management to medical devices ISO 5356-1:2004 Anaesthetic and respiratory equipment – Conical connectors - Part 1: Cones and sockets BS EN 980:2008 Graphical symbols for use in the labelling of medical devices ISO 15223-1:2012 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplies – Part 1: General requirements BS EN 1041:2008 Information supplied by the manufacturer with medical devices
49 | Regional/local standards | Any national standards equivalent/harmonized to International standards.
50 | Regulations | Directive 93/68/EEC (CE Mark) Directive 2001/104/EC Directive 2007/47/EC