## Foglio1
Electrocardiograph, portable
i | Version No. | 2
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2020-11-13
iv | Date of publication | 2020-11-19
v | Completed / submitted by | WHO working group
Name, category and coding
1 | WHO Category / Code | (under development)
2 | Generic name | Electrocardiograph
3 | Specific type or variation (optional) | multi-channel
4 | CND code (https://ec.europa.eu/health/ md_topics-interest/overview_en) | Z120503, Z12050301
5 | CND nomenclature | ELECTROCARDIOGRAPHS, GENERAL PURPOSE ELECTROCARDIOGRAPHS
Purpose of use
6 | Clinical or other purpose | Detect, measure, and display a patient's electrocardiogram (ECG) through leads and sensors attached to the patient
7 | Level of use (if relevant) | Health centre, district hospital, provincial hospital, specialized hospital and outreach (mobile clinics)
8 | Clinical department/ ward(if relevant) | Emergency room (ER), intensive care unit (ICU), surgery, outpatient, hospital
9 | Overview of functional requirements | Recorder of patient ECG and heart rate with printing and displaying capabilities. 12 electrographic standard leads measuring with simultaneous acquisition. Settable alarms for low or high heart rate, ventricular fibrillation and other cases. Operates from mains voltage or from internal rechargeable battery. Patient connectors that are sterilizable and reusable are preferred, though reusable cables that attach to disposable connection patches are also acceptable. Hard copy printout of traces will be included. Display to be either clear screen, readable in direct sunlight, and/or durable paper printout.
Technical characteristics
10 | Detailed requirements | Automatic equipment calibration. 12 electrographic standard leads measuring with simultaneous acquisition. Visualization of at least one group of 3 leads simultaneously. At least 1 rest ECG patient cable with at least 5 terminals. LCD or TFT colour monitor display of at least 7 inches with visualization of analogical curves, alphanumeric values measured and the related physiological limits. Frequency response range at least: 0.05–150 Hz. Equipment compatible with patients with pacemakers. Protection against defibrillation. Automatic internal data storing up to at least 40 ECG records. Input impedance not less than 5 MΩ (preferably higher than 10 MΩ). Patient leakage current not higher than 10 uA. Common mode rejection (CMR) not less than + 100 dB (preferably + 120 dB). ECG signal measurement range not smaller than -2 mV to +2 mV. At least the following adjustable alarms: a) heart failure b) ventricular fibrillation c) tachycardia d) bradycardia e) electrode disconnection. Equipment provided with filters at least for: baseline instability, AC 50/60 Hz interference, rumours, low- pass, etc. Equipment provided with at least the following software applications: a) arrhythmias detection and analysis b) ventricular fibrillation detection and analysis c) ventricular tachycardia detection and analysis d) ST segment analysis. Integrated printer using standard A4 format paper. Selectable printing paper speed at least of 5, 25 and 50 mm/sec.
11 | Displayed parameters | Arrhythmia detection basic. Numeric display. Graphical and numerical trends and physiological curves display (preferable).
12 | User adjustable settings | N/A
Physical/chemical characteristics
13 | Components(if relevant) | There must be no sharp edges on the unit. Equipment provided with data management capabilities: RS232 or USB port for PC connection and data transmission. Provided with case for transport when available.
14 | Mobility, portability(if relevant) | Easy and safe portable equipment
15 | Raw Materials(if relevant) | N/A
Utility requirements
16 | Electrical, water and/or gas supply (if relevant) | Powered by both mono-phase electrical source and internal, rechargeable, replaceable batteries. 110–220 V, 60–50 Hz, ± 10% mono-phase electrical source with line-conection plug type in compliance with specific national standards and regulations. Rechargeable batteries with at least the following characteristics: a) automatic switch from electric-line mode to battery operating mode and vice-versa; b) battery capacity: at least 30 ECGs or 30 min of continuous recording; c) low battery visual alarm; d) 100% high capacity batteries with re-charging time not greater than 6 hours. e) battery charger to operate from input supply 110–220 V, 60–50 Hz, ± 10% (battery charger built-in or external). Appropriate external device to protect the equipment against over-voltage and over-current line conditions (between plug and socket). Voltage corrector/stabilizer to allow operation at ± 30% of local rated voltage.
Accessories, consumables, spare parts, other components
17 | Accessories (if relevant) | 12 leads ECG patient cable. Sets of disposable electrodes for adult and paediatric patients. Sets of reusable electrodes (suction ball-type chest electrodes and/or extremity clamp type). Set of Limb electrodes. Rest ECG patient cable with at least 5 terminals. Any accessory or dedicated device necessary to the proper functioning and utilization of the equipment is included. Any accessory, consumable and/or part needed to the proper equipment installation, calibration, training, testing and commissioning is included.
18 | Sterilization process for accessories (if relevant) | Not required
19 | Consumables / reagents (if relevant) | Bottle/tubes of conductive gel for body-connection electrodes. Pocket/sets of recording thermal paper.
20 | Spare parts (if relevant) | Sets of rechargeable batteries, supplied empty of charge. Sets of spare fuses.
21 | Other components (if relevant) | N/A
Packaging
22 | Sterility status on delivery (if relevant) | N/A
23 | Shelf life (if relevant) | N/A
24 | Transportation and storage (if relevant)
25 | Labelling (if relevant) | N/A
Environmental requirements
26 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0–50 °C and relative humidity of 15–85%, preferably 90%. Capable of operating continuously in ambient temperature of 10–40 °C and relative humidity of 15–85%, preferably 90%.
Training, installation and utilization
27 | Pre-installation requirements(if relevant) | Not required.
28 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation. Supplier to perform installation, safety and operation checks before handover.
29 | Training of user/s (if relevant) | Training of users in operation and technicians in basic maintenance shall be provided.
30 | User care(if relevant) | The unit has to be cleanable and disinfectable with alcohol or chlorine wipes.
Warranty and maintenance
31 | Warranty | Not less than 2 years. Specific inclusions and exclusions to be listed. Contact details of manufacturer, supplier and local service agent to be provided
32 | Maintenance tasks | List of equipment and procedures required for local routine maintenance should be provided. Advanced maintenance tasks required shall be documented.
33 | Type of service contract | Costs and types of post-warranty service contract available shall be described.
34 | Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty should be pointed out.
35 | Software / Hardware upgrade availability | Software updates Included during warranty. Upgradable device with additional programs and software fuctions (when available). Guaranteed time period of support and updates (when necessary) availability post-warranty shall be described.
Documentation
36 | Documentation requirements | User/technical manual to be supplied in English (provision of versions in other UN languages, if available, will be an asset). Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration, if necessary, and routine maintenance. Battery disposal according local laws. Contact details of manufacturer, supplier and local service agent to be provided.
Decommissioning
37 | Estimated Life Span | Not less than 7 years
Safety and standards
38 | Standards, for the manufacturer and the equipment | Certified quality management system for medical devices (e.g. ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory purposes). General quality management (e.g. ISO 9001:2015 Quality management systems – Requirements). Application of risk management to medical devices (e.g. ISO 14971:2019 Medical devices – Application of risk management to medical devices).
39 | Regulatory Approval / Certification | Free sales certificate (FSC). Certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
40 | International standards | Compliance to the following international standards, when applicable, or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party). Reference to the last available version is recommended but compliance to previous standards versions could be considered (as/when applicable). IEC 60601-1:2012 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-2:2007 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. IEC 60601-1-8:2012 Medical electrical equipment – Part 1-8: General requirements for basic safety and essential performance – Collateral standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems. IEC 60601-2-25:2011 Part 2-25: Particular requirements for the basic safety and essential performance of electrocardiographs. IEC 60601-2-27:2011 Part 2-27: Particular requirements for the basic safety and essential performance of electrocardiographic monitoring equipment. IEC 60601-2-49:2011 Part 2-49: Particular requirements for the basic safety and essential performance of multifunction patient monitoring equipment.