## Ophthalmoscope, direct
OPHTHALMOSCOPE, DIRECT
i | Version No. | 1
ii | Date of initial version | 2012-06-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
NAME, CATEGORY AND CODING
1 | WHO Category / Code | (under development)
2 | Generic name | Ophthalmoscope
3 | Specific type or variation (optional) | direct, battery-powered
4 | GMDN name | Direct ophthalmoscope, battery-powered
5 | GMDN code | 46786
6 | GMDN category | 08 Ophthalmic and optical devices
7 | UMDNS name | Ophthalmoscopes, Direct
8 | UMDNS code | 12817
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Ophthalmic ophthalmoscope
11 | Alternative code/s (optional) | MS 46785
12 | Keywords (optional) | funduscope, funduscopy, fundoscope, fundoscopy, retina, diagnostic set.
13 | GMDN/UMDNS definition (optional) | Visualizes the interior of the eye, with the instrument relatively close to the subject's eye and the observer viewing an upright magnified image.
PURPOSE OF USE
14 | Clinical or other purpose | Designed to visualize the interior of the eye, with the instrument relatively close to the subject's eye and the observer viewing an upright magnified image.
15 | Level of use (if relevant) | Health post, Health center, District hospital, Provincial hospital, Specialized hospital
16 | Clinical department/ward(if relevant) | ophthalmology
17 | Overview of functional requirements | A light source illuminates the eye and the image is viewed through a magnifying lens. The lens is selectable either within the unit or be exchanging eyepieces. The unit is handheld with internal batteries.
TECHNICAL CHARACTERISTICS
18 | Detailed requirements | At least 2.5 V Xenon or Halogen light source. Magnification up to x15, in several selectable stages from direct vision to maximum magnification. Anti-reflection lens. At least 3 apertures and fixation star. Range of lenses not smaller than -20D to +20D with steps not greater than 1D. Dust free sealed optics and aspherical optical system. Red-free, blue and polarization filters.
19 | Displayed parameters
20 | User adjustable settings
PHYSICAL/CHEMICAL CHARACTERISTICS
21 | Components(if relevant) | Dust-free sealed optics. Supplied in protective, reclosable container. On/off switch to be robust and easy to use. External material to be non-ferrous.
22 | Mobility, portability(if relevant) | Handheld unit, single piece when in use.
23 | Raw Materials(if relevant) | N/A
UTILITY REQUIREMENTS
24 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug. Battery compartment (if reusable) to be sealed against liquid ingress, yet easily opened. Internal batteries, rechargeable preferred compatible with both 2.5 V and 3.5 V batteries or handles provided; Led display indicating the charging status. Charger electrical source requirements: Amperage: ______; Voltage: ______.
ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS
25 | Accessories (if relevant) | Two spare bulbs.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant)
29 | Other components (if relevant)
PACKAGING
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
ENVIRONMENTAL REQUIREMENTS
34 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of 10 to 40 deg C and relative humidity of 15 to 90%., Liquid splash resistant.
TRAINING, INSTALLATION AND UTILISATION
35 | Pre-installation requirements(if relevant) | Local clinical staff to affirm completion of installation
36 | Requirements for commissioning (if relevant)
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant)
WARRANTY AND MAINTENANCE
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
DOCUMENTATION
44 | Documentation requirements | Advanced maintenance tasks required shall be documented User, technical and maintenance manuals to be supplied in ************** language. List to be provided of equipment and procedures required for local calibration and routine maintenance, List to be provided of important spares and accessories, with their part numbers and cost. Certificate of calibration and inspection to be provided. Contact details of manufacturer, supplier and local service agent to be provided
DECOMMISSIONING
45 | Estimated Life Span | 10 years
SAFETY AND STANDARDS
46 | Risk Classification | Class A (GHTF Rule 12);Class II (USA); Class I (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 10942:2006 Ophthalmic instruments -- Direct ophthalmoscopes ISO 15004-1:2006 Ophthalmic instruments -- Fundamental requirements and test methods -- Part 1: General requirements applicable to all ophthalmic instruments ISO 15004-2:2007 Ophthalmic instruments -- Fundamental requirements and test methods -- Part 2: Light hazard protection
49 | Reginal / Local Standards
50 | Regulations | US regulations 21 CFR part 820 21 CFR section 886.1570 ophthalmoscope, battery-powered JP regulations MHLW Ordinance No.169 12817000 Direct ophthalmoscope