## Foglio2
Computed Tomography (CT) System (multi-slice)
i | Version No. | 1
ii | Date of initial version | 2020-11-19
iii | Date of last modification
iv | Date of publication
v | Completed / submitted by | WHO and IAEA working groups
Name, category and coding
1 | WHO Category / Code | (under development)
2 | Generic name | Computed Tomography scanning system
3 | Specific type or variation (optional)
4 | UNSPS code (optional) | N.A.
5 | EMDN name | COMPUTED TOMOGRAPHS (CT)
6 | EMDN code | Z110306
7 | Alternative name/s (optional) | CT scanner, computed tomography unit, CT system, helical CT scanner, multi-detector CT scanner, spiral CT scanner, volume CT scanner
8 | Alternative code/s (optional)
9 | Keywords (optional) | tomography, imaging, slices, CT, computed, scanner, scanning
Purpose of use
10 | Clinical or other purpose | Computed tomography technology figures centrally in evidence-based patient care algorithms for many clinical indications. Applications of a multi-slice unit range from the complex imaging of manifestations of infections like COVID-19 and tuberculosis to cardiovascular conditions, chronic lung diseases, trauma, diabetes complications, the most common types of cancer, and other pathologies.
11 | Level of use (if relevant) | CT systems are used primarily in imaging departments of district general hospitals and specialized hospitals.
12 | Clinical department/ward(if relevant) | Not relevant
13 | Overview of general functional requirements | CT scanning system to be used as a complete standalone solution for acquisition, review, display, storage (CT console and workstation) and transfer of images in resource-limited settings. It mainly consists of: an X-ray system, a patient table, a gantry and a PC control. A high-voltage X-ray generator supplies power to the X-ray tube, which has a rotating anode and is capable of withstanding high heat loads generated during multi-detector acquisition. The gantry includes the X-ray generator, detector system, X-ray tube, collimators and rotational frame.
Technical characteristics and minimum requirements
14 | Functional and performance requirements and scanning parameters | CT scanning system with at least 64 slices suitable to perform a wide range of scans in patients of all ages and a diverse range of diseases, with a special focus on lung diseases.
The CT scanner should be suitable for all relevant cross-sectional imaging requirements in diagnostic radiology.
Number of actual slices per 360° rotation should be not less than 64 slices at all speeds. Reconstructed slice width options range from at least 0.625 mm to 10 mm.
Minimum rotation time (360°) not higher than 0.5 sec.
Retrospective reconstruction should be possible of raw data files with the capacity to modify parameters such as field of view (FOV).
At least the following scanning modes should be possible: scan projection radiograph (SPR), axial and spiral.
The SPR length should be at least 1500 mm and the minimum width 500 mm. It must be possible to obtain the SPR from anteroposterior (AP) or posteroanterior (PA) or left to right or right to left directions.
Multiple volumetric studies capability.
The system must have automated dose control and milliampere (mA) control software that automatically adjusts mA for patient size, adjusts mA along the z-axis, and modulates mA during rotation.
The CT scanner should have conventional built-in lasers or light beams, which indicate the coincidence of the centre of rotation and scan position.
The system should be interconnected (all the workstations, any laser systems, printers etc.) and the CT scanner should be able to be networked at the site to allow transfer of CT data sets in Digital Imaging and Communications in Medicine (DICOM) format to a hospital information system (HIS), radiology information system (RIS) and picture archiving and communication system (PACS) system.
Dose computation and display: the system should display CT dose index (CTDI) – volumetric (CTDIvol) or weighted (CTDIw) – and dose-length product (DLP) and have the capacity to transfer this information to the exam record.
Dose optimization tools should be available.
15 | Patient table | The CT scanner should have a tabletop made of resistant and radiotransparent material. Table with minimum dimensions of 235 x 40 cm (preferable).
Vertical moving minimum range 44–90 cm (preferable).
The speed of horizontal movement must be variable with a maximum speed of at least 100 mm per sec. The scannable range should be at least 160 cm.
Maximum loading capacity of not less than 200 kg without any change in stated performance technical specifications (such as the positioning accuracy).
Capability of automatic patient isocentring (preferable).
16 | Gantry | Number of rows not less than 64.
A gantry aperture of at least 70 cm. Scan FOV of at least 50 cm.
Balancing system and laser lights to support centring.
Preferable: gantry tilt at least ± 30 degrees (preferable), controllable by the console; also (remote).
17 | X-ray system (tube and generator) and detectors | High-frequency generator with gantry implanted microprocessor.
The X-ray tube shall have dual focal spots. The contractor shall state the size of the focal spots. Tube focal spots not higher than 0.9 x 0.8 mm and 1.1 x 1.2 mm.
Power rating of at least 70 kW.
kVp range of at least 80 kV to 135 kV. mA range of at least 20 mA to 550 mA.
Maximum anode cooling rate of at least 1000 kHU/min. Anode heat storage capacity of not less than 7 MHU.
High performance, low noise, high data density, active response data acquisition system. Solid state detectors.
At least 38 mm detector coverage with no less than 64 rows Free from repeated calibrations.
There shall be at least 600 elements per row and 64 or more detectors for acquisition of a minimum of 64 slices at a time.
High detection efficiency (to be specified by the contractor). High discharging speed (to be specified by the contractor).
18 | Hardware, console and software | Raw data reconstruction hardware:
· Latest generation system with high performance processors.
· 4 GB of RAM.
· Raw data hard disk of at least 300 GB.
· High multi-tasking level.
Double monitor acquisition (operation) console with at least the following characteristics:
· Capability for simultaneous scanning and reconstruction.
· At least the following functions should be available: simultaneous scanning and routine analysis, simultaneous scanning and archiving and/or hard copying and simultaneous scanning and transfer to second console/workstation.
· Last generation processor.
· At least two 19’’ thin-film transistor liquid-crystal display (TFT LCD or equivalent or better technologies) colour monitors medical safety requirements compliant.
· Keyboard and mouse.
· At least 2 GB of RAM.
· At least 250 GB image storage hard disk.
· CD/DVD driver with integrated CD reviewing software for DICOM stored images.
· The CT system should be fully DICOM compliant. The DICOM should support at least the following: DICOM
o print, storage, send/receive, and query/retrieve.
· DICOM compliance statement should be provided.
· DICOM data high transmission speed to and from the workstation.
· Verbal/audio bi-directional communication system must be provided between the operator and the patient.
· At least local area network (LAN) connection driver.
Acquisition (operation) console’s software at least with the following functions:
· Easy access software interface.
· Positioning digital radiography with a length of at least 160 cm.
· Initial definition of the exam’s protocol with real-time availability to modify it.
· List of defined exam’s protocols.
· Spiral and axial (preferable) scan.
· For contrast-enhanced exams, scan synchronization with the arrival of the administered contrast (bolus test and bolus tracking).
· Multi-perspective and cardiovascular examination software with at least the following requirements:
· best reconstruction timing choice;
· electrocardiogram (ECG) tracing and image-gating procedure;
· automatic multi-phase reconstruction of cardiac kinetic function.
· Software for low dose protocols included.
Independent double (preferable) monitor workstation with at least the following characteristics:
· High-performance workstation for post-processing and advanced clinical applications.
· Latest generation processor.
· One (preferably two) monitor, at least 21”, LCD or equivalent or better technologies, medical safety requirements compliant.
· Keyboard and mouse.
· At least 4 GB of RAM.
· Not less than 1 TB image storage hard disk.
· In the hard disk for image storage, the number of uncompressed 512 x 512 pixel images that can be stored should be at least 200 000. The maximum possible hard disk capacity should be provided.
· CD/DVD driver with integrated CD reviewing software for DICOM stored images.
· The CT system should be fully DICOM compliant. The DICOM should support at least the following: DICOM print, storage, send/receive, and query/retrieve.
· DICOM compliance statement should be provided.
· DICOM data high transmission speed to and from the workstation.
· At least LAN connection driver.
Software features for the independent workstation described in the previous point:
· Volume, partial, package or radial maximal intensity projection (MIP) and minimum intensity projection (MinIP).
· Volume, partial, package or radial volume rendering technique (VRT).
· Volume, package, radial multiplanar reformation (MPR).
· Single curve, package or radial MPR.
· Surface 3-D.
· Dynamic scan with or without patient table movement.
· Angio package with bolus-CT capability.
· CT cardiovascular package, with at least the following functions available: calculation of the calcium density/quantity in the coronaries, functional/anatomic and morphologic analysis of cardiac system (coronaries and left ventricle) and coronary plaques characterization software.
· Vessel structure analysis package with 2-D and 3-D measurement (angio-CT software) (preferable).
· CT neuro package (neuro digital subtraction angiography [DSA] CT and neuro perfusion CT)(preferable).
· Lung examination/analysis full package provided. Moreover, software for COVID-19 detection/assessment should be considered as an option, if available.
· CT oncology package (lung nodules detection with computer assisted detection [CAD] capability and colonography with CAD capability) (preferable).
· Automatic bones subtraction package (preferable).
· Analysis and comparison of tumours for serial exams on the same patient (preferable).
· All the software should be provided for reporting procedure with DICOM, pdf, rtf format export availability to enable transfer of images for tele-radiology to electronic health records, to other medical facilities or to others as warranted.
19 | Image quality | The reconstruction matrix should be at least 512 x 512 pixel. Display matrix not less than 1024 x 1024 pixel.
In plane spatial resolution values shall be provided by the contractor to be evaluated. Low-contrast detectability (or resolution) 5 mm or less at 0.3% at no more than 20 mGy.
High-contrast spatial resolution: at least 18 lp/cm at 0% modulation transfer function (MTF). Metal artefacts reduction algorithm (preferable).
Reconstruction index of at least 190 slice/sec (3 frames/sec).
Reconstruction speed, in both spiral and axial scan, with 512 x 512 pixel matrix, with every artefact correction, of at least 16 images/sec VRT.
Table and bone removal (image subtraction) function available. Virtual endoscopic view available (preferable).
Utility requirements
20 | Power supply (voltage, frequency and plug vary across countries) | Power input to be approximately 380 V, 50/60 Hz, 90 kVA, triphasic electrical source.
The provider will supply direct electric connection to the three-phase power supply network and the connection will be effectuated directly to the network with thermomagnetic disconnecting switch.
Protectors against power surge (over-voltage and over-current) line conditions.
UPS provided: online UPS with maintenance-free batteries for the backup of the entire system for at least 30 minutes.
Accessories, consumables, spare parts, other components
21 | Accessories and spare parts | Patient table provided with complete accessories set for patient positioning for any exam type.
Necessary quality assurance phantoms to check the image quality and calibration of the CT scanner should be provided.
Phantom holder provided. Contrast injector system.
At least two complete sets of necessary protective equipment for the staff/users. Each set will include lead: apron, thyroid collar, specific gloves, glasses and a mask.[1]
List of important spares and accessories to be provided with their part numbers and cost. Spare parts availability for the equipment lifespan (not less than 10 years).
Packaging
22 | Packaging, transport and environmental requirements | The system should have all safety markings in English.
System packaging: for the shipment by air to the end user, should be packed in accordance with international standards that are applicable for the shipment by air of this kind of equipment.
Capable of being stored continuously in ambient temperature of 0–50 °C and relative humidity (RH) of 15–90%.
Capable of operating continuously in ambient temperature of 10–40 °C and RH of 15–90% (preferable). Labelling on the primary packaging to include:
· Name and/or trademark of the manufacturer.
· Production year.
· Model or product reference.
· Information for particular storage conditions (temperature, pressure, light, humidity).
Safety, standards and warranty
23 | Standards, for the manufacturer and the equipment | Certified quality management system for medical devices (e.g. ISO 13485:2016).
General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
24 | Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country.
Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
25 | Standards, for product performance | Compliance with the following international standards or to regional or national equivalent, as applicable (including the technical tests for safety and performance from accredited laboratory or third party).
Reference to the last available version is recommended but compliance to previous standards versions could be considered .
· IEC 60601-1:2005+AMD1:2012 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance.
· IEC 60601-1-2:2014 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests.
· IEC 60601-2-44 Particular requirements for the basic safety and essential performance of X-ray equipment for computed tomography.
· IEC 60601-1-3:2008+AMD1:2013 Part 1-3: General requirements for basic safety and essential performance – Collateral standard: Radiation protection in diagnostic X-Ray equipment.
· IEC 60336:2005 X-ray tube assemblies for medical diagnosis – Characteristics of focal spots.
· IEC 60601-2-28:2010 Part 2-28 Particular requirements for the basic safety and essential performance of X-ray tube assemblies for medical diagnosis.
· IEC 60613:2010 Electrical and loading characteristics of X-ray tube assemblies for medical diagnosis.
26 | Warranty and post- warranty | The system should be covered by at least 1 year’s warranty including spare or replacement parts, tube and labour, starting as of the date of successful on-site acceptance, as per testing and acceptance below.
Warranty should include all necessary spare parts, X-ray tube, shipment to site, cost of replacement work, personnel, disposal of faulty parts and software (patches, upgrades and updates).
Specific warranty inclusions and exclusions should be clearly listed. At least two service visits per year should be made under warranty.
List to be provided of important spares and accessories, with their part numbers and cost. Spare parts availability for the equipment lifespan (not less than 10 years).
List of equipment and procedures required for local calibration and routine maintenance should be provided, if any.
Expected life of the tube, as dependent upon utilization and local conditions, to be specified by the contractor (2).
Advanced maintenance tasks required should be documented.
The contractor should provide detailed annual costs related to up to 5 additional years of preventive and corrective maintenance services, following the initial year of full warranty, including parts, X-ray tube and labour costs.
Any variation to be indicated in the offer.
Additional services required Installation requirements • The contractor should install the system at the site requested in the contract. • The contractor will supply direct electric connection of the equipment to the three-phase power supply network. The connection will be effectuated directly to the network with thermomagnetic disconnecting switch. • The contractor will connect the equipment to the radiology DICOM network, when available. The contractor should review the site layout room to ensure compliance to the technical requirements before supplying the system. • The contractor should communicate with the beneficiary/user and verify the appropriateness of the designated area for the installation of the system. Furthermore, the contractor should provide timely drawings detailing all electrical, structural or any additional requirement to be performed by the beneficiary for the installation of the system. • The completion date of the site preparation will be communicated by the contracting authority to the contractor in due time to start the execution of the contract activities. • Notwithstanding any authorization given by the contracting authority, the contractor should visit, inspect and ascertain that all necessary conditions are met at the beneficiary/user site before starting any activities. Any comments or suggestions as regards the conditions of the site should be made at least 4 weeks before initiating installation activities. Testing and acceptance • Factory acceptance test (FAT): the system, prior to shipment, should be tested for conformity of the system with the manufacturer’s performance specifications and the minimum requirements specified. The FAT report should be provided by the supplier. • The supplier should perform, when applicable: installation, calibration, safety and operation checks before handover. • On-site acceptance test (SAT): the system, after delivery, shall be tested by the contractor together with the user to demonstrate that the performance meets the manufacturer’s performance specifications and the minimum requirements specified as determined by WHO, IAEA and users. The results of the SAT shall be documented in an acceptance test protocol that shall be signed by the end user (after consultation with the hospital medical physicist) and the manufacturer. • Recommended constancy checks provided. • The results of the testing of the system should be documented by the contractor in an acceptance protocol that should be signed and dated by the user. Training and manuals • User manuals, hard and soft copies, to be supplied in English (provision of versions in other UN languages, if available, will be an asset; manuals shall be provided in a language understandable to the user). The supplier has to describe any materials contained in the device that are classified as hazardous under local regulations. • Technical and maintenance manuals, hard and soft copies, to be supplied in English (provision of versions in other UN languages, if available, will be an asset; manuals shall be provided in a language understandable to the user). Supplier has to describe any materials contained in the device that are classified as hazardous under local regulations. • Specific cleaning and disinfection instructions for IPC should be also included. • Training of users in operation should include at least the following topics (a detailed training programme should be provided): – CT technology; – description of all settings, parameters; – optimization procedures of protocols on a patient selective basis; – dose considerations for different patient sizes, in particular for paediatric patients; – image quality and techniques to be used for different clinical indications; – steps on how to adapt to different noise textures; – CT dose quantities; – relevance of CT diagnostic reference levels (DRLS); – warning and reference levels in CT; – CT dose metrics and patient dose estimation; – quality assurance and quality control tests; – radiation safety issues; – importantly, a specific training on the “dose optimization methods” available should be provided. • Training to technical and specifically appointed staff for basic maintenance should be provided (provided also in an online format, if available). • The training provided for user will last not less than 40 hours. The training sessions will be organized for not fewer than four persons/staff. A follow-up training phase should be scheduled. • The training for maintenance staff will last not less than 30 hours. The training session will be organized for not fewer than four persons/staff. A follow-up training phase should be scheduled. • Contact details of manufacturer, supplier and lists of local service agents to be provided together with the documentation. Maintenance requirements • The contractor should include full maintenance services during the standard warranty period. Full maintenance services during the warranty period should include: – preventive maintenance; – on-call interventions; – any software update/upgrade for the system that will be become available; – all necessary replacement and spare parts. • As part of the on-site acceptance, the contractor should provide to the local engineer and to the hospital medical physicist a plan for preventative maintenance and the name and contacts of a service representative/office for on-call maintenance intervention. • Intervention time should be clearly defined and should comply with the uptime requirements defined as follows: – The contractor should guarantee that the CT scanner will have an uptime of at least 95% (excluding outages for maintenance or causes external to the system). – Uptime is calculated on a basis of 250 operating days per year (weekly working days). – Should the downtime exceed 2 working days cumulative on a 6-month basis (i.e. summing up the hours), then the warranty and/or maintenance (as applicable) will be extended for a corresponding period. – The records of downtime of the CT scanner will be kept by a representative of the beneficiaries/ users; the contractor should have the right to request copies of such records. • The contractor should provide an offer up to 5 additional years of preventive and corrective maintenance services as part of the main contract, following the initial year of full warranty, including parts, X-ray tube and labour costs. The related annual costs should be stated in the offer. • The contractor shall ensure that a suitable qualified person can be on-site within 48 hours following an unexpected breakdown and to resolve any problem within 5 working days, throughout the warranty period. • The contractor should preferably include, when available, in the maintenance proposal (separating the periods during warranty and post-warranty) the feature related to remote diagnosis of the equipment. The costs of this option should be separately detailed and fully described in the offer.