## Foglio1
Pulse oximeter, handheld
Published on: 19 November 2020 - Version: 4
Version 3: September 2019 (WHO-UNICEF)
Version 2: 18/06/2014 (modification of the initial version)
Version 1: 13/06/2012 (initial version)
Technical Specification | Minimum requirements
General technical requirements | SpO2 and pulse rate monitor, for adults, children, for all skin pigmentations. Plethysmography waveform (preferable). SpO2 detection to include the range: 70–100%. SpO2 resolution: 1% or less. SpO2 accuracy (in the range at least 70–100%): within ± 2% under ideal conditions of use, and within ± 3% for all patients and perfusion/movement conditions. If equipment is capable of a wider SpO2 detection range, the accuracy over that wider range shall be stated. Pulse rate detection to include the range: 30–240 bpm. Pulse rate resolution: 1 bpm or less. Pulse rate accuracy within ± 3 bpm. Data update period for valid data displayed ≤ 10 s. Display with main parameters: % SpO2, pulse rate, plethysmography waveform (preferable) (and possibly other indicators of signal quality), alarm messages, battery state indication. Suitable for detection in low perfusion conditions (as per ISO 80601-2-61, test method must be described). Automatic correction for movement and ambient light artefacts (as per ISO 80601-2-61, test method must be described). Design must enable use in demanding environments, e.g. shock, vibration and free fall tests as per tests in ISO 80601-2-61. Capable of working with adult, paediatric and neonatal reusable probes. Enclosure to have ingress protection level IPX2 or better. Overall device and probe weight < 400 g. Any aspects of usability as per IEC 62366-1 must be described. Suitable for cleaning and disinfection. Continuous operation within specification in ambient temperature of at least 5–40 °C, relative humidity of at least 10–85% non-condensing (90% preferable). Function for continuous monitoring. Internal data storage for patient trends and event log (preferable).
Displayed parameters | Display with main parameters: % SpO2, pulse rate, plethysmography waveform (preferable) (and possibly other indicators of signal quality), alarm messages, battery state indication. Display must allow easy viewing in all ambient light levels.
User adjustable settings | Audio-visual adjustable alarms: high/low SpO2 and high/low pulse rate (operator variable settings). Alarm override and temporary silencing.
Alarms | Audible and visual alarms for low/high saturation and pulse rate, threshold set by user. Audible and visual alarms for sensor error or disconnected, system errors, low battery. Alarm temporary silencing function.
Accessories | Carry case: 1 per equipment. Reusable probes, adult (finger clip): 2 per equipment. Reusable probes, paediatric: 2 per equipment. Extender cable to achieve probe cable length > 1 m (if applicable): 3 per equipment. Battery charger (if applicable): 1 per equipment.
Spare parts | 1-year spare parts kit as per preventive maintenance programme.
Portability | Portable handheld.
Power supply, voltage, frequency and plug vary across the countries | Operated by replaceable battery power supply, either rechargeable or single use. Rechargeable batteries are preferred. External or built-in AC battery charger, if rechargeable type. Power connection requirements as per local power supply. Charger, if used, to have protection against over-voltage and over-current line conditions, and preferably to be certified to IEC 60601-1. Protection against defibrillator discharges and electrosurgical units. Automatic switch between battery and mains powered modes, when recharging or in mains power failure (preferable). The display shall show which power source is in use. Running time on battery only ≥ 12 hours. Operates from AC power electric line: 100–240 V~/50–60 Hz. Main power cable and plug adapted for various countries.
Documentation requirements (English language mandatory) | Set: user and maintenance manuals, hard and soft copies, in English (mandatory) and other languages (preferable). Certificate of calibration and inspection (other means of assurance may be considered). Troubleshooting, calibration and routine maintenance. List of all spare parts and accessories, with part numbers and contact details for parts supply. Document with contact details of manufacturer, supplier and local service agent.
Primary packaging label | Name and/or trademark of the manufacturer. Electrical power input requirements (voltage, frequency and socket type) and safety use and storage (keep away from oil, grease and petroleum-based or flammable products as well as smoking or open flames). Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: IEC 60601-1 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-1 Medical electrical equipment – Part 1-1: General requirements for safety – Collateral standard: Safety requirements for medical electrical systems. IEC 60601-1-2 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. ISO 80601-2-61 Medical electrical equipment – Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment. ISO/IEEE 11073-10404 Health informatics – Personal health device communication – Part 10404: Device specialization – Pulse oximeter.
Warranty | 2 years with regards efficiency and quality of the product. Availability of accessories and spare parts for at least 2 years.