## Foglio1
Infusion pump
Published on: 19 November 2020 - Version: 1
Alternative naming: infusion multi-therapy pumps; injection pumps
Technical Specification | Minimum requirements
General technical requirements | One channel (at least). Capable of accepting any kind of fluids (solutions and medications preferable). Pump capabilities: • flow range 0.1 to ≥ 999 mL/hr • increments 0.1–100 mL/hr • keep vein open (KVO) rate 1–5 mL/hr • volume to be infused selector (VTBI) 1–9999 mL • flow rate accuracy of ± 5% or better • when multiple channel automatic piggybacking. Ingress protection not less than IPX2. Front panel lockout. Self-check carried out on powering on. Events stored system – log book. Pause infusion facility required. Anti-bolus system to reduce pressure on sudden release of occlusion. IV set: • free-flow protection • air trapping capability • needleless IV connection. “Dose error” reduction system (preferable). Drug library software available, including updates (free during warranty). Air bubble detector with single and cumulative functions (preferable). Clearly visible optical alarms. Acoustic alarms not less than 45dB. Easy set up and operation. Large easy to read display. Real-time display. Availability of a nurse call system connectable to a staff alerting system, 24 V/0.2 A; static or dynamic (preferable). Compatibility with standard infusion sets commonly distributed in the market (desirable at least by the leading brands). Equipment provided “ready to use”. Continuous operation within specification in ambient temperature of at least 5–40 °C, relative humidity of at least 10–90% non-condensing. A limit of 85% relative humidity will be possibly considered as “acceptable” in the current COVID-19 pandemic context. Any accessory or dedicated device necessary to the proper functioning and utilization of the equipment is included. Shock-fall protection.
Monitored and displayed parameters (colour and graphic preferable) | Flow. Pressure. Dose. Availability of software to monitor the delivery of drugs (preferable).
Alarms | Audible alarm required with volume control. Momentary silence less than 2 min. Occlusion upstream. Occlusion downstream. Air in line. System malfunction. Set loaded improperly. Door open. Infusion complete. Loss of mains power. Low battery. Depleted battery (preferable). Clinical advisory messages.
Consumables, labelled “single use” (included and mentioned in a disaggregated list) | Compatible administration sets: 100. Compatible administration sets microbore and macrobore: 100. Compatible long administration sets both microbore/small and macrobore or long extension sets: 100.
Accessories, reusable (included and mentioned in a disaggregated list) | Clamp for mounting pump on IV stand. Clamp for external transportation (preferable) (if applicable).
Spare parts (included and mentioned in a disaggregated list) | As per manufacturer. Include calibration software and hardware (kit). Include list of spare parts with their part numbers and costs.
Portability | Data port required, at least RS232 and/or USB interface. Log analysis software and updates provided for free during warranty period. Wireless connectivity (preferable). Event log required and recording, 1 week at least of operations. Software upgrade required. Software to diagnose and calibrate the equipment with access to calibration settings.
Power supply (voltage, frequency and plug vary across the countries) | Operates from AC mains power: 100–240 V~/50–60 Hz. In-built rechargeable battery. Battery with operating time at least 4 hours at 25 mL/hr. Automatic switch from AC mains power mode to battery operating mode and vice versa. Total re-charging time not greater than 6 hours. Appropriate external device to protect the equipment against over-voltage and over-current line conditions (between plug and socket). 12 V DC socket for recharging during outside transportation (preferable). Equipment must be connected to a reliable and continuous source of energy.
Documentation (included) | Instruction for use; service manual and product information to be provided in English, at least. Certification of calibration. List of procedures for calibration and routine maintenance.
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) and certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]) National local regulatory approval (of recipient country, as applicable).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: Standards specific for infusion equipment: ISO 8536-8:2004 Infusion equipment for medical use – Part 8: Infusion equipment for use with pressure infusion apparatus. ISO 8536-9:2004 Infusion equipment for medical use – Part 9: Fluid lines for use with pressure infusion equipment. ISO 8536-10:2004 Infusion equipment for medical use – Part 10: Accessories for fluid lines for use with pressure infusion equipment. ISO 8536-11:2004 Infusion equipment for medical use – Part 11: Infusion filters for use with pressure infusion equipment. ISO 8536-12:2007 Infusion equipment for medical use – Part 12: Check valves. IEC 60601-2-24 Ed. 2.0:2012 (b) Medical electrical equipment – Part 2-24: Particular requirements for the basic safety and essential performance of infusion pumps and controllers. Other general standards for medical equipment: IEC 60601-1:2012 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-1:2000 Medical electrical equipment – Part 1-1: General requirements for safety – Collateral standard: Safety requirements for medical electrical systems. IEC 60601-1-2:2007 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests.
Warranty | Minimum 2 years. Availability of accessories, consumables and spare parts for at least 5 years.