## Foglio1
Non-wearable urine collection bag, non-sterile/sterile (The technical specification detailed below is general and for an urine collection bag non-sterile)
Published on: 19 November 2020 - Version: 1
WHO Item name: | Required technical specifications: | Dimensions / Sizes / Presentation: | Material: | Certificates for product safety and quality assurance, as requested or equivalent
Bag, collecting, urine, with outlet tap, with non-return valve, 2000 mL, adult, non-sterile, single use | Used to fit gastric tube (aspirating/feeding tube) or incontinence condom. Drainable bag for collecting urine with permanent and legible graduations every 100 mL. With reinforced eyelets for hanging. With a drainage valve (outlet tap) which permits the bag to be emptied without disconnecting, maintaining sterility. It is fitted to an outlet tube. With a non-return valve, located inside the urine collector bag at the upper part (urine entry point), which prevents urine backflow into the indwelling urinary catheter. It is fitted with a kink resistant | Capacity, bag: 2000 mL. Length, tube: 85–95 cm. Diameter, tube: 6.5 mm, approximately | Bag: polyvinyl chloride (PVC), polypropylene or ethylene vinyl acetate (EVA); medical grade. Tube, connector and protective cap: polyvinyl chloride (PVC); medical grade. | ISO 8669-2:1996: Urine collection bags – Part 2: Requirements and test methods.