## Foglio1
Ventilator for intensive care unit
Published on: 19 November 2020 - Version: 1
Technical Specification | Minimum requirements
General technical requirements | Medical oxygen and air high-pressure input ports (> 35 psi [2.4bar]) provide a means to limit reverse gas flowrate (leakage). Each high-pressure input port with a filter and water trap, if applicable, for air input port. Medical air compressor or turbine in-built preferred, alternatively external air compressor. Possibility for using external low-pressure oxygen (approx. 20 psi), as source (preferable). Mechanical safety valve. Internal function testing/leak testing. At least IP21 degree of protection to the harmful ingress of water (fluid spill resistance). Mechanical shock resistance, mechanical vibration, electromagnetic compatibility and electrical safety testing. Compatible active humidifying system. Event log for errors traceability (preferable). Operating temperature and humidity 5–40 °C and 0–95% relative humidity (RH). Storage temperature and humidity -20–60 °C, 0–95% RH.
Ventilation modes | Pressure control ventilation (PCV). Volume control ventilation (VCV). Pressure support ventilation (PSV). Synchronized intermittent mandatory ventilation (SIMV) (preferable). Pressure regulated volume control (PRVC) or similar (preferable). Non-invasive ventilation (CPAP or BiPAP).
Monitored and controlled parameters (by user) | FiO2: 21–100%. Tidal volume: 20–1500 mL. Pressure setting: 0–40 cmH2O. I:E ratio. I:E inverse ratio. RR: 10–60 breaths/min, minimum. Inspiratory pause manoeuvre capability to measure plateau pressure. Adjustable peak pressure limitation/pressure-cycling mechanism above measured peak pressure. Internal PEEP capability/range: 0–20 cmH2O, minimum.
Displayed parameters (colour and graphics preferable) | Display with brightness and contrast control capabilities and/or any additional and equivalent technological solution to improve the readability in low ambient light and in sunlilght. 3 scalar waveforms: pressure, volume and flow (preferable). Loop (axis) displays: pressure-volume, flow-volume and pressure-flow (preferable). Status indicators for ventilator mode, battery status, patient data, alarm settings. FiO2. Airway pressures (peak, plateau mean and PEEP). Tidal volume (expired and inspired preferable). Minute volume (inspired and expired). I:E ratio. RR (spontaneous and mechanical). End-tidal CO2.
Alarms, related to gas delivered (adjustable, visual and audible) | High/low FiO2. High/low inspiratory pressure and PEEP. High/low tidal volume (not achieved or exceeded). Apnoea. High/low RR. Continuously high pressure/occlusion. Breathing circuit disconnect. Low minute volume.
Alarms, related to equipment operation (visual and audible) | Gas supply failure. Power failure. Low battery. Self-diagnostics failure alarm.
Consumables, labelled “single use” (included and mentioned in a disaggregated list) | Breathing circuits for adult: double-limb with standard outlet/inlet connectors with 22 mm outside diameter (at least 60). Breathing circuits for paediatric (if needed): double-limb with standard outlet/inlet connectors with 22 mm outside diameter (at least 40). Exhaled gas filter (when applicable). Heat moisture exchanger filters (HMEF) (at least 90). Non-invasive ventilation masks, for adult (at least 50 in different sizes). Non-invasive ventilation masks, for paediatric (if needed) (at least 12). Viral/bacterial filters 99.99% efficiency minimum; inspiratory and expiratory, as applicable (at least 180).
Reusable accessories (included and mentioned in a disaggregated list) | Breathing circuits for adult: double-limb with standard outlet/inlet connectors with 22 mm outside diameter (at least 2). Breathing circuits for paediatric (if needed): double-limb with standard outlet/inlet connectors with 22 mm outside diameter (at least 2). Expiratory housing with in-built bacteria filters; as well as the possibility to accommodate heat moisture exchangers (HMEs). Exhalation valve. CO2 sensors. Active humidifier with relevant connectors. Air compressor if external to the unit. Standard hoses and connectors (i.e. DISS/NIST as applicable for the country) for oxygen and medical air wall outlets and cylinders. Pressure regulators (from wall outlet to ventilator) to avoid damaging ventilator, as required to operate.
Spare parts (included and mentioned in a disaggregated list) | 1 year’s spare parts kit as per preventive maintenance programme (preferable).
Portability | Mounting tray and support stand (cart for transport with at least 2 castors fitted with breaks).
Power supply (voltage, frequency and plug vary across the countries) | Operates from AC power electric line: 100–240 V AC ± 10% / 50–60 Hz ± 10% of nominal value. In-built rechargeable battery. Automatic switch from AC power electric-line mode to battery operating mode and vice versa. Continuous in battery operating mode with standard ventilation not less than 1 hour. Total re-charging time not greater than 6 hours. Equipment must be connected to a reliable and continuous source of energy.
Documentation (included) | Instruction for use; service manual and product information to be provided in English language, at least.
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for (latest version recommended but compliance to previous standards versions could be accepted): IEC 60601-1:2005 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 60601-1-2:2014 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests. ISO 80601-2-12:2020 Medical electrical equipment – Part 2-12: Particular requirements for basic safety and essential performance of critical care ventilators. ISO 80601-2-74:2017 Medical electrical equipment – Part 2-74: Particular requirements for basic safety and essential performance of respiratory humidifying equipment. ISO 18562-1:2017 Biocompatibility evaluation of breathing gas pathways in healthcare applications – Part 1: Evaluation and testing within a risk management process (if applicable). ISO 20789:2018 Anaesthetic and respiratory equipment – Passive humidifiers (if applicable).
Warranty | Minimum 2 years. Availability of accessories, consumables and spare parts for at least 2 years.