## Foglio1
Resuscitation bag (the technical specification below refers to a "Compressible self-refilling ventilation bag for adult with masks (small, medium, large)")
Published on: 19 November 2020 - Version: 3
Version 2: 16/09/2013 (modification of the initial version)
Version 1: 25/06/2013 (initial version)
Technical Specification | Minimum requirements
General technical requirements | Hand-operated resuscitator used for mechanical ventilation of adult and paediatric patients. Easy to disassemble and reassemble. Easy to clean and disinfect. Reusable. All parts must be manufactured from high-strength, long-life materials and require no special maintenance or storage conditions. Ventilation can be done with ambient air or with oxygen. Resuscitator shall be supplied as a complete set with: • Compressible self-refilling ventilation bag with maximum capacity not less than 1300 mL. Dead volume: < 7 mL. • One-way valve with or without pressure limiting. Pressure-limiting system: the elasticity of the bag’s outer cover limits the airway pressure to approximately 7 kPa (70 cmH2O) when squeezed normally with one hand. • Patient connector, outside diameter: 22 mm; inside diameter: 15 mm. • Inlet valve with nipple for oxygen tubing. • Oxygen reservoir bag capacity according to patient size.
Sizes | Adult: for small adult, adult standard, and large adult. Infant: for small infant, infant standard, and large infant.
Material | Compressible self-refilling ventilation bag: silicone rubber, or other materials specified in ISO10651-4 or equivalent. One-way valve: polycarbonate; polysulfide; silicone, or any other material fulfilling ISO 10651-4 or equivalent. Inlet valve: polycarbonate; polysulfide, or any other material fulfilling ISO 10651-4 or equivalent. Oxygen reservoir bag: bag is made of silicone and valve of polycarbonate/polysulfone or any materials fulfilling ISO10651-4 or equivalent. Oxygen masks: silicone rubber, transparent or equivalent. Materials must be compatible with steam sterilization.
Primary packaging label | Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity, etc.) as appropriate (or equivalent harmonized symbol). If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging. Lot number prefixed by the word “LOT” (or equivalent harmonized symbol). Information for handling, if applicable (or equivalent harmonized symbol).
Documentation requirements (English language mandatory) | Set: user manuals, hard and soft copies, in English (mandatory) and other languages (preferable).
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485). General quality management (e.g. ISO 9001). Application of risk management to medical devices (e.g. ISO 14971).
Regulatory approval/ certification | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | Compliance to the following international standards or to regional or national equivalent (including the technical tests for safety and performance from accredited laboratory or third party) for: ISO 10651-4:2002* Lung ventilators – Part 4: Particular requirements for operator- powered resuscitators (*EN 13544-2 implied), oxygen related clauses are optional for face mask (if not made of silicone). ISO 10993-1:2009 Biological evaluation of medical devices – Part 1: Evaluation and testing for mask. ISO 10993-5:2009 Biological evaluation of medical devices – Part 5: Tests for in vitro cytotoxicity. ISO 10993-10:2010 Biological evaluation of medical devices – Part 10: Tests for irritation and delayed- type hypersensitivity (or classified as USP class V).
Warranty | 1 years with regards efficiency and quality of the product.