## Foglio1
Sterilizer steam autoclave
i | Version No. | 2
ii | Date of initial version | 2012-08-13
iii | Date of last modification | 2014-06-18
iv | Date of publication
v | Completed / submitted by | WHO working group
Name, category and coding
1 | WHO Category / Code | (under development)
2 | Generic name | Steam sterilizer
3 | Specific type or variation (optional)
4 | GMDN name | Unwrapped-device steam sterilizer
5 | GMDN code | 40547
6 | GMDN category | 03 Dental devices 05 Hospital hardware 15 Healthcare facility products and adaptations
7 | UMDNS name | Sterilizing Units, Steam, Tabletop
8 | UMDNS code | 16142
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Autoclave; High pressure steam sterilizer
11 | Alternative code/s (optional) | MS 42162; MS 34639
12 | Keywords (optional) | benchtop, autoclave
13 | GMDN/UMDNS definition (optional) | A mains electricity (AC-powered) device designed for total elimination and/or inactivation of microorganisms from medical devices and related products, not placed in sterilization wraps/packaging, using pressurized steam (i.e., moist heat) as the sterilizing agent; it is used for products non-sensitive to high temperature, water, or steam. It typically includes a treatment chamber with shelves for device placement, usually after cleaning of gross debris; a means to introduce the steam into the chamber; and controls to regulate the time and/or temperature of the procedure. The device is available in a variety of shapes and sizes, including stand-alone (bulk) and tabletop units.
Purpose of use
14 | Clinical or other purpose | Total elimination and/or inactivation of microorganisms from medical devices and related products, not placed in sterilization wraps/packaging
15 | Level of use (if relevant) | health centre, district hospital, provincial hospital, specialized hospital
16 | Clinical department/ ward(if relevant) | operating theater, central sterilization, clinical laboratory
17 | Overview of functional requirements | Uses pressurized steam to kill microorganisms on medical devices and products Allows the user to control time and temperature of procedure Stands on table top with easy access to chamber Generates heat using integral electric heater
Technical characteristics
18 | Detailed requirements | Microprocessor control Overheat shutoff and overpressure safety valve to be incorporated Vacuum air removal facility is not required, gravity removal valve is sufficient Pressure lock to be incorporated to prevent door opening at pressure Temperature range to include at least 100 to 132 deg C. Chamber capacity to be at least 20 liters. Required water level to be clearly indicated. External surfaces to remain at safe temperatures even when in use Internal steam electrical generator. At least the following cycles available: • Solids. • Glassware materials. • Liquids. • Vacuum test. • Bovie-Dick test. Adjustable temperature working range not smaller than from 115 °C up to at least 121 °C. Temperature measure precision not greater than +/- 3%. Vacuum pump and vacuum sustainability diagnostic system. Safety systems, at least: Thermostat, Pressure switch, Valves. Protection system for high pressure risks. Filters for air intake system. Automatic block in high and low pressure conditions.
19 | Displayed parameters | Temperature. Pressure. Working Time. Equipment status and alarms. Alarms for at least: • Power failure. • Low water. • Door not closed. • Pressure and/or Temperature out-limits. • Sterilization cycle failure. • End of sterilization cycle.
20 | User adjustable settings | User-resettable time elapsed indicator to be incorporated.
Physical/chemical characteristics
21 | Components(if relevant) | Desktop mounted unit, standing stable in all door positions Supplied with internal trays that are removed and replaced easily, of perforated or meshed construction Chamber drain to be secure in operation at pressure but easy to open after use Chamber door to include gasket and closure handles that are easy to operate Supplied with cover for protection from spray and dust
22 | Mobility, portability(if relevant) | Fixed, non mobile
23 | Raw Materials(if relevant) | All metal parts to be constructed of stainless steel
Utility requirements
24 | Electrical, water and/or gas supply (if relevant) | 1) Electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. 2) Electrical source with line connection plug ______ type. 3) Protections against over-voltage and over-current line conditions. 4) Compliance with ______ electrical standards and regulations. Voltage corrector / stabilizer and UPS to allow operation at ± 30% of local rated Electrical protection provided by fuses in both live and neutral supply lines
Accessories, consumables, spare parts, other components
25 | Accessories (if relevant) | Mains cable to be at least 3m in length Stainless steel stand designed to support the autoclave. At least n. 2 system compatible baskets for different sterilization applications.
26 | Sterilization process for accessories (if relevant)
27 | Consumables / reagents (if relevant)
28 | Spare parts (if relevant) | Two sets of spare fuses (if non-resettable fuses used) Replacement door gasket to be supplied. Replacement heating element to be supplied
29 | Other components (if relevant) | 1 air intake spare filter. 2 sets of compatible printing paper (if integrated printer provided). 2 complete Bovie-Dick test kits for daily monitoring. 4 testing complete sets for standard cycles control/tests.
Packaging
30 | Sterility status on delivery (if relevant) | N/A
31 | Shelf life (if relevant) | N/A
32 | Transportation and storage (if relevant) | N/A
33 | Labelling (if relevant) | N/A
Environmental requirements
34 | Context-dependent requirements | The case is to be cleanable with alcohol or chlorine wipes. Interior chamber to be accessible for easy cleaning
Training, installation and utilization
35 | Pre-installation requirements(if relevant) | Supplier to perform installation, safety and operation checks before handover
36 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided
38 | User care(if relevant) | The case is to be cleanable with alcohol or chlorine wipes. Interior chamber to be accessible for easy cleaning
Warranty and maintenance
39 | Warranty
40 | Maintenance tasks
41 | Type of service contract
42 | Spare parts availability post-warranty
43 | Software / Hardware upgrade availability
Documentation
44 | Documentation requirements | User, technical and maintenance manuals to be supplied in ************** language. Certificate of inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost.
Decommissioning
45 | Estimated Life Span | 5 to 15 years
Safety and standards
46 | Risk Classification | Class B (GHTF Rule 15);Class II (USA); Class II (EU, Japan, Canada and Australia)
47 | Regulatory Approval / Certification
48 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes (Australia, Canada and EU) ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests IEC 61010-2-040:2005 (Safety requirements for electrical equipment for measurement, control and laboratory use - Part 2-040: Particular requirements for sterilizers and washer-disinfectors used to treat medical materials)
49 | Reginal / Local Standards | AAMI/ANSI ST8:2008 Hospital steam sterilizers AAMI/ANSI ST55:2010 Table-top steam sterilizers JIS K 3605:1992 General rules for autoclaving JIS T 7322:2005 High-pressure steam sterilizers for medical use JIS T 7324:2005 High-pressure steam sterilizers for medical use (Small size)
50 | Regulations | US regulations 21 CFR part 820 21CFR section 880.6880 sterilizer, steam CDC, Guideline for Disinfection and Sterilization in Healthcare Facilities, 2008 JP regulations MHLW Ordinance No.169 38671010 Wrapped-device steam sterilizer 40547010 Unwrapped-device steam sterilizer