## Content
Technical specifications for selection of essential in vitro diagnostics (WHO-EDL)
This document presents technical specifications for one in vitro diagnostic test listed in the 3rd WHO-EDL (2021)
Each technical specification is presented in a standard template organized as follows: name of the IVD test or laboratory equipment and then 88 technical specifications organized into 13 categories: Name, category and coding; intended use; performance characteristics; technical and operational characteristics; instrument physical and technical characteristics; infrastructure required; accessories, consumables, spare parts and other components; documentation; environmental and safety requirements; training, installation and use; warranty and maintenance; decommissioning; quality and registration.
To review the technical specifications for a specific IVD category, go to the test menu tab and click on the name of the test category. The selected item will be displayed on the screen. Scroll down to review the generic technical characteristics.

## Test menu
Technical specifications for selection of essential in vitro diagnostics: WHO-EDL 3
No. | Page in TRS 1031/EDL 3 |  |  | Diagnostic test name | Test purpose | Assay format
1 | P288 for disease-specific IVDs for use in community settings and health facilities without laboratories | 51707/ Beta-haemolytic Group A Streptococcus antigen IVD, kit, rapid ICT, clinical | W0105090103/ STREP. A - RT & POC | Group A Streptococcus antigen | To aid in the diagnosis of Group A streptococcal pharyngitis | Rapid diagnostic test (RDT)

## GroupAStAgRDTnolab
Group A Streptococcus antigen
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51707/ Beta-haemolytic Group A Streptococcus antigen IVD, kit, rapid ICT, clinical
3 | EMDN code/ Term name | W0105090103/ STREP. A - RT & POC
4 | Generic name | Streptococcus Group A antigen
5 | Specific type or variation (optional) | Rapid chromatographic immunoassay for qualitative detection of Streptococcus A antigens
6 | Alternative name/s (optional) | Streptococcus Group A antigen rapid diagnostic test
7 | Keywords (optional) | Streptococcal pharyngitis, Streptococcus pyogenes Group A antigen RDT, Streptococcus pyogenes Group A antigen rapid diagnostic test
Intended use
8 | Detection target | Group A streptococcus antigen
9 | Test purpose | To aid in the diagnosis of Group A streptococcal pharyngitis
10 | Specific disorder / condition | Streptococcal pharyngitis
11 | Test format | Rapid diagnostic test (RDT)
12 | Specimen type | Throat swab
13 | Testing population | Adults and children (> 3 years)
14 | Intended users | Trained health-care providers, trained lay and peer workers
15 | Level of health-care system | For use in community settings and health facilities without laboratories
Performance characteristics
16 | Clinical sensitivity | In comparison with swab culture: ≥ 91%
17 | Clinical specificity | In comparison with swab culture: ≥ 97%
18 | Analytical specificity | Assay does not cross-react with other human infections relevant to the local setting or common interfering substances. If any cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | 1.0 x 10 000 CFU per test
21 | Rate of invalidity, error or unreturnable | < 5%
22 | Precision (repeatability/reproducibility) | Percentage agreement: ≥ 98%
23 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
24 | Principle of the assay | Lateral flow immunoassay
25 | Specimen(s) stability | Testing should be performed as soon as possible after specimen collection. Specimens can be stored at 2–8 °C for up to 24 h or at 15–30 °C for up to 4 h in a dry, sterile, tightly sealed plastic tube.
26 | Specimen(s) volume | Single swab
27 | Type of result | Qualitative
28 | Time to result | ≤ 12 min
29 | End-point stability | 0–10 min
30 | Ease of use (for point-of-care tests only): number of steps that require precision | ≤ 6 steps
31 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 5 specimens per hour per operator
32 | Test limitations | The manufacturer should state any conditions that have not been validated with the test, such as testing population, sample types and users. This assay does not distinguish between viable and nonviable organisms and may produce a positive result in the absence of living organisms. The assay does not differentiate asymptomatic carriers of Group A Streptococcus from those with streptococcal infection. Additional follow-up testing with the culture method is required if the result is negative and clinical symptoms persist or in the event of an outbreakof acute rheumatic fever (ARF) .
33 | Internal quality control | Built-in quality control line that must appear when specimen is added plays a role of Internal quality control. Absence of control line makes the test invalid. A recommendation for re-testing should be provided in the procedure in case of invalid results. Also, a description should be included in the package insert regarding how positive, negative, or invalid results are read and determined.
34 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme. Proficiency of each user is best confirmed by their participation in an external quality proficiency programme. Please note that internal and external controls should be designed to demonstrate that the device is operating as intended with the same thresholds for detectability of an analyte. The concentration of any positive control should be at the lower end of 100% detectability.
35 | Transport stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected against humidity and direct sunlight
36 | Storage stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected against humidity and direct sunlight
37 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
38 | In-use stability of kit/reagents (temperature and humidity) | At 15–30 °C as a required temperature for testing
39 | Transport stability of specimen (temperature and humidity) | Specimens can be stored at 2 - 8°C up to 24 h in a dry, sterile and tightly sealed plastic tube
40 | Storage stability of specimen (temperature and humidity) | Specimens can be stored at 2 - 8°C up to 24 h in a dry, sterile and tightly sealed plastic tube or at 15 - 30°C up to 4 h in a dry, sterile and tightly sealed plastic tube
41 | In-use stability of specimen (temperature and humidity) | Specimens can be stored at 2 - 8°C up to 24 h in a dry, sterile and tightly sealed plastic tube or at 15 - 30°C up to 4 h in a dry, sterile and tightly sealed plastic tube
42 | Transport stability of controls/calibrators (temperature and humidity) | At 2–30 °C
43 | Storage stability of controls/calibrators (temperature and humidity) | At 2–30 °C
44 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
45 | In-use stability of controls/calibrators (temperature and humidity) | At 15–30 °C as a required temperature for testing
46 | Shelf-life of kit/reagents at time of manufacture (months) | 18 months from date of manufacture
47 | Remaining shelf-life of kit/reagents at delivery (months) | At least 6 months shelf-life remaining when product arrives at point of use Note: Please see page 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
48 | Size of device (height x width x depth) | Not applicable
49 | Weight of the device (kg) | Not applicable
50 | Power requirements | Not applicable
51 | Time for battery to charge | Not applicable
52 | Battery duration | Not applicable
53 | Alternative charging options | Not applicable
54 | Operating conditions (temperature and humidity) | Not applicable
55 | User interface | Not applicable
56 | Display parameters | Not applicable
57 | Display languages | Not applicable
58 | Built-in memory storage capacity | Not applicable
59 | Diagnostic connectivity | Not applicable
60 | Open or closed system | Not applicable
61 | Multi-disease testing capability / Test menu | Not applicable
Infrastructure requirements
62 | Water requirements | Not required
63 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature will require electricity.
64 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits
65 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
66 | Kit component – test / reagents / consumables (if relevant) | Test strips or devices or cassettes or card, positive control, negative control, reagents, swabs, instruction for use or package insert
67 | Reagent kit size (number of tests) | ≥ 10 tests
68 | Consumables required but not provided in the test kit | A permanent marker or pen, biohazard bags, disposable gloves and sodium hypochlorite (bleach) for sterilization before disposal of infectious waste are required. If disposable swabs are not provided, swabs are required.
69 | Other auxiliary laboratory equipment required but not provided | A biohazard container and a timer are required. If several tests are being performed, > 1 timer would be an advantage.
70 | Spare parts (if relevant) | Not applicable
Documentation
71 | Instrument operator manual | Not applicable
72 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
73 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. It should contain the actual results obtained during final quality control for release for each lot of a product. This certificate should be retained by the testing centre with the dates on which the lot number was in use.
74 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environment and safety requirements
75 | Hazard classification | Potential source of harm classified according to the Global Harmonized System, such as health hazard, physical hazard or environmental hazard. Usually stated on the material safety data sheet
76 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g. simple garbage or incineration
Training, installation and use
77 | Pre-installation requirements (if relevant) | Not applicable
78 | Requirements for installation and calibration | Not applicable
79 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
80 | Warranty | Not applicable
81 | Preventive maintenance | Not applicable
82 | Corrective maintenance | Not applicable
83 | Type of service contract (including leasing or reagent rental) | Not applicable
84 | Availability of spare parts post-warranty | Not applicable
82 | Availability of technical support in the country | Availability of a local authorized service provider or distributor is recommended if the product was not procured directly from the manufacturer.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification (PQ) status | This assa, in an RDT format, is not within the scope of WHO Prequalification. Please see WHO Public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback on complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## Acknowledgement and Purpose
Purpose and aim
The WHO Model List of Essential In-vitro Diagnostics (EDL) is a list of recommended types of tests to support diagnosis of infectious diseases and noncommunicable diseases at each level of the health-care system. The EDL serves as an evidence-based reference point for countries to develop their own national lists to guide how they choose and use IVDs. The EDL is also published in an electronic format as eEDL. The EDL do not provide detailed specifications to guide selection and procurement of specific IVDs. Therefore detailed technical specification are required to guide countries in procuring appropriate tests. Technical specifications for IVDs consist of criteria and baseline requirements to ensure good quality, safety, performance and efficacy. Technical specifications for IVDs listed in the EDL will be published in an electronic format. The purpose of technical specifications for IVDs in the EDL is to advance universal health coverage, address health emergencies and promote healthier populations by helping countries to select and procure appropriate IVD products. The technical specifications provide detailed technical information to support decisions on procurement of good-qualit, safe IVDs to improve access to in vitro diagnostic testing in low- and middle-income countries in particular.
Acknowledgements
Technical contributions: Nikhil Moorchung (WHO consultant, India), Emma English (WHO consultant, United Kingdom), Jitendra Kumar (WHO consultant, India), Olga Perovic (WHO consultant, South Africa), Ehab Saad (WHO consultant, South Africa), Francis Gabriel Moussy (WHO headquarters), Ana Aceves Capri (WHO headquarters), Naofumi Hashimoto (WHO headquarters), Adriana Velazquez Berumen (WHO headquarters) and Jorge Matheu (WHO headquarters)
Conflicts of interest and declarations of interest
All the WHO consultants involved in preparing these technical specifications provided declarations of interests, and no conflicts were identified.
Funding
Funding and in-kind contributions to this project were provided by the following organizations: • The Bill & Melinda Gates Foundation • The Foreign, Commonwealth and Development Office, United Kingdom • The Economic Research Institute for the Association of Southeast Asian Nations and East Asia
Updating
The document will be updated regularly, at least every 3 years.

## Process
Process
Technical specifications define the minimum requirements to ensure good quality, safety and efficacy of the products. Development of technical specifications consists of the six steps listed below.
•Step 1: Review existing WHO guidance and specifications that support the procurement, donation or lease of good-quality IVDs to develop a common template of generic specifications. The guidance documents analysed were: Guidance for procurement of in vitro diagnostics and related laboratory items and equipment, Laboratory diagnosis of group A streptococcal infections, WHO Model List of Essential In Vitro Diagnostics and WHO internal procurement templates.
•Step 2: WHO commissions a team of external consultants to develop technical specifications for the IVDs listed in the WHO-EDL. Development of specifications comprises: horizon scanning of existing tests and platforms; analysis of products on the market after approval by stringent regulatory authorities; analysis of international, regional and country standards; review of WHO guidance and guidelines for all the diseases included in the EDL; review of peer-reviewed publications, textbooks and published qualitative and quantitative comparative studies of tests and platforms; comparison of test specifications (i.e. accuracy, precision, positive and negative predictive value, likelihood ratios), comparison and review of the package inserts and user manuals of products approved by stringent regulatory authorities; comparison and review of documents and data on the safety and performance of approved products on the websites of stringent regulatory authorities; results of external quality assessments and technical expert opinion. Each template is drafted or reviewed by at least two qualified technical experts. All consultants provided declarations of Interest, and no conflicts were identified.
•Step 3: Internal review of all draft technical specifications by the EDL secretariat. •Step 4: Review of draft technical specifications by relevant disease programmes. •Step 5: 3-week public consultation on the draft technical specifications by posting them on the WHO website, sharing them on relevant networks and with non-State actors in official relations with WHO, such as the International Federation of Biomedical Laboratory Science. •Step 6: Compile all feedback, and finalize the technical specifications before adding them to the electronic EDL.

## Evidence and Input
Evidence and Input
Sources of evidence:
1. Laboratory diagnosis of group A streptococcal infections. Geneva: World Health Organization; 1996 (https://apps.who.int/iris/bitstream/handle/10https://apps.who.int/iris/bitstream/handle/10665/41879/9241544953_eng.pdf?sequence=1&isAllowed=y). 2. Guidance for procurement of in vitro diagnostics and related laboratory items and equipment. Geneva: World Health Organization; 2017 (https://apps.who.int/iris/handle/10665/255577). 3. Stewart EH, Davis B, Clemans-Taylor BL, Littenberg B, Estrada CA, Centor RM. Rapid antigen group A ttreptococcus test to diagnose pharyngitis: a systematic review and meta-analysis. PLoS One. 2014;9(11):e111727. doi:10.1371/journal.
4. Cohen JF, Bertille N, Cohen R, Chalumeau M. Rapid antigen detection test for group A streptococcus in children with pharyngitis. Cochrane Database Syst Rev. 2016(7): CD010502. doi:10.1002/14651858.CD010502.pub2. 5. Point-of-care tests for group A streptococcus. Horizon scan report. Oxford: Community Health Care MedTech and In Vitro Diagnostics Co-operative, National Institute for Health Research; 2015 (https://www.community.healthcare.mic.nihr.ac.uk/reports-and-resources/horizon-scanning-reports/point-of-care-tests-for-group-a streptococcus).
6. Cadamuro J, Lippi G, von Meyer A, Ibarz M, van Dongen E, Lases5 et al. European survey on preanalytical sample handling – Part 2: Practices of European laboratories on monitoring and processing haemolytic, icteric and lipemic samples. On behalf of the European Federation of Clinical Chemistry and Laboratory Medicine (EFLM) Working Group for the Preanalytical Phase (WG-PRE). Biochem Med (Zagreb). 2019;29(2):020705. doi:10.11613/BM.2019.020705. 7. ISO 15189:2012 Medical laboratories – Requirements for quality and competence. Geneva: International Standards Organization; 2022 (https://www.iso.org/standard/56115.html). 8. Cheng CM, Kuan CM, Chen CF. In-vitro diagnostic devices: Introduction to current point-of-care diagnostic devices. Berlin: Springer International Publishing; 2015 (https://www.springerprofessional.de/en/introduction-to-in-vitro-diagnostic-devices/2468988). 9. Ogrodnik PJ. In vitro diagnostic devices: Class 1. Sterile and measuring devices. Amsterdam: Elsevier Science; 2014. ISBN-10. 0128000252. 10. FDA guidance for industry: sterile drug products produced by aseptic processing – current good manufacturing practice, September 2004. College Park (MD): Department of Health and Human Services, Food and Drug Administration, Center for Drug Evaluation and Research; 2004 (https://www.gmp-compliance.org/guidelines/gmp-guideline/fda-guidance-for-industry-sterile-drug-products-produced-by-aseptic-processing-current-good-manufacturing-practice-september-2004).
11. Pisano DJ, Mantus DS, editors. FDA regulatory affairs: A guide for prescription drugs, medical devices, and biologics. Boca Raton (FL): CRC Press; 2008. doi:10.3109/9781420073553. 12. Fiedler BA, editor. Managing medical devices within a regulatory framework. Amsterdam: Elsevier Science; 2016 (https://www.elsevier.com/books/managing-medical-devices-within-a-regulatory-framework/fiedler/978-0-12-804179-6). 13. Tobin JJ, Walsh G. Medical product regulatory affairs: Pharmaceuticals, diagnostics, medical Devices. Hoboken (NJ): Wiley; 2011 (https://www.wiley.com/en-us/Medical+Product+Regulatory+Affairs:+Pharmaceuticals,+Diagnostics,+Medical+Devices-p-9783527318773). 14. Wong J, Tong RKY, editors. Handbook of medical device regulatory affairs in Asia (second edition). Singapore: Jenny Stanford Publishing; 2018 (https://www.jennystanford.com/9789814774314/handbook-of-medical-device-regulatory-affairs-in-asia-second-edition/). 15. Mantus D, Pisano DJ, editors. FDA Regulatory Affairs, third edition. Boca Raton (FL): CRC Press; 2014 (https://www.routledge.com/FDA-Regulatory-Affairs-Third-Edition/Mantus-Pisano/p/book/9781841849195).