## Simulator childbirth advanced
Simulator, childbirth, torso
i | Version No.
ii | Date of initial version
iii | Date of last modification | October 2023
iv | Date of publication
v | Completed / submitted by | WHO working group
Name, category and coding
1 | WHO Category / Code | Under development
2 | Generic name | Simulator, childbirth, advanced
3 | Specific type or variation (optional) | Advanced childbirth simulator
Purpose of use
4 | Clinical or other purpose | Simulator for obstetric training in normal and complicated deliveries including breech presentations and C-section.
5 | Level of use (if relevant) | General outpatient and outreach services for primary care; First-level (district) hospital services; Second-level and third-level hospital services and specialized outpatient services
6 | Clinical department/ ward(if relevant) | Outpatient care, Inpatient care, Labor and delivery services
7 | Overview of functional requirements | N/A
Technical characteristics
10 | Detailed requirements | Childbirth simulator for training in normal as well as abnormal deliveries including complete, frank, and footling breech. Should consist of a full-size adult female lower torso with a life-size pelvic cavity and major anatomical landmarks. Able to simulate multiple pregnancies in vertex/vertex, vertex/breech, breech/vertex, or breech/breech presentations. To be used to demonstrate placenta praevia and prolapse of the umbilical cord. Exhibit soft vulval inserts for episiotomy exercises with an inflatable uterus insert that allows for the performance of Leopold’s maneuvers. Able to realistically simulate cesarean-section delivery, Ritgen’s maneuver, and intrauterine manipulation. Allows training in normal delivery of umbilical cord and placenta with a full-term fetal baby with fontanelles, umbilical cord, and placental. Anterior and posterior foetal fontanelles can be palpated. Removable diaphragm end plate for manual positioning of fetal baby/babies. Complete unit to be easily washable with mild soap and water without bringing deterioration in the manikin, non-reactive to manufacturer-specified disinfectants/alcohol wipes.
11 | Displayed Alarms | N/A
12 | User adjustable settings | N/A
Physical/chemical characteristics
13 | Components(if relevant) | Material of the simulator: Thermoplastic polyurethane - TPU (primary material for body skin), Polyvinylchloride - PVC, silicone rubber, Polyoxymethylene - POM, Polypropylene - PP.
14 | Mobility, portability(if relevant) | N/A
15 | Raw Materials(if relevant) | Latex-free manikins - ideally made of silicone rubber, if not, made of plastics. Hinges or joints if any should be of noncorrodible and durable metals/plastics or silicone rubber. Lubrication for joints or hinges if any should be of non-damaging substances and safe to use chemicals.
Utility requirements
16 | Electrical, water and/or gas supply (if relevant) | Relevant fluid specifications for inflatable manikins (fluid type, quantity, temperature, etc) Optional: Electrical connections as recommended by the manufacturer of manikins for alarms, simulation devices and software
Accessories, consumables, spare parts, other components
17 | Accessories (if relevant) | Can be ordered individually. To be used together with the Advanced childbirth simulator. Fetal baby, w/umbilical cord & placenta Labor delivery module
18 | Sterilization process for accessories (if relevant) | Easily disinfectable by either rubbing vigorously with alcohol wipes or by use of chemical agents, as recommended by the manufacturer and in compliance with regional/organizational regulatory protocols for manikin disinfection. The product should be cleaned with soap and water prior to any disinfection methods.
19 | Consumables / reagents (if relevant) | Alcohol wipes/disinfectants for cleaning the manikin, lubricants (if any), and puncture repair patches (for inflatable manikins).
20 | Spare parts (if relevant) | Detachable faces, airway systems
21 | Other components (if relevant) | N/A
Packaging
22 | Sterility status on delivery (if relevant) | N/A
23 | Shelf life (if relevant) | Shelf life of at least 5 years on storage at 0˚C to +40˚C, relative humidity of 15 to 95%
24 | Transportation and storage (if relevant) | Capable of being transported/stored continuously in an ambient temperature of -30˚C to +60˚C and relative humidity of 15 to 95%. Capable of operating continuously in an ambient temperature of 0˚C to +40˚C and relative humidity of 15 to 95%
25 | Labelling (if relevant) | Labeling on the primary packaging to include: Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity). Manufacturer’s product reference Manufacture and expiry date, lot number Main characteristics (dimensions, capacity) Recommended storage conditions Information for handling
Environmental requirements
26 | Context-dependent requirements | Capable of being stored continuously in an ambient temperature of at least 0–50 °C and relative humidity of 15–85%, preferably 90%. Capable of operating continuously in an ambient temperature of at least 10–40 °C and relative humidity of 15–85%, preferably 90%.
Training, installation and utilization
27 | Pre-installation requirements(if relevant) | Not required.
28 | Requirements for commissioning (if relevant) | Local clinical staff to affirm completion of installation. Supplier to perform installation, safety and operation checks before handover.
29 | Training of user/s (if relevant) | Source of useful information; Training of users in handling, disinfection, and basic maintenance should be provided: WHO Guidelines on basic newborn resuscitation https://apps.who.int/iris/handle/10665/75157 WHO technical specifications of neonatal resuscitation devices. https://apps.who.int/iris/handle/10665/206540 ASTM F1453 - 92 Standard Guide for Training and Evaluation of First Responders Who Provide Emergency Medical Care (USA). https://www.astm.org/Standards/F1453.htm
30 | User care(if relevant) | User manual for handling, instructions for disinfection and basic maintenance should be provided.
Warranty and maintenance
31 | Warranty | At least 2 years from date of purchase
32 | Maintenance tasks | A list shall be provided of procedures required for routine maintenance. A calibration certificate and calibration plan are mandatory for manikins with electronic sensors and alarms/mechanical systems for feedback.
33 | Type of service contract | N/A
34 | Spare parts availability post-warranty | 5 years from the date of purchase or more as per manufacturers offer. List of spares/ accessories, relevant part numbers, MOQ and price list to be submitted along with the quotation
35 | Software / Hardware upgrade availability | N/A
Documentation
36 | Documentation requirements | The instruction manual should include the Standard Operating Procedure for using the product, day-to-day maintenance, a maintenance schedule, storage condition, cleaning, troubleshooting, safety details, warranty details with a scope, and postal, electronic, and telephone contacts of authorized dealer/country representative as appropriate. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
Decommissioning
37 | Estimated Life Span | At least 5 years
Safety and standards
38 | Standards, for the manufacturer and the equipment | N/A
39 | Regulatory Approval / Certification | Free sales certificate (FSC). Certificate for exportation of medical device provided by the authority in the manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
40 | International standards | In the absence of a product standard for manikins, relevant International standards for testing of raw material and/or the finished product (or other internationally recognized or national equivalent standards) may be used to qualify the manikins. The qualification approach should be as per base guidelines given in section 2.4 of this document. Silicone rubber: The proof of raw material qualification certificate from a nationally or Internationally approved laboratory. Some standards that are suggested are: Plastics: ISO 11403-3: Plastics -- Acquisition and presentation of comparable multipoint data - Part 3: Environmental influences on properties. https://www.iso.org/standard/55462.html Weathering of adhesive joints: ISO 14615: Durability of structural adhesive joints - Exposure to humidity and temperature under load https://www.iso.org/standard/24971.html General compliance: ISO/IEC 17050-1: Conformity assessment. Supplier’s declaration of conformity. General requirements and regional regulatory guidelines should be complied with for waste disposal and recovery management. https://www.iso.org/standard/29373.html