## Foglio1
Voltage stabilizers
i | Version No. | 1
ii | Date of initial version
iii | Date of last modification
iv | Date of publication | September 2019
v | Completed / submitted by | UNICEF / WHO
Name, category and coding
1 | WHO Category / Code
2 | Generic name | Voltage stabilizer.
3 | Specific type or variation (optional) | Electronic, solid state.
4 | GMDN name | Voltage regulator.
5 | GMDN code | 35934
6 | GMDN category
7 | UMDNS name | Power supplies, Line-voltage stabilization
8 | UMDNS code | 16934
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Voltage stabilizer, power conditioner.
11 | Alternative code/s (optional) | N.A.
12 | Keywords (optional) | Voltage regulator, stabilizer, conditioner, power supply, electrical supply.
13 | GMDN/UMDNS definition (optional) | GMDN A device that helps to maintain a stable mains voltage, i.e. with the help, or as part, of a suitable secondary power supply. Should the voltage drop below a pre-set limit the device will increases the voltage in correspondence with the particular type of device it is protecting. Some types may, when there is a sufficient reduction, automatically switch on a boost transformer. UMDNS Power supplies designed to deliver an almost steady, predetermined value of AC voltage (e.g. 120 V) for a limited range of input line voltage variations (e.g. plus or minus 10%). These devices typically include an electronic line voltage monitor to detect line voltage fluctuations and some electric or electronic stabilization circuit or device to compensate for these variations. Line voltage stabilizing power supplies are mostly used to maintain normal operation of electrical equipment (e.g. motors, pumps, lights, refrigerators) under reduced-voltage conditions (brownouts).
Purpose of use
14 | Clinical or other purpose | The voltage regulator is a power supply that is installed between the mains and the electrical medical equipment. It stabilizes the mains voltage to a pre-determined output, eliminating its fluctuations and protecting the electrical equipment also from surges. Various capacity and configurations of voltage regulators can be installed according to the needs. The solid state voltage regulator is fully electronic, without moving parts. Compared with the servo-electronic voltage regulator (which is electro-mechanic with motors and coiled transformers), the solid state version is a smaller and lower maintenance device, and, due to its full electronic design, it has faster reaction. On the other side, the servo-electronic regulator is more accurate and resistant to input voltage and current fluctuations. For equipping normal department, such as NICU, with no special requirements in terms of electrical supply, a solid state voltage regulator is the optimal solution. For high-demanding departments, such as radiology, that needs more accurate voltage stability and improved protection from external events, a servo-electronic voltage regulator may be the best solution. In harsh environments and with low maintenance resources, often found in low-resource settings, the solid state technology is more reliable and offers the best long-term performance and value for money.
15 | Level of use (if relevant) | Health post; health centre; district/general hospital; regional hospital, secondary care; specialized hospital, tertiary care; emergency vehicles; home care.
16 | Clinical department/ ward(if relevant) | All sites where electric equipment, requiring stable mains power supply and sensitive to voltage fluctuations and surges, is installed.
17 | Overview of functional requirements | Electronic voltage regulator, solid state tap-switching technology, microprocessor controlled and without moving parts.Single phase, available for 110–120 VAC and 220–240 VAC networks, 50–60 Hz.Availability of multiple capacities and configurations, suitable for portable use and for fixed installations.Availability of models with built-in surge suppressor, circuit breakers and isolating transformer.Availability of various configurations and connections to fit all electrical standards for networks and equipment.
Technical characteristics
18.1 | Detailed requirements | Electronic, microprocessor controlled voltage stabilizer. Solid state tap-switching technology, without moving parts. Single phase, one or two output socket(s) of local standard design.If output is wired or IEC style, a convertor to local standard socket must be supplied. 110–120 VAC or 220–240 VAC ± 10% nominal input voltage, depending upon local requirement. Frequency range at least 50–60 Hz ± 10%, exact frequency will depend on local requirement. Load rating at least 1 kVA. Voltage input and output indicators.Input resettable circuit breaker for overload or short circuit. (Single-use fuses not acceptable.) Spike, surge, transient and lightning protection. Suitable for short term overload working conditions, at least 135% for 4 minutes and 110% for 6 minutes.Independent and filtered (non-interfering) outlets (applicable to multiple outlet voltage regulators). Automatic over-voltage and under-voltage protection, with output power supply disconnection. Suitable for short-term extended under/over input voltage (at least additional ± 2%) without activation of the protection and with output swing < 10%.Automatic start-up after disconnection, with start-up delay between 2 s and 10 s. Outlet voltage 110 VAC or 220 VAC (as per specified local requirement) ± 8% or narrower accuracy. Protection of the regulator and the load against switch on/off over current and over voltage. Manual bypass function. Visual and/or audible alarm for operational conditions and system status. Efficiency at least 80% at 25% load, 90% at 50% load and 95% at 75–100% load.Correction speed > 750 V/s.Unaffected by load power factor. Power cord length > 1.5 m. Low noise < 45 dB.
18.2 | Configurations / Options | Indoor operation, enclosure protection at least IP20. Suitable for use on grounded and non-grounded mains, with the possibility to use an alternative grounding. Optional voltage and power input and output metering. Optional built-in isolating transformer (power conditioner), available at least for capacity > 1.5 kVA. Optional built-in output circuit breaker, independent for each output. Availability of versions with and without output circuit breaker, one for each output. Optional output spike, surge and transient protection (including lightning), with availability of Type I and Type II IEC rating lightning surge protection. Availability of different delay on start-up and on automatic reconnection, including (but not limited to): 0 s/immediate on power reconnection or start-up, 3–4 minutes, adjustable. Capacity range (depending on the models) 1–30 kVA: - models up to 1 kVA shall be lightweight and portable (at least 0.5 and 1 kVA); - models from 1 kVA to 10 kVA shall be mobile/wheeled (at least 1, 1.5–2, 2.5–3.5, 5, 7–10 kVA); - models from 10 kVA to 30 kVA shall be available in wheeled and fixed version (at least 7–10, 10–15, 20–25, 30 kVA); - wired and plugged version availability, with customizable plug (inlet) and sockets (outlets) matching all commercial electrical standards (at least, but not limited to, European/German, British, French, American, Australian, Indian).
19 | Displayed parameters | Input and output voltage, load (indication/measurement/metering options depending on the configuration).
20 | User adjustable settings | Input/output voltage (depending on the configuration)
Physical/chemical characteristics
21 | Components(if relevant) | Voltage regulator body, sockets/plugs or wiring (fixed installations), switches, display, fuses and circuit breakers, electronic boards, solid state electronics for voltage regulation, isolating transformer (if applicable), surge protection electronics (if applicable).
22 | Mobility, portability(if relevant) | Portable/mobile or fixed installation, depending on the configuration
23 | Raw Materials(if relevant) | N.A.
Utility requirements
24 | Electrical, water and/or gas supply (if relevant) | The requirements for power input voltage/frequency and plug type of the equipment must be chosen according to the local electrical supply. Electrical protection by resettable circuit breakers in both live and neutral supply lines
Accessories, consumables, spare parts, other components
25 | Accessories (if relevant) | Power cables, plugs and sockets, matching the applicable electric standard in the installation site.
26 | Sterilization process for accessories (if relevant) | N.A.
27 | Consumables / reagents (if relevant) | N.A.
28 | Spare parts (if relevant) | Sockets/plugs and wiring, switches, fuses/circuit breakers, display, control panel, external housing, electronic boards, isolating transformer (if applicable).
29 | Other components (if relevant) | N.A.
Packaging
30 | Sterility status on delivery (if relevant) | Non sterile.
31 | Shelf life (if relevant) | N.A.
32 | Transportation and storage (if relevant) | Sealed container. Capable of being transporterd and stored in ambient temperature of at least 5–50 °C, relative humidity of at least 15–95% non condensing. Specific requirements for altitude may be required, depending on the installation site.
33 | Labelling (if relevant) | Primary packaging: Unit of use: one (1) voltage regulator in a box or case with manufacturer’s instruction for use, spare parts and accessories (when applicable). Labelling on the primary packaging: Name and/or trademark of the manufacturer; manufacturer’s product reference; type of product and main characteristics; lot number prefixed by the word “LOT” (or equivalent harmonized symbol) (if applicable); information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonized symbol); information for handling, if applicable (or equivalent harmonized symbol). Over packaging: Packaging unit. Labelling on the packaging unit: Labelling to be the same as primary packaging. Extra information required: number of units
Environmental requirements
34 | Context-dependent requirements | Capable of being stored in ambient temperature of at least 5–50 °C, relative humidity of at least 15–95% non condensing. Suitable for continuous operation in ambient temperature of at least 5–45 °C, relative humidity of at least 15–90% non condensing. Specific requirements for altitude may be required, depending on the installation site.
Training, installation and utilization
35 | Pre-installation requirements(if relevant) | Compatibility with the electrical system (design and specifications).
36 | Requirements for commissioning (if relevant) | Recommended that the supplier performs installation, safety and operation checks before handover.Local technical staff to affirm completion of installation.
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided upon request. Training of technical staff in advanced maintenance tasks shall be provided upon request.
38 | User care(if relevant) | Pre-use checks (for portable and mobile versions). Proper connection, avoid ungrounded outlet, do not use with inappropriate extension cords or adapters (for portable and mobile versions). Avoid exceeding the maximum power output. Cleaning with compatible products. Periodic maintenance, functionality and electrical safety checks with electrical testers.
Warranty and maintenance
39 | Warranty | 3 years recommended (2 years at least), the product shall be in production and fully supported when procured.
40 | Maintenance tasks | Preventive maintenance, safety checks and functionality tests at least once per year.
41 | Type of service contract | Not required.
42 | Spare parts availability post-warranty | 8 years in total, including the warranty period
43 | Software / Hardware upgrade availability | N.A.
Documentation
44 | Documentation requirements | User and maintenance manuals, hard and soft copies, to be supplied in local language, or agreed other language if local not available. English version must also be available. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance. List to be provided of common spares and accessories, with part numbers. Contact details of manufacturer, supplier and local service agent to be provided.
Decommissioning
45 | Estimated Life Span | 10 years.
Safety and standards
46 | Standards, for the manufacturer and the equipment | N.A.
47 | Regulatory Approval / Certification | Product must be approved by the FDA, CE, IEEE, NFPA, OSHA and/or other internationally recognized regulatory body
48 | International standards (for reference) | Standards applicable to the manufacturer and the manufacturing process (compliance to the last available version is recommended, proof of compliance must be provided): ISO 13485 Medical devices – Quality management systems – Requirements for regulatory purposes or ISO 9001 Quality management systems – Requirements. Standards applicable to the product (where applicable, compliance to the last available version is recommended, proof of compliance must be provided): IEC 60601-1 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance. IEC 61643-11 Low-voltage surge protective devices – Part 11: Surge protective devices connected to low-voltage power systems – Requirements and test methods (replaces IEC 61643-1). IEC 61643-21 Low voltage surge protective devices – Part 21: Surge protective devices connected to telecommunications and signalling networks – Performance requirements and testing methods. IEC 61643-22 Low-voltage surge protective devices – Part 22: Surge protective devices connected to telecommunications and signalling networks – Selection and application principles. UL 1449 Standard for surge protective devices. Telcordia Technologies technical reference TR-NWT-001011. ANSI/IEEE C62.xx, including (but not limited to) C62.72 Application of surge devices for low voltage AC circuits, C62.47 Electrostatic discharge, C62.41.2 Recommended practice on characterization of surges in low voltage. IEEE-1100 Powering and grounding electronic equipment. NFPA 780 Lightning protection systems. EN 62305 Lightning protection. AS/NZS 1768 Lightning protection. UL 60950-1 Information technology equipment – Safety – Part 1: General requirements
49 | Regional/local standards | Country-specific and regional standards apply and must be listed
50 | Regulations | Country-specific and regional standards apply and must be listed