## Foglio1
Paediatric foley catheters (The specification detailed below is for general reference and it refers to different sizes, not only to paediatric)
Published on: 19 November 2020 - Version: 1
WHO Item name: | Required technical specifications: | Dimensions / Sizes / Presentation: | Material: | Certificates for product safety and quality assurance, as requested or equivalent
Catheter, urethral, Foley, 2-way, sterile, single use. Set of different sizes | A soft thin rubber tube with a balloon at the nelaton tip, designed for insertion in the bladder cavity, via the urethra, in order to drain off urine, instil a liquid or irrigate the bladder. A standard catheter consists of a hollow 2-way cylindrical tube with one central channel for urinary drainage; bladder side ending, Foley type, with a rounded atraumatic end tip (nelaton tip), with two opposing eyelets and one balloon; collector side ending that is a universal and hollow truncated cone (funnel) to connect the urine bag, spigot, syringe or irrigating device; and the balloon port, side channel ending with a non-return valve and a Luer tip connector. Catheter size is expressed in French gauge or Charrier (Fr or CH) and colour coded; balloon expansion capacity is expressed in mL. Dimensions and colour code must be legible and visible on the connectors. Preferable intended use for this size: children, medium- or long-term catheterization. Initial sterilization method: ethylene oxide gas or equivalent if applicable. Individually peel-packed in paper and/or plastic. Double-packaged: protected with an interior layer and an outer peel pack. Each set provided is composed by at least the following external diameter sizes: 10 Fr, 12 Fr, 14 Fr, 16 Fr and 18 Fr. | Length: 30 cm. External diameter sizes provided: 10 Fr, 12 Fr, 14 Fr, 16 Fr and 18 Fr. Balloon expansion capacity: 3 to 5 mL. | Silicone-coated natural latex. | ISO 20696:2018: Sterile urethral catheters for single use. ISO 11135:2014: Sterilization of health-care products – Ethylene oxide – Requirements for the development, validation and routine control of a sterilization process for medical devices.