## Oxygen analyzer, ultrasonic
Oxygen analyzer, ultrasonic, handheld
Published on:
Version 3:
Version 2:
Version 1:
Technical Specification | Hand held, battery powered device that measures the oxygen concentration in a flow of gas from a medical gas source or, with adapters, through a medical gas-flow device such mum requirementsHand held, battery powered device that measures the oxygen concentration in a flow of gas from a medical gas source or, with adapters, through a medical gas-flow device such as a ventilator or anaesthesia system.
General technical requirements | Ultrasonic oxygen sensing technology, with permanent sensing chamber. O2 measurement to include a range of at least between 21 and 96%. O2 resolution: at least within 0.1%. O2 accuracy: at least within ± 1.5% of full scale. Flow rate measurement: at least up to 10 LPM Flow rate accuracy at least within ± 0.2 LPM Flow rate resolution at least within 0.1 LPM Suitable for measuring gas supply with pressure at least up to 50 psi (344 kPa) Response time of a maximum of < 20 s. Warm-up time of at least < 1 s. Casing and environment IPX1 ingress protection. Suitable for cleaning and disinfection with hospital-grade cleaning products.
Displayed parameters | Display visualizing O2 concentration, flow rate, system messages, and battery status.
User adjustable settings | N/A
Alarms | N/A
Accessories (included and mentioned in a disaggregated list) | Protective case
Sterilization process for accessories (if relevant) | Supplier to describe any sterilization process required for accessories, if any.
Consumables / reagents (if relevant) | Connector tube for gas sampling
Spare parts (included and mentioned in a disaggregated list) | Set of single-use batteries Supplied empty of charge. 1-year spare parts kit as per preventive maintenance programme, preferable.
Mobility, portability | Portable
Power supply, voltage, frequency and plug vary across countries | Operated by alkaline replaceable single use batteries. Battery life > 1000 hours continuous use. Automatic power-off when not in use.
Documentation requirements (English language mandatory) | User/technical manual to be supplied in English (provision of versions in other UN languages, if available, will be an asset). Certificate of calibration and inspection. Troubleshooting, calibration and routine maintenance. List of all spare parts and accessories, with part numbers and contact details for parts supply. Document with contact details of manufacturer, supplier and local service agent.
Environmental requirements | Capable of being stored continuously in ambient temperature of 0–50 °C and relative humidity of 15–85%, preferably 90%. Capable of operating continuously in ambient temperature of 10–40 °C and relative humidity of 15–85%, preferably 90%.
Primary packaging label | Labelling on the primary packaging to include: Name and/or trademark of the manufacturer. Model or product’s reference. Information for particular storage conditions (temperature, pressure, light, humidity). Manufacturer’s product reference Manufacture and expiry date, lot number Main characteristics (dimensions, capacity) Recommended storage conditions Information for handling
Standards, for the manufacturer | Certified quality management system for medical devices (e.g. ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory purposes). General quality management (e.g. ISO 9001:2015 Quality management systems – Requirements). Application of risk management to medical devices (e.g. ISO 14971:2019 Medical devices – Application of risk management to medical devices).
Regulatory approval/ certification | Free sales certificate (FSC). Certified quality management system for medical devices (e.g. ISO 13485, or good manufacturing practice (GMP)). Certificate for exportation of medical device provided by the authority in the manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
Standards, for product performance | ISO 13485:2003 Medical devices -Quality management systems -- Requirements for regulatory purposes EN ISO 14971:2012 Medical devices — Application of risk management to medical devices EN 1041:2008 Information supplied by the manufacturer of medical devices EN ISO 15223-1:2016 Medical devices — Symbols to be used with medical device labels, labelling and information to be supplied — Part 1: General requirements
Estimated Life Span | Not less than 5 years.
Spare parts availability post-warranty | Guaranteed time period of availability of spare parts post-warranty shall be described.
Warranty | Not less than 24 months with specific inclusions and exclusions to be clearly listed. Contact details of manufacturer, supplier and local service agent to be provided.
Maintenance tasks | List shall be provided of equipment and procedures required for local calibration and routine maintenance.
Training of user/s (if relevant) | Training of users in operation and technicians in basic maintenance shall be provided.