## Content
Technical specifications for selection of essential in vitro diagnostics (WHO-EDL)
This tool presents technical specifications for 15 in vitro diagnostic tests listed in the 3rd WHO-EDL (2021)
Each technical specification is presented in a standard template organized as follows: name of the IVD test or laboratory equipment and then 88 technical specifications organized into 13 categories: Name, category and coding; intended use; performance characteristics; technical and operational characteristics; instrument physical and technical characteristics; infrastructure required; accessories, consumables, spare parts and other components; documentation; environmental and safety requirements; training, installation and use; warranty and maintenance; decommissioning; quality and registration.
To review technical specifications for a specific IVD category, go to the test menu tab and click on the name of the category. The selected item will be displayed on the screen. Scroll down to review the generic technical characteristics.

## Test menu
Technical specifications for selection of essential in vitro diagnostics: WHO-EDL 3
No. | Page in EDL-3 + laboratory setting | GMDN code/ Term name | EMDN code/ Term name | Diagnostic test name | Test purpose | Assay format
1 | P288 for no laboratory | 51801/ Treponema pallidum total antibody IVD, kit, rapid ICT, clinical | W0105090105/ SYPHILIS - RT & POC | Antibodies to Treponema pallidum | For diagnosis or as an aid in the diagnosis of T. pallidum | Rapid diagnostic test (RDT)
2 | P288 for no laboratory | 63667/ HIV1/HIV2/Treponema pallidum antibody IVD, kit, rapid ICT, clinical | Not available | Combined antibodies to T. pallidum and to HIV-1/2 | For diagnosis or as an aid in the diagnosis of HIV-1/2 and/or T. pallidum | Rapid diagnostic test (RDT)
3 | P339 for clinical laboratory | 51801/ Treponema pallidum total antibody IVD, kit, rapid ICT, clinical | W0105090105/ SYPHILIS - RT & POC | Antibodies to Treponema pallidum | For diagnosis or as an aid in the diagnosis of syphilis | Rapid diagnostic test (RDT)
4 | P339 for clinical laboratory | 51798/ Treponema pallidum total antibody IVD, kit, enzyme immunoassay (EIA) | W0105010302/ SYPHILIS ANTIBODY ASSAYS TOTAL | Antibodies to Treponema pallidum | For diagnosis or as an aid in the diagnosis of syphilis | Immunoassay (manual method) | EIA reader
5 | P339 for clinical laboratory | 51798/ Treponema pallidum total antibody IVD, kit, enzyme immunoassay (EIA) | W0105010302/ SYPHILIS ANTIBODY ASSAYS TOTAL | Antibodies to Treponema pallidum | For diagnosis or as an aid in the diagnosis of syphilis | Immunoassay (automated method)
6 | P339 for clinical laboratory | 63667/ HIV1/HIV2/Treponema pallidum antibody IVD, kit, rapid ICT, clinical | Not available | Combined antibodies to T. pallidum and to HIV-1/2 | For diagnosis or as an aid in diagnosis of HIV-1/2 infection and/ or syphilis | Rapid diagnostic test (RDT)
7 | P340 for clinical laboratory | 51819/ Treponema pallidum reagin antibody IVD, kit, agglutination | W0105010399/ SYPHILIS REAGENTS OTHER | Non-treponemal rapid plasma reagin (RPR) test | For screening for syphilis and monitoring treatment effectiveness | Particle/charcoal agglutination assay (manual method)
8 | P340 for clinical laboratory | 51819/ Treponema pallidum reagin antibody IVD, kit, agglutination | W0105010399/ SYPHILIS REAGENTS OTHER | Non-treponemal rapid plasma reagin (RPR) test | For screening for syphilis and monitoring treatment effectiveness | Particle/charcoal agglutination assay (automated method) | Automated Hemagglutination analyzer
9 | P340 for clinical laboratory | 51819/ Treponema pallidum reagin antibody IVD, kit, agglutination | W0105010399/ SYPHILIS REAGENTS OTHER | Non-treponemal Venereal Disease Research Laboratory (VDRL) test | For screening, diagnosis and confirmation of neurosyphilis | Flocculation test (manual method)
10 | P340 for clinical laboratory | 51800/ Treponema pallidum total antibody IVD, kit, agglutination | W0105010399/ SYPHILIS REAGENTS OTHER | T. pallidum haemagglutination (TPHA) test | For confirmation of syphilis infection and diagnosis of early and late syphilis infection | Red cell agglutination assay (manual method)
11 | P340 for clinical laboratory | 51800/ Treponema pallidum total antibody IVD, kit, agglutination | W0105010399/ SYPHILIS REAGENTS OTHER | T. pallidum haemagglutination (TPHA) test | For confirmation of syphilis infection and diagnosis of early and late syphilis infection | Red cell agglutination assay (automated method) | Automated Hemagglutination analyzer
12 | P340 for clinical laboratory | 51800/ Treponema pallidum total antibody IVD, kit, agglutination | W0105010399/ SYPHILIS REAGENTS OTHER | T. pallidum particle agglutination (TPPA) test | For confirmation of syphilis infection and diagnosis of early and late syphilis infection | Particle agglutination assay (manual method)
13 | P340 for clinical laboratory | 51800/ Treponema pallidum total antibody IVD, kit, agglutination | W0105010399/ SYPHILIS REAGENTS OTHER | T. pallidum particle agglutination (TPPA) test | For confirmation of syphilis infection and diagnosis of early and late syphilis infection | Particle agglutination assay (automated method)
14 | P350 for blood-screening laboratory | 51798/ Treponema pallidum total antibody IVD, kit, enzyme immunoassay (EIA) | W0105010302/ SYPHILIS ANTIBODY ASSAYS TOTAL | Antibodies to T. pallidum | For screening blood donations for syphilis | Immunoassay (manual method) | EIA reader
15 | P350 for blood-screening laboratory | 51798/ Treponema pallidum total antibody IVD, kit, enzyme immunoassay (EIA) | W0105010302/ SYPHILIS ANTIBODY ASSAYS TOTAL | Antibodies to T. pallidum | For screening blood donations for syphilis | Immunoassay (automated method)
Excel cells in cream show disease-specific IVDs for use in community settings and health facilities without a laboratory.
Excel cells in green show disease-specific IVDs for use in clinical laboratories.
Excel cells in pink show disease-specific IVDs for blood-screening laboratories.

## 1
Antibodies to Treponema pallidum
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51801/ Treponema pallidum total antibody IVD, kit, rapid ICT, clinical
3 | EMDN code/ Term name | W0105090105/ SYPHILIS - RT & POC
4 | Generic name | Antibodies to Treponema pallidum
5 | Specific type or variation (optional) | Qualitative, immunochromatographic assay
6 | Alternative name/s (optional) | Rapid diagnostic test
7 | Keywords (optional) | Treponema pallidum, syphilis, antibodies to Treponema pallidum, RDT
Intended use
8 | Detection target | Anti-treponemal antibodies
9 | Test purpose | Diagnosis or aid in the diagnosis of syphilis
10 | Specific disorder / condition | Syphilis, sexually or congenitally transmitted infections
11 | Test format | Rapid diagnostic test
12 | Specimen type | Capillary whole blood, venous whole blood, plasma and serum
13 | Testing population | Adults, including pregnant women
14 | Intended users | Trained health-care providers, trained lay and peer workers
15 | Level of health-care system | For use in community settings and health facilities without laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 90%
17 | Clinical specificity | ≥ 97%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6: https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Not applicable
20 | Rate of invalidity, error or unreturnable | < 5%
21 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): agreement ≥ 95%; inter-assay precision (reproducibility): agreement ≥ 95%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Rapid lateral flow immunochromatographic assay (test strips or devices coated with recombinant T. pallidum antigen to produce a colored line in case of reactive test)
24 | Specimen(s) stability | Whole blood collected by fingerstick should be tested immediately. Whole blood can be stored at 2–8 °C for 24 h. Do not freeze whole blood specimens. Plasma and serum specimens can be stored at 2–8 °C for 72 h. For prolonged storage, plasma and serum specimens can be stored at –20 °C.
25 | Specimen(s) volume | ≤ 80 µL
26 | Type of result | Qualitative (reactive or non-reactive)
27 | Time to result | ≤ 20 min
28 | End-point stability | ≥ 5 min
29 | Ease of use for point-of-care tests only: number of steps that require precision | Maximum 2 steps
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≧10 specimens per hour per operator
31 | Test limitations | The manufacturer should state any limitations.
32 | Internal quality control | Built-in quality control line that must appear when specimen is added
33 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected against humidity and direct sunlight
38 | Transport stability of specimen (temperature and humidity) | Whole blood specimen at 2–8 °C for 24 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 72 h Plasma and serum at –20 °C for > 72 h Avoid multiple freeze–thaw cycles.
39 | Storage stability of specimen (temperature and humidity) | Whole blood specimen at 2–8 °C for 24 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 72 h Plasma and serum at –20 °C for > 72 h Avoid multiple freeze–thaw cycles.
40 | In-use stability of specimen (temperature and humidity) | Whole blood specimen at 2–8 °C for 24 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 72 h Avoid multiple freeze–thaw cycles.
41 | Transport stability of controls/calibrators (temperature and humidity) | Not applicable
42 | Storage stability of controls/calibrators (temperature and humidity) | Not applicable
43 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability of controls/calibrators (temperature and humidity) | Not applicable
45 | Shelf-life of kit/reagents at time of manufacture (months) | At least 12 months after date of manufacture
46 | Remaining shelf-life of kit/reagents at delivery (months) | At least 6 months’ shelf-life remaining when product arrives at point of use. See p. 16 of the Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Display parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability / Test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Not required
62 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature will require electricity.
63 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits
64 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | Test strips, devices or cartridges and instructions for use or a package insert
66 | Reagent kit size (number of tests) | ≥ 10 tests
67 | Consumables required but not provided in the test kit | For collecting finger-prick whole blood, EDTA capillary tubes or tips for micropipette and lancets are required. For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if controls (reactive and non reactive) are not provided in the kits, it is desirable to prepare controls.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection or for collecting finger-prick whole blood is required. For testing, a timer is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environment and safety requirements
74 | Hazard classification | Potential source of harm classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. Usually stated on the material safety data sheet
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g., simple trash or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of technical support in the country | Availability of a local authorized service provider or distributor is recommended.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification (PQ) status | This assay with RDT format is within the scope of WHO PQ. There is a WHO PQ product and prequalification is recommended as of July 2022. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 2
Combined antibodies to Treponema pallidum and HIV-1/2
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 63667/ HIV1/HIV2/Treponema pallidum antibody IVD, kit, rapid ICT, clinical
3 | EMDN code/ Term name | Not available
4 | Generic name | Combined antibodies to Treponema pallidum and to HIV-1/2
5 | Specific type or variant (optional) | Qualitative, immunochromatographic assay
6 | Alternative name/s (optional) | Rapid diagnostic test
7 | Keywords (optional) | T. pallidum and HIV-1/2 antibodies, rapid diagnostic test, RDT
Intended use
8 | Detection target | Antibodies to T. pallidum and to HIV-1/2
9 | Test purpose | To diagnose or to aid in the diagnosis of HIV-1/2 infection and/or syphilis
10 | Specific disorder / condition | Syphilis and HIV, sexually or congenitally transmitted infections
11 | Test format | Rapid diagnostic test
12 | Specimen type | Capillary whole blood, whole blood, venous whole blood, plasma and/or serum
13 | Testing population | Adults, including pregnant women
14 | Intended users | Trained health-care providers, trained lay and peer workers
15 | Level of the health-care system | For use in community settings and health facilities without laboratories
Performance characteristics
16 | Clinical sensitivity | HIV antibodies: ≥ 99% T. pallidum antibodies: ≥ 87%
17 | Clinical specificity | HIV antibodies : ≥ 98% T. pallidum antibodies : ≥ 93%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1) and WHO TSS-1: Human immunodeficiency virus (HIV) rapid diagnostic tests for professional use and/or self-testing (https://apps.who.int/iris/bitstream/handle/10665/251857/9789241511742-eng.pdf?sequence=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Not applicable
20 | Rate of invalidity, error or unreturnable | < 5%
21 | Precision (repeatability/reproducibility) | HIV: Intra-assay precision (repeatability): agreement ≥ 99%; inter-assay precision (reproducibility): agreement ≥ 93% T. pallidum and syphilis: Intra-assay precision (repeatability): agreement ≥ 95%; inter-assay precision (reproducibility): agreement ≥ 95%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Rapid flow immunochromatographic assay
24 | Specimen(s) stability | Whole blood collected by finger-stick should be tested immediately. Blood specimens should be tested as soon as possible after collection, at least within 1 h. Do not freeze whole-blood specimens. Venous whole blood can be stored at 2–8 °C for 24 h. Plasma and serum specimens should be stored at 2–8 °C for 3 days. Plasma and serum specimens should be stored at - 20 C for prolonged storage,
25 | Specimen(s) volume | ≤ 80 µL
26 | Type of result | Qualitative
27 | Time to result | 15–20 min
28 | End-point stability | Fixed reading time. Follow the manufacturer's instructions for use. Generally, 5–15 min
29 | Ease of use for point-of-care tests only: number of steps that require precision | 2–4 steps
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 4 specimens per hour per operator
31 | Test limitations | The manufacturer should state any limitations.
32 | Internal quality control | Built-in quality control line that must appear when specimen is added
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–30 °C protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–30 °C protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–30 °C protected against humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated.
38 | Transport stability of specimen (temperature and humidity) | Venous whole blood specimen at 2–8 °C for 24 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 3 days Plasma and serum at –20 °C for > 3 days Avoid multiple freeze–thaw cycles.
39 | Storage stability of specimen (temperature and humidity) | Venous whole blood specimen at 2–8 °C for 24 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 3 days Plasma and serum at –20 C for > 3 days Avoid multiple freeze–thaw cycles.
40 | In-use stability of specimen (temperature and humidity) | Venous whole blood specimen at 2–8 °C for 24 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 3 days Plasma and serum at –20 C for > 3 days Avoid multiple freeze–thaw cycles.
41 | Transport stability of controls/calibrators (temperature and humidity) | Not applicable
42 | Storage stability of controls/calibrators (temperature and humidity) | Not applicable
43 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability of controls/calibrators (temperature and humidity) | Not applicable
45 | Shelf-life of kit/reagents at time of manufacture (months) | At least 12 months for RDT kits and 6 months for controls (if used) from date of manufacture
46 | Remaining shelf-life of kit/reagents at time of delivery (months) | At least 6 months’ shelf-life remaining when product arrives at point of use. See p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Displayed parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability / Test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Not applicable
62 | Electricity requirements | Electricity not required for conducting the test; however, storage of some kits at low temperature will require electricity.
63 | Refrigeration or cold chain for storage of kit/reagents | Yes if the temperature in the storage area for the kits exceeds the upper limit of the storage temperature of the test kits.
64 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | Test strips, devices or cartridge, buffer, consumables necessary for finger-pricking and instructions for use or package insert
66 | Reagent kit size (number of tests) | ≥ 10 tests
67 | Consumables required but not provided in the test kit | For collecting finger-prick whole blood, EDTA capillary tubes or tips for micropipette and lancets are required. For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if controls (reactive and non reactive) are not provided in the kits, it is desirable to prepare controls.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection or for collecting finger-prick whole blood is required. For testing, a timer is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specification. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | Potential source of harm classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. Usually stated on the material safety data sheet
75 | Disposal requirements | Description of disposal techniques according to the hazard classification, e.g., simple trash or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of technical support in the country | A local authorized service provider or distributor is recommended.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification (PQ) status | The rapid diagnostic test format of the assay is within the scope of WHO PQ. There are three WHO PQ products as of July 2022, and prequalification is recommended. See WHO Public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 3
Antibodies to Treponema pallidum
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51801/ Treponema pallidum total antibody IVD, kit, rapid ICT, clinical
3 | EMDN code/ Term name | W0105090105/ SYPHILIS - RT & POC
4 | Generic name | Antibodies to Treponema pallidum
5 | Specific type or variation (optional) | Qualitative, immunochromatographic assay
6 | Alternative name/s (optional) | Rapid diagnostic test
7 | Keywords (optional) | Treponema pallidum, syphilis, antibodies to Treponema pallidum, RDT
Intended use
8 | Detection target | Anti-treponemal antibodies
9 | Test purpose | Diagnosis or aid in the diagnosis of syphilis
10 | Specific disorder / condition | Syphilis, sexually or congenitally transmitted infections
11 | Test format | Rapid diagnostic test
12 | Specimen type | Capillary whole blood, venous whole blood, plasma and serum
13 | Testing population | Adults, including pregnant women
14 | Intended users | Trained laboratory scientists, technologists, technicians or specialists with expertise
15 | Level of health-care system | For use in clinical laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 90%
17 | Clinical specificity | ≥ 97%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6: https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Not applicable
20 | Rate of invalidity, error or unreturnable | < 5%
21 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability): agreement ≥ 95%; inter-assay precision (reproducibility): agreement ≥ 95%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Rapid lateral flow immunochromatographic assay (test strips or devices coated with recombinant T. pallidum antigen to produce a colored line in case of reactive test)
24 | Specimen(s) stability | Whole blood collected by fingerstick should be tested immediately. Whole blood can be stored at 2–8 °C for 24 h. Do not freeze whole blood specimens. Plasma and serum specimens can be stored at 2–8 °C for 72 h. For prolonged storage, plasma and serum specimens can be stored at –20 °C.
25 | Specimen(s) volume | ≤ 80 µL
26 | Type of result | Qualitative (reactive or non-reactive)
27 | Time to result | ≤ 20 min
28 | End-point stability | ≥ 5 min
29 | Ease of use for point-of-care tests only: number of steps that require precision | Maximum 2 steps
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≧10 specimens per hour per operator
31 | Test limitations | The manufacturer should state any limitations.
32 | Internal quality control | Built-in quality control line that must appear when specimen is added
33 | Compatibility with external quality control material | Testing known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–30 °C, protected against humidity and direct sunlight
38 | Transport stability of specimen (temperature and humidity) | Whole blood specimen at 2–8 °C for 24 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 72 h Plasma and serum at –20 °C for > 72 h Avoid multiple freeze–thaw cycles.
39 | Storage stability of specimen (temperature and humidity) | Whole blood specimen at 2–8 °C for 24 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 72 h Plasma and serum at –20 °C for > 72 h Avoid multiple freeze–thaw cycles.
40 | In-use stability of specimen (temperature and humidity) | Whole blood specimen at 2–8 °C for 24 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 72 h Avoid multiple freeze–thaw cycles.
41 | Transport stability of controls/calibrators (temperature and humidity) | Not applicable
42 | Storage stability of controls/calibrators (temperature and humidity) | Not applicable
43 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability of controls/calibrators (temperature and humidity) | Not applicable
45 | Shelf-life of kit/reagents at time of manufacture (months) | At least 12 months after date of manufacture
46 | Remaining shelf-life of kit/reagents at delivery (months) | At least 6 months’ shelf-life remaining when product arrives at point of use. See p. 16 of the Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Display parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability / Test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Not required
62 | Electricity requirements | Electricity is not required to conduct the test; however, storage of some kits at low temperature will require electricity.
63 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kit exceeds the upper limit of the storage temperature of the test kits
64 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | Test strips, devices or cartridges and instructions for use or a package insert
66 | Reagent kit size (number of tests) | ≥ 10 tests
67 | Consumables required but not provided in the test kit | For collecting finger-prick whole blood, EDTA capillary tubes or tips for micropipette and lancets are required. For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if controls (reactive and non reactive) are not provided in the kits, it is desirable to prepare controls.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection or for collecting finger-prick whole blood is required. For testing, a timer is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environment and safety requirements
74 | Hazard classification | Potential source of harm classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. Usually stated on the material safety data sheet
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g., simple trash or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of technical support in the country | Availability of a local authorized service provider or distributor is recommended.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification (PQ) status | This assay with RDT format is within the scope of WHO PQ. There is a WHO PQ product and prequalification is recommended as of July 2022. Please see WHO Public Reports for In Vitro Diagnostics https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr.
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 4
Antibodies to Treponema pallidum
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code | Antibodies to Treponema pallidum, immunoassay
2 | GMDN code/ Term name | 51798/ Treponema pallidum total antibody IVD, kit, enzyme immunoassay (EIA)
3 | EMDN code/ Term name | W0105010302/ SYPHILIS ANTIBODY ASSAYS TOTAL
4 | Generic name
5 | Specific type or variation (optional) | Qualitative assay
6 | Alternative name/s (optional)
7 | Keywords (optional) | T. pallidum, syphilis, antibodies to T. pallidum, microtitre plate enzyme immunoassay
Intended use
8 | Detection target | Anti-treponemal antibodies
9 | Test purpose | To diagnosis or to aid in the diagnosis of syphilis
10 | Specific disorder / condition | Syphilis, sexually or congenitally transmitted infections
11 | Test format | Immunoassay (manual method)
12 | Specimen type | Plasma and serum
13 | Testing population | Adults, including pregnant women
14 | Intended users | Trained laboratory scientists, technologists, technicians and specialists with expertise
15 | Level of the health-care system | For use in clinical laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 98%
17 | Clinical specificity | ≥ 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Not applicable
20 | Rate of invalidity, error and unreturnable | < 5%
21 | Precision (repeatability/reproducibility) | With reactive specimens: Intra-assay precision (repeatability): coefficient of variation ≤ 10% inter-assay precision (reproducibility): coefficient of variation ≤ 20%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Enzyme immunoassay, microwells coated with T. pallidum antigen, specific for T. pallidum
24 | Specimen(s) stability | Specimens can be stored at 2–8 °C for 2 days. For prolonged storage, plasma and serum specimens can be stored at – 20 °C.
25 | Specimen(s) volume | ≤ 100 µL
26 | Type of result | Qualitative (reactive or non-reactive)
27 | Time to result | 3–4 h
28 | End-point stability | Fixed reading time. Follow the manufacturer's instructions on use or package insert.
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 40 specimens per day per operator
31 | Test limitations | The manufacturer should state any limitations.
32 | Internal quality control | Antibodies to T. pallidum, positive control, negative control and blank control
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C protected against humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated.
38 | Transport stability for specimen collection (temperature and humidity) | Whole blood specimen at 2–8 °C for 8 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 48 h Plasma and serum at – 20 °C for > 48 h Avoid multiple freeze–thaw cycles.
39 | Storage stability for specimen collection (temperature and humidity) | Whole blood specimen at 2–8 °C for 8 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 48 h Plasma and serum at – 20 °C for > 48 h Avoid multiple freeze–thaw cycles.
40 | In-use stability for specimen collection (temperature and humidity) | Whole blood specimen at 2–8 °C for 8 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 48 h Plasma and serum at – 20 °C for more 48 h Avoid multiple freeze–thaw cycles.
41 | Transport stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
42 | Storage stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
43 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight. Any specific requirement once controls are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents at the time of manufacture (months) | ≥ 12 months for kits/reagents and 6 months for controls/calibrators from the date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months’ shelf-life remaining when product arrives at point of use. See p. 16 of Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Display parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability / Test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Distilled or deionized water with adequate quality for enzyme immunoassays and enzyme-linked immunosorbent assays. See the instruction for use or package insert of kit or reagents.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | Coated microwells or coated microwell strips, calibrator, substrate, positive and negative controls, conjugate, wash buffer and stop solution
66 | Reagent kit size (number of tests) | ≥ 96 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues and some reagents that the kits do not contain are required. See the package insert or instruction for use of the kit or reagent.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection is required. For testing, a microwell plate or strip reader, an incubator or water bath, a heating block, a microwell plate or strip washer, a vacuum pump, vortex, calibrated single and multi-channel micropipettes, an electronic scale with hoods and a cover for calibrating the micropipettes, a micropipette stand, a timer, a cover for plate, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | Potential source of harm classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. Usually stated on the material safety data sheet
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g., simple trash or incineration.
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of technical support in the country | A local authorized service provider or distributor is recommended.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approval | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
88 | WHO prequalification (PQ) status | The manual method format of this assay is not within the scope of WHO PQ as of July 2022. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 5 Auto
Antibodies to Treponema pallidum |  | The technical specification below is for a particular test to be run on an automated analyser. As automated analysers are usually used to run various tests, selection of the equipment should take into account the other tests performed in the facility and not just that in the technical specification.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51798/ Treponema pallidum total antibody IVD, kit, enzyme immunoassay (EIA)
3 | EMDN code/ Term name | W0105010302/ SYPHILIS ANTIBODY ASSAYS TOTAL
4 | Generic name | Antibodies to Treponema pallidum, immunoassay
5 | Specific type or variation (optional) | Qualitative
6 | Alternative name/s (optional)
7 | Keywords (optional) | T. pallidum, syphilis, antibodies to T. pallidum, chemiluminescence immunoassay (CLIA)
Intended use
8 | Detection target | Anti-treponemal antibodies
9 | Test purpose | To diagnose or to aid in the diagnosis of syphilis
10 | Specific disorder / condition | Syphilis, sexually or congenitally transmitted infections
11 | Test format | Immunoassay (automated method)
12 | Specimen type | Plasma and serum
13 | Testing population | Adults, including pregnant women
14 | Intended users | Trained laboratory scientists, technologists, technicians and specialists with expertise
15 | Level of the health-care system | For use in clinical laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 93%
17 | Clinical specificity | ≥ 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Not applicable
20 | Rate of invalidity, error or unreturnable | Not applicable
21 | Precision (repeatability/reproducibility) | With reactive samples: intra-assay precision (repeatability): coefficient of variation ≤ 5% inter-assay precision (reproducibility): coefficient of variation ≤ 6%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Chemiluminescent immunoassay
24 | Specimen(s) stability | Specimens can be stored at 2–8 °C for 4 days. For prolonged storage, specimens can be stored at –10 °C or colder.
25 | Specimen(s) volume | ≤ 150 µL
26 | Type of result | Qualitative
27 | Time to result | ≤ 30 min
28 | End-point stability | Not applicable
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | In a laboratory with a low sample volume: Specimens per day per analyser : ≤ 500 (Sample throughput per hour per analyser: ≤ 62) In a laboratory with a medium sample volume: Specimens per day per analyser: 500–1000 (Sample throughput per hour: 63–125) In a laboratory with a high sample volume: Specimens per day per analyser : ≥ 1000 (Sample throughput per hour : ≥ 126 )
31 | Test limitations | The manufacturer should state any limitations.
32 | Internal quality control | Quality control material should be run as defined by the manufacturer.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected against humidity and all heat and light sources
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected against humidity and all heat and light sources
36 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of the analyser's measuring unit
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected against humidity and all heat and light sources. Any specific requirement for reagents once they are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
38 | Transport stability of specimen (temperature and humidity) | At 2–8 °C for 4 days and at –20 °C or lower for > 48 h
39 | Storage stability of specimen (temperature and humidity) | At 2–8 °C for 4 days. At –20 °C or lower for > 4 days
40 | In-use stability of specimen | At 2–8 °C for 48 h and at –10 °C or lower for > 48 h
41 | Transport stability of controls/calibrators (temperature and humidity) | At 2–8 °C, in an upright position, protected against humidity and all heat and light sources
42 | Storage stability of controls/calibrators (temperature and humidity) | At 2–8 °C, in an upright position, protected against humidity and all heat and light sources
43 | On-board stability of controls/calibrators (temperature and humidity) | At the temperature of the analyser's measuring unit
44 | In-use stability of controls/calibrators (temperature and humidity) | At 2–8 °C, in an upright position, protected against humidity and all heat and light sources. Any specific requirement for controls or calibrators once opened should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents at time of manufacture (months) | At least 9 months for kits/reagents and 6 months for controls/calibrators from the date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months’ shelf-life remaining when the product arrives at the point of use. See p. 16 of the Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Depends on the specific brand or model and whether bench-top or floor-standing
48 | Weight of the device (kg) | Depends on the specific brand or model. Many analysers weigh 150–800 kg.
49 | Power requirements | 100–240 VAC / 50–60 Hz or 200–240 VAC / 50–60 Hz
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Ideally, an uninterrupted power source
53 | Operating conditions (temperature and humidity) | Temperature range: 18–30 °C. Humidity: < 80%. Altitude: ≤ 2000 m
54 | User interface | Simple test menu, simple PC as provided by the manufacturer, option for connection to laboratory computer and middle-ware
55 | Display parameters | Digital display of test results, electronic and printed readout
56 | Display languages | English or local language
57 | Built-in memory storage capacity | Analyser for low-volume sample laboratory: ≥ 800 tests Analyser for medium-volume sample laboratory: ≥ 2000 tests Analyser for high-volume sample laboratory: ≥ 10000 tests
58 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programmed database to allow information-sharing. Capability for data transmission, e.g., via USB, electronic health records and health programme databases
59 | Open or closed system | Both open and closed systems available
60 | Multi-disease testing capability / Test menu | Multi-disease testing
Infrastructure requirements
61 | Water requirements | Distilled or deionized water with adequate quality for chemiluminescence immunoassays. See the instructions for use or package insert of kit or reagents.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | In general, buffer/diluent, antigen and conjugate.
66 | Reagent kit size (number of tests) | ≥ 40 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, it depends on the test/reagents and please see the package insert or instruction for use of the test/reagents. In general, controls, calibrators, microplates (if the kit/reagents does not contains), specimen collection tubes and pipette tips are required.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette's stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection is required. For testing, an analyser (if specified by the kits/reagent, the specified one is required), a timer, calibrated micropipettes, a micropipette stand, a refrigerator, a freezer and an autoclave are required. Please see the package insert or instruction for use of the test/reagents. If the room temperature goes over the upper limit of the temperature for operating condition, an air conditioner is required.
69 | Spare parts (if relevant) | Basic spare parts that customers can replace themselves during maintenance. The manufacturer should supply a list of spare parts.
Documentation
70 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their reference numbers and a list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | The manufacturer should clearly document potential sources of harm on kit inserts. Harm should be classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. More details will usually be provided in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g., simple waste disposal, hazardous waste, incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range must be assessed.
77 | Requirements for installation and calibration | The manufacturer should be responsible for installation and initial setup of the analyser. The analyser should be in good working order, calibrated and quality-controlled before being handed over to users.
78 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuous training should be offered if appropriate.
Warranty and maintenance
79 | Warranty | The manufacturer must provide a minimum 1-year warranty from the date of handover and specify the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty.
80 | Preventive maintenance | The minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either a fixed period or according to the number of tests processed, which will be dictated by the workload.
81 | Corrective maintenance | Specific requirements for calibration and test procedures must be described in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or the supplier will specify the maintenance support they will supply.
82 | Type of service contract (including leasing or reagent rental) | Must cover costs of labour, repair, spare parts, shipping, logistics and training. Possible lending of small instruments
83 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
84 | Availability of technical support in the country | Required. The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
85 | Estimated lifespan | 5 years after the start of use/installation under the conditions of regular maintenance / inspection specified in the instruction manual, replacement of parts and necessary repairs or maintenance under normal use and operating conditions
Quality and registration
86 | Global regulatory approval | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
88 | WHO prequalification (PQ) status | The automated method format of this assay is not within the scope of WHO PQ as of August 2022. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 6
Combined antibodies to Treponema pallidum and HIV-1/2
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 63667/ HIV1/HIV2/Treponema pallidum antibody IVD, kit, rapid ICT, clinical
3 | EMDN code/ Term name | Not available
4 | Generic name | Combined antibodies to Treponema pallidum and to HIV-1/2
5 | Specific type or variant (optional) | Qualitative, immunochromatographic assay
6 | Alternative name/s (optional) | Rapid diagnostic test
7 | Keywords (optional) | T. pallidum and HIV-1/2 antibodies, rapid diagnostic test, RDT
Intended use
8 | Detection target | Antibodies to T. pallidum and to HIV-1/2
9 | Test purpose | To diagnose or to aid in the diagnosis of HIV-1/2 infection and/or syphilis
10 | Specific disorder / condition | Syphilis and HIV, sexually or congenitally transmitted infections
11 | Test format | Rapid diagnostic test
12 | Specimen type | Capillary whole blood, whole blood, venous whole blood, plasma and/or serum
13 | Testing population | Adults, including pregnant women
14 | Intended users | Trained laboratory scientists, technologists, technicians and specialists with expertise
15 | Level of the health-care system | For use in clinical laboratories
Performance characteristics
16 | Clinical sensitivity | HIV antibodies: ≥ 99% T. pallidum antibodies: ≥ 87%
17 | Clinical specificity | HIV antibodies : ≥ 98% T. pallidum antibodies : ≥ 93%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1) and WHO TSS-1: Human immunodeficiency virus (HIV) rapid diagnostic tests for professional use and/or self-testing (https://apps.who.int/iris/bitstream/handle/10665/251857/9789241511742-eng.pdf?sequence=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Not applicable
20 | Rate of invalidity, error or unreturnable | < 5%
21 | Precision (repeatability/reproducibility) | HIV: Intra-assay precision (repeatability): agreement ≥ 99%; inter-assay precision (reproducibility): agreement ≥ 93% T. pallidum and syphilis: Intra-assay precision (repeatability): agreement ≥ 95%; inter-assay precision (reproducibility): agreement ≥ 95%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Rapid flow immunochromatographic assay
24 | Specimen(s) stability | Whole blood collected by finger-stick should be tested immediately. Blood specimens should be tested as soon as possible after collection, at least within 1 h. Do not freeze whole-blood specimens. Venous whole blood can be stored at 2–8 °C for 24 h. Plasma and serum specimens should be stored at 2–8 °C for 3 days. Plasma and serum specimens should be stored at - 20 C for prolonged storage,
25 | Specimen(s) volume | ≤ 80 µL
26 | Type of result | Qualitative
27 | Time to result | 15–20 min
28 | End-point stability | Fixed reading time. Follow the manufacturer's instructions for use. Generally, 5–15 min
29 | Ease of use for point-of-care tests only: number of steps that require precision | 2–4 steps
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 4 specimens per hour per operator
31 | Test limitations | The manufacturer should state any limitations.
32 | Internal quality control | Built-in quality control line that must appear when specimen is added
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–30 °C protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–30 °C protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–30 °C protected against humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated.
38 | Transport stability of specimen (temperature and humidity) | Venous whole blood specimen at 2–8 °C for 24 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 3 days Plasma and serum at –20 °C for > 3 days Avoid multiple freeze–thaw cycles.
39 | Storage stability of specimen (temperature and humidity) | Venous whole blood specimen at 2–8 °C for 24 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 3 days Plasma and serum at –20 C for > 3 days Avoid multiple freeze–thaw cycles.
40 | In-use stability of specimen (temperature and humidity) | Venous whole blood specimen at 2–8 °C for 24 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 3 days Plasma and serum at –20 C for > 3 days Avoid multiple freeze–thaw cycles.
41 | Transport stability of controls/calibrators (temperature and humidity) | Not applicable
42 | Storage stability of controls/calibrators (temperature and humidity) | Not applicable
43 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability of controls/calibrators (temperature and humidity) | Not applicable
45 | Shelf-life of kit/reagents at time of manufacture (months) | At least 12 months for RDT kits and 6 months for controls (if used) from date of manufacture
46 | Remaining shelf-life of kit/reagents at time of delivery (months) | At least 6 months’ shelf-life remaining when product arrives at point of use. See p. 16 in Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Displayed parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability / Test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Not applicable
62 | Electricity requirements | Electricity not required for conducting the test; however, storage of some kits at low temperature will require electricity.
63 | Refrigeration or cold chain for storage of kit/reagents | Yes if the temperature in the storage area for the kits exceeds the upper limit of the storage temperature of the test kits
64 | Refrigeration or cold chain for reconstituted reagents and controls | Not applicable
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | Test strips, devices or cartridge, buffer, consumables necessary for finger-pricking and instructions for use or package insert
66 | Reagent kit size (number of tests) | ≥ 10 tests
67 | Consumables required but not provided in the test kit | For collecting finger-prick whole blood, EDTA capillary tubes or tips for micropipette and lancets are required. For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if controls (reactive and non reactive) are not provided in the kits, it is desirable to prepare controls.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection or for collecting finger-prick whole blood is required. For testing, a timer is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specification. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | Potential source of harm classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. Usually stated on the material safety data sheet
75 | Disposal requirements | Description of disposal techniques according to the hazard classification, e.g., simple trash or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of technical support in the country | A local authorized service provider or distributor is recommended.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subject to the contract is required.
88 | WHO prequalification (PQ) status | The rapid diagnostic test format of the assay is within the scope of WHO PQ. There are three WHO PQ products as of July 2022 and prequalification is recommended. See WHO Public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 7
Non-treponemal rapid plasma reagin (RPR) test
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51819/ Treponema pallidum reagin antibody IVD, kit, agglutination
3 | EMDN code/ Term name | W0105010399/ SYPHILIS REAGENTS OTHER
4 | Generic name | Non-treponemal rapid plasma reagin test
5 | Specific type or variation (optional) | Qualitative and quantitative assay
6 | Alternative name/s (optional)
7 | Keywords (optional) | Rapid plasma reagin, RPR, syphilis, particle/charcoal agglutination assay, non-treponemal RPR
Intended use
8 | Detection target | Reagin antibodies to non-treponemal antigens
9 | Test purpose | To screen for syphilis and monitor treatment effectiveness
10 | Specific disorder / condition | Syphilis, sexually or congenitally transmitted infections
11 | Test format | Particle/charcoal agglutination assay (manual method)
12 | Specimen type | Plasma and/or serum
13 | Testing population | Sexually active population, including, e.g., men who have sex with men, sex workers, attendees at STI clinics/services. Also, people living with HIV and pregnant women
14 | Intended users | Trained laboratory scientists, technologists, technicians and specialists with expertise
15 | Level of the health-care system | For use in clinical laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 96%
17 | Clinical specificity | ≥ 98%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Not applicable
20 | Rate of invalidity, error, unreturnable | < 5%
21 | Precision (repeatability/reproducibility) | Not applicable
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Rapid plasma reagin and reagin antibodies to non-treponemal antigens such as cardiolipin can be detected in serum or plasma from people with syphilis with an antigen–antibody reaction known as flocculation. The results are seen macroscopically.
24 | Specimen(s) stability | Specimens can be stored at 2–8 °C for 48 h. For prolonged storage, serum and plasma can be stored at –20 °C. Avoid successive freezing and thawing.
25 | Specimen(s) volume | ≤ 50 µL
26 | Type of result | Qualitative (reactive or non-reactive) or quantitative (reactive or non-reactive)
27 | Time to result | 10 min
28 | End-point stability | Immediately, in the wet stage
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 30 specimens per day per operator
31 | Test limitations | The manufacturer should state any test limitations, i.e., conditions that have not been validated for the test, such as testing population, sample types and users.
32 | Internal quality control | Test control cards containing graded reactivity specimens.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight. Any specific requirement for reagents once opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
38 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 48 h and at –20 °C for > 48 h
39 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 48 h and at –20 °C for > 48 h
40 | In-use stability for specimen collection (temperature and humidity) | At –20 °C for 12 months; avoid repeated freezing and thawing (maximum twice)
41 | Transport stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
42 | Storage stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
43 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight. Any specific requirement for controls once opened should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents at the time of manufacture (months) | At least 12 months after the date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months’ shelf-life remaining when the product arrives at the point of use. See p. 16 of the Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Display parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability / Test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Deionized or distilled water
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | RPR carbon antigen, reactive control, RPR test cards, dropper bottle, 20-gauge dispensing needle, disposable stirrer pipettes
66 | Reagent kit size (number of tests) | ≥ 25 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, if controls (reactive and non reactive) are not provided in the kits, it is desirable to prepare controls.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection is required. For testing, a mechanical rotator set at 100 ± 2 to 5 rpm with a humidity cover, a timer, calibrated micropipettes, a micropipette stand, a refrigerator, a freezer and an autoclave are required. If the room temperature is too high for testing, an air conditioner is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | Potential source of harm classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. Usually stated on the material safety data sheet
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification e.g. simple trash or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of technical support in the country | A local authorized service provider or distributor is recommended.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment. (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subjected to the contract is required.
88 | WHO prequalification (PQ) status | The manual method format of this assay is not within the scope of WHO PQ as of July 2022. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | International standards and certifications | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 8 Auto
Non-treponema rapid plasma reagin (RPR) test |  | The technical specification below is for a particular test to be run on an automated analyser. As automated analysers are usually used to run various tests, selection of the equipment should take into account the other tests performed in the facility and not just that in the technical specification.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51819/ Treponema pallidum reagin antibody IVD, kit, agglutination
3 | EMDN code/ Term name | W0105010399/ SYPHILIS REAGENTS OTHER
4 | Generic name | Non-treponemal rapid plasma reagin test
5 | Specific type or variation (optional) | Qualitative assay
6 | Alternative name/s (optional)
7 | Keywords (optional) | Rapid plasma reagin, RPR, syphilis, particle/charcoal agglutination assay, non-treponemal RPR
Intended use
8 | Detection target | Non-treponemal rapid plasma reagin
9 | Test purpose | To screen for syphilis and monitor treatment effectiveness
10 | Specific disorder / condition | Syphilis, sexually and congenitally transmitted infections
11 | Test format | Particle/charcoal agglutination assay (automated method)
12 | Specimen type | Plasma and serum
13 | Testing population | Sexually active population, including, e.g., men who have sex with men, sex workers, attendees at STI clinics and services. Also, people living with HIV and pregnant women
14 | Intended users | Trained laboratory scientists, technologists, technicians and specialists with expertise
15 | Level of the health-care system | For use in clinical laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 92%
17 | Clinical specificity | ≥ 97%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | ≤1 reagin unit
20 | Rate of invalidity, error, unreturnable | Not applicable
21 | Precision (repeatability/reproducibility) | For reactive samples: Intra-assay precision (repeatability): coefficient of variation ≤ 12% Inter-assay precision (reproducibility): coefficient of variation ≤ 13%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Rapid plasma reagin, reagin antibodies to non-treponemal antigens, such as cardiolipin, can be detected in serum or plasma from people with syphilis by an antigen–antibody reaction known as flocculation.
24 | Specimen(s) stability | Specimens may be stored at 2–8 °C for 5 h. For prolonged storage, serum and plasma may be stored at –20 °C. Avoid successive freezing and thawing.
25 | Specimen(s) volume | ≤ 300 µL
26 | Type of result | Qualitative (reactive or non-reactive)
27 | Time to result | ≤ 15 min
28 | End-point stability | Immediately, in the wet stage
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 1500 specimens per h/day
31 | Test limitations | The manufacturer should state any test limitations, i.e., conditions that have not been validated with the test such as testing population, sample type, user.
32 | Internal quality control | Quality control material should be run as defined by the manufacturer.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | ≥ 4 weeks
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight. Any specific requirement for reagents once opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
38 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 5 days and at –20 °C for > 5 days
39 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at –20 °C for > 72 h
40 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 72 h and at –20 °C for > 72 h. Ideally, a sample can be frozen at –20 °C and thawed once.
41 | Transport stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
42 | Storage stability for controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
43 | On-board stability of controls/calibrators (temperature and humidity) | Ideally, ≥ 4 weeks at 2–8 °C
44 | In-use stability of controls/calibrators (temperature and humidity) | Ideally, ≥ 4 weeks at 2–8 °C
45 | Shelf-life of kit/reagents at the time of manufacture (months) | ≥ 12 months after date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | ≥ 6 months’ shelf-life remaining when product arrives at point of use. See p. 16 of the Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577)
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Depends on the specific brand or model.
48 | Weight of the device (kg) | Depends on the specific brand or model.
49 | Power requirements | 100–240 VAC / 50–60 Hz
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Ideally, uninterrupted power source required
53 | Operating conditions (temperature and humidity) | Operating temperature range, at least 18–30 °C Humidity, ≤ 80%
54 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle-ware
55 | Displayed parameters | Digital display of test results, electronic and printed readout.
56 | Display languages | English or local language
57 | Built-in memory storage capacity | Ideally, ≥ 5000 tests
58 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programmed database to allow information-sharing. Capability for data transmission, e.g., via USB, electronic health records and health programme databases
59 | Open or closed system | Both open and closed systems are available.
60 | Multi-disease testing capability / Test menu | Multi-disease testing and single testing are both possible.
Infrastructure requirements
61 | Water requirements | Deionized water with adequate quality. See the analyser operational manual or instructions for use or package insert of the kit and reagents.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | In general, buffer or diluent, antigen and conjugate
66 | Reagent kit size (number of tests) | ≥ 100 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, it depends on the test/reagents and please see the package insert or instruction for use of the test/reagents. In general, controls, calibrators (if the kit/reagents does not contains) and specimen collection tubes are required.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection is required. For testing, an analyser (if specified by the kits/reagent, the specified one is required), a timer, calibrated micropipettes, a micropipette stand, a refrigerator, a freezer and an autoclave are required. Please see the package insert or instruction for use of the test/reagents. If the room temperature goes over the upper limit of the temperature for operating condition, an air conditioner is required.
69 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. The manufacturer should provide a list of spares.
Documentation
70 | Instrument operator manual | Operating and service manuals in local language(s), including lists of important spares and accessories with their reference numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also contain recommended procedures for disposal and any probable hazards to the environment and/or community.
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | The manufacturer should clearly document potential sources of harm on kit inserts, which should be classified according to the Global Harmonized System, e.g.,. health hazard, physical hazard or environmental hazard. More details will usually be provided on the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g., simple waste disposal, hazardous waste, incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range must be assessed.
77 | Requirements for installation and calibration | The manufacturer is responsible for installation and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
78 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
79 | Warranty | The manufacturer must provide a minimum 1-year warranty from the date of handing over; the date of commencement, duration of the warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty must be specified.
80 | Preventive maintenance | The minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be based on either a fixed period or the number of tests processed by the instrument, which will de dictated by the workload.
81 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed on the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures from information provided by the manufacturer. The manufacturer or supplier will specify the maintenance support they will supply.
82 | Type of service contract (including leasing or reagent rental) | The contract must cover the costs of labour, repair, spare parts, shipping, logistics and training. Possible lending of smaller devices
83 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
84 | Availability of technical support in the country | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
85 | Estimated lifespan | 5 years after the start of use or installation, on the condition that the regular maintenance and inspection specified in the instruction manual has been provided, necessary parts have been replaced and repairs or maintenance measures performed under normal use and operating conditions.
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subjected to the contract is required.
88 | WHO prequalification (PQ) status | The automated method format of this assay is not within the scope of WHO PQ as of July 2022. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 9
Non-treponemal Venereal Diseases Research Laboratory (VDRL) test
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51819/ Treponema pallidum reagin antibody IVD, kit, agglutination
3 | EMDN code/ Term name | W0105010399/ SYPHILIS REAGENTS OTHER
4 | Generic name | Non-treponemal Venereal Diseases Research Laboratory (VDRL) test
5 | Specific type or variation (optional) | Qualitative and quantitative assay
6 | Alternative name/s (optional)
7 | Keywords (optional) | VDRL, syphilis, non-treponemal VDRL
Intended use
8 | Detection target | Reagin antibodies to non-treponemal antigens
9 | Test purpose | To screen for and diagnose syphilis and confirm neurosyphilis
10 | Specific disorder / condition | Syphilis, sexually or congenitally transmitted infections
11 | Test format | Flocculation assay against non-treponemal antigens (manual method)
12 | Specimen type | Serum, plasma, serum and cerebrospinal fluid
13 | Testing population | Sexually active population, including, e.g., men who have sex with men, sex workers, attendees at STI clinics and services. Also, people living with HIV and pregnant women.
14 | Intended users | Trained laboratory scientists, technologists, technicians and specialists with expertise
15 | Level of the health-care system | For use in clinical laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 88%
17 | Clinical specificity | ≥ 98%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Not applicable
20 | Rate of invalidity, error, unreturnable | < 5%
21 | Precision (repeatability/reproducibility) | Not applicable
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Antibodies to non-treponemal antigens, such as cardiolipin, can be detected in serum or plasma from people with syphilis with an antigen–antibody reaction known as flocculation. The results can be seen macroscopically.
24 | Specimen(s) stability | Specimens may be stored at 2–8 °C for 3 days. For prolonged storage, serum and plasma may be stored at –20 °C. Avoid successive freezing and thawing.
25 | Specimen(s) volume | ≤ 50 µL
26 | Type of result | Qualitative (reactive or non-reactive) or quantitative (reactive or non-reactive)
27 | Time to result | 8–10 min
28 | End-point stability | Immediately, in the wet stage
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 30 specimens per day per operator
31 | Test limitations | The manufacturer should state any test limitation, i.e., conditions that have not been validated with the test such as testing population, sample type, user.
32 | Internal quality control | Test control cards containing graded reactivity specimens
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight or At 15–30 °C, protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight or At 15–30 °C, protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C or 15–30° C, protected against humidity and direct sunlight. Any specific requirement for reagents once opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
38 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 3 days and at –20 °C for > 3 days
39 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 3 days and at –20 °C for > 3 days
40 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 3 days and at –20 °C for > 3 days, avoid repeated freezing and thawing (maximum twice)
41 | Transport stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight or At 15–30 °C, protected against humidity and direct sunlight
42 | Storage stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight or At 15–30 °C, protected against humidity and direct sunlight
43 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability of controls/calibrators (temperature and humidity) | At 2–8 °C or at 15–30 °C protected against humidity and direct sunlight. Any specific requirement for controls once opened should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents at time of manufacture (months) | 12 months after date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months’ shelf-life remaining when the product arrives at point of use. See p. 16 of the Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Display parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability / Test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Deionized or distilled water
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes, if the temperature in the storage area of the kits exceeds the storage temperature of the kits
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit components – test / reagents / consumables (if relevant) | VDRL carbon antigen
66 | Reagent kit size (number of tests) | ≥ 100 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, disposable syringes, 1 mL, disposable calibrated needles without bevel. Serum test: 18 gauge; cerebrospinal fluid test: 21 or 22 gauge, 0.9% Saline, tips for 50 μL pipettes and tips for micropipette are required. If controls (reactive and non reactive) are not provided in the kits, it is desirable to prepare controls.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection is required. For testing, bottles, 30 mL, round, narrow-mouthed, 35 mm in diameter with glass stoppers and a flat inner bottom surface (Properly washed 25-mL glass-stoppered Erlenmeyer flasks may be substituted if 30-mL bottles are not available.); micropipette, serological, graduated to tip: 1.0 mL, graduated in 1/100 mL, 5.0 mL, graduated in 1/10 mL, 10.0 mL, graduated in 1/10 mL. Slides: serum test: 25 x 50 mm with paraffin or ceramic rings approximately 14 mm in diameter and high enough to prevent spillage during rotation; cerebrospinal fluid test: Kline concavity slides, 50 × 25 × 3 mm, 12 concavities measuring 16 mm in diameter and 1.75 mm in depth, a slide holder for 50 × 100 mm slides, a mechanical rotator adjustable to 180 ± 2 rpm, circumscribing a circle 19 mm in diameter on a horizontal plane, a water bath, 56 °C, a light microscope with 10× ocular and 10× objective, a timer, a calibrated micropipette, a micropipette stand, pipettes, a refrigerator, a freezer and an autoclave are required. If room temperature is high (for example, ≥30 °C), an air conditioner is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | Potential source of harm classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. This information is usually provided on the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques according to the hazard classification, e.g., simple trash or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of technical support in the country | A local authorized service provider or distributor is recommended .
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subjected to the contract is required.
88 | WHO prequalification (PQ) status | The manual method format of this assay is not within the scope of WHO PQ as of July 2022. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 10
Treponema pallidum hemagglutination (TPHA) test
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51800/ Treponema pallidum total antibody IVD, kit, agglutination
3 | EMDN code/ Term name | W0105010399/ SYPHILIS REAGENTS OTHER
4 | Generic name | T. pallidum hemagglutination test
5 | Specific type or variation (optional) | Qualitative or semi-quantitative assay
6 | Alternative name/s (optional)
7 | Keywords (optional) | T. pallidum hemagglutination assay, TPHA, red cell agglutination assay, treponemal antibody detection, syphilis
Intended use
8 | Detection target | Anti-treponemal antibodies (IgG and IgM)
9 | Test purpose | To confirm syphilis infection and diagnose early and late syphilis infection
10 | Specific disorder / condition | Syphilis, sexually and congenitally transmitted infections
11 | Test format | Hemagglutination assay (manual method)
12 | Specimen type | Serum or serum and plasma
13 | Testing population | Blood donors, sexually active population, including, e.g., men who have sex with men, sex workers, attendees at STI clinics and services. Also, people living with HIV and pregnant women
14 | Intended users | Trained laboratory scientists, technologists, technicians and specialists with expertise
15 | Level of the health-care system | For use in clinical laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 98%
17 | Clinical specificity | ≥ 95%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Not applicable
20 | Rate of invalidity, error, unreturnable | < 5%
21 | Precision (repeatability/reproducibility) | Not applicable
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | The test is a passive hemagglutination assay based on hemagglutination of erythrocytes sensitized with T. pallidum antigen by antibodies found in the patient’s serum or plasma. It is used for both qualitative and semi-quantitative detection of anti-treponemal antibodies.
24 | Specimen(s) stability | Specimens may be stored at 2–8 °C for 24 h. For prolonged storage, serum and plasma may be stored at –20 °C. Avoid successive freezing and thawing.
25 | Specimen(s) volume | 10–100 µL
26 | Type of result | Qualitative or semi-quantitative (reactive, indeterminate or non-reactive)
27 | Time to result | 70–130 min
28 | End-point stability | Not applicable
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 48 specimens per day per operator
31 | Test limitations | The manufacturer should state any test limitations, e.g., conditions that have not been validated with the test, such as testing population, sample type, user.
32 | Internal quality control | Positive control to T. pallidum antibodies and negative control to T. pallidum antibodies
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight. Any specific requirement for reagents once opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
38 | Transport stability for specimen collection (temperature and humidity) | Specimens may be stored at 2–8 °C for 24 h. For prolonged storage, samples may be stored at –20 °C. Avoid successive freezing and thawing.
39 | Storage stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C for > 24 h
40 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C for > 24 h and avoid repeated freezing and thawing
41 | Transport stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
42 | Storage stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
43 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight. Any specific requirement for controls once opened should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents at time of manufacture (months) | At least 12 months after the date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months’ shelf-life remaining when product arrives at point of use. See p. 16 of the Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Display parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability / Test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Deionized or distilled water
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other consumables
65 | Kit component – test / reagents / consumables (if relevant) | Test cells, control cells, sample diluent, positive control and negative control if possible
66 | Reagent kit size (number of tests) | ≥ 100 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, U-shaped microtitration plates, specimen collection receptacles, pipette tips, disinfectants are required.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection is required. For testing, a plate shaker, a plate cover, a centrifuge, pipettes, a timer, calibrated micropipettes capable of delivering: 10, 25, 75 and 190μL, a micropipette stand, a refrigerator, a freezer and an autoclave are required. If room temperature is high(e.g., over 30°C), an air conditioner is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | Potential source of harm. Classification according to the Global Harmonized System, e.g. health hazard, physical hazard or environmental hazard. This is usually contained in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g., simple trash or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of technical support in the country | A local authorized service provider or distributor is recommended.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subjected to the contract is required.
88 | WHO prequalification (PQ) status | The manual method format of this assay is not within the scope of WHO PQ as of July 2022. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 11 Auto
Treponema pallidum hemagglutination (TPHA) test |  | The technical specification below is for a particular test to be run on an automated analyser. As automated analysers are usually used to run various tests, selection of the equipment should take into account the other tests performed in the facility and not just that in the technical specification.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51800/ Treponema pallidum total antibody IVD, kit, agglutination
3 | EMDN code/ Term name | W0105010399/ SYPHILIS REAGENTS OTHER
4 | Generic name | T. pallidum particle agglutination test
5 | Specific type or variation (optional) | Qualitative
6 | Alternative name/s (optional)
7 | Keywords (optional) | T. pallidum particle agglutination test, TPPA, particle agglutination assay, syphilis, treponemal antibody detection
Intended use
8 | Detection target | Anti-treponemal antibodies (IgG and IgM)
9 | Test purpose | To confirm syphilis infection and diagnose early and late syphilis infection
10 | Specific disorder / condition | Syphilis, sexually and congenitally transmitted infections
11 | Test format | Hemagglutination assay (automated method) The intended use of some automated T. pallidum haemagglutination products on the market is for screening blood donations. When countries develop their own EDL, they should take note of the intended use claimed by the manufacturer.
12 | Specimen type | Serum (preferred), plasma
13 | Testing population | Blood donors, sexually active population, e.g., men who have sex with men, sex workers, attended at STI clinics and services Also, people living with HIV and pregnant women.
14 | Intended users | Trained laboratory scientists, technologists, technicians and specialists with expertise
15 | Level of the health-care system | For use in clinical laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 99%
17 | Clinical specificity | ≥ 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Ideally, ≤ 0.1 IU/mL
20 | Rate of invalidity, error, unreturnable | < 5%
21 | Precision (repeatability/reproducibility) | Intra-assay precision (repeatability) with reactive samples, coefficient of variation ≤ 18% Reproducibility (among several sites) with medium or highly reactive samples, coefficient of variation ≤ 28%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Passive haemagglutination assay based on haemagglutination of erythrocytes sensitized with T. pallidum antigen by antibodies in the patient’s serum or plasma. Used for both qualitative and semi-quantitative detection of anti-treponemal antibodies
24 | Specimen(s) stability | Specimens may be stored at 2–8 °C for 7 days. For prolonged storage, may be stored at –20 °C
25 | Specimen(s) volume | ≤ 50 µL
26 | Type of result | Qualitative (positive or negative)
27 | Time to result | ≤ 60 min
28 | End-point stability | Not applicable
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 2400 specimen per 8-h day
31 | Test limitations | The manufacturer should state any limitations.
32 | Internal quality control | Quality control material should be run as defined by the manufacture.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C in an upright position, protected against high humidity and direct sunlight. See manufacturer’s instructions for details.
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C in an upright position, protected against high humidity and direct sunlight. See manufacturer’s instructions for details.
36 | On-board stability of kit/reagents (temperature and humidity) | Ideally, ≥ 12 h
37 | In-use stability of kit/reagents (temperature and humidity) | Ideally, ≥ 72 h at 2–8 °C once opened
38 | Transport stability for specimen collection media (temperature and humidity) | At 2– 8°C for 7 days; at –20 °C for > 7 days
39 | Storage stability for specimen collection media (temperature and humidity) | At 2–8 °C for 72 h; at –20 °C beyond 72 h
40 | In-use stability for specimen collection media (temperature and humidity) | At 2–8 °C for 72 h; at –20 °C beyond 72 h. Avoid freezing and thawing specimens.
41 | Transport stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
42 | Storage stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
43 | On-board stability of controls/calibrators (temperature and humidity) | Ideally, ≥ 1 month
44 | In-use stability of controls/calibrators (temperature and humidity) | Ideally, > 7 days at 2–8 °C
45 | Shelf-life of kit/reagents at time of manufacture (months) | ≥ 12 months after date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months’ shelf-life remaining when product arrives at point of use. See p. 16 of the Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Depends on the specific brand or model.
48 | Weight of the device (kg) | Depends on the specific brand or model.
49 | Power requirements | 100–240 VAC / 50–60 Hz or 200–240 VAC / 50–60 Hz
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Uninterrupted power source required
53 | Operating conditions (temperature and humidity) | Operating temperature range: 18–30 °C Humidity: ≤ 80%
54 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle-ware
55 | Display parameters | Digital display of test results, electronic and printed readout.
56 | Display languages | English or local language
57 | Built-in memory storage capacity | Ideally, 5000 tests
58 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programme database to allow information-sharing. Capability for data transmission via, e.g., USB, electronic health records and health programme databases.
59 | Open or closed system | Closed system
60 | Multi-disease testing capability / Test menu | Multi-disease testing
Infrastructure requirements
61 | Water requirements | Deionized water. See the analyses operational manual or instructions for use of kit/reagents.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | Reagent (test cell suspension) and buffer/diluent
66 | Reagent kit size (number of tests) | ≥ 1000 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, it depends on the test/reagents and please see the package insert or instruction for use of the test/reagents. In general, reactive and non reactive controls, calibrator, microplates (if the kit/reagents does not contains), specimen collection tube, pipette tips are required.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection is required. For testing, an analyser (if specified by the kits/reagent, the specified one is required), a timer, a calibrated micropipette, a micropipette stand, a refrigerator, a freezer and an autoclave are required. Please see the package insert or instruction for use of the test/reagents. If the room temperature goes over the upper limit of the temperature for operating condition, an air conditioner is required.
69 | Spare parts (if relevant) | Basic spares that the customer can replace themselves during maintenance. A list of spares should be provided by the manufacturer.
Documentation
70 | Instrument operator manual | Operating and service manuals in local language(s), including lists of important spares and accessories with their reference numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also contain recommended procedures for disposal and any probable hazards to the environment and/or community.
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | The manufacturer should clearly document potential sources of harm on kit inserts, which should be classified according to the Global Harmonized System, e.g.,. health hazard, physical hazard or environmental hazard. More details will usually be provided on the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g., simple waste disposal, hazardous waste, incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range must be assessed.
77 | Requirements for installation and calibration | The manufacturer is responsible for installation and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
78 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
79 | Warranty | The manufacturer must provide a minimum 1-year warranty from the date of handing over; the date of commencement, duration of the warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty must be specified.
80 | Preventive maintenance | The minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be based on either a fixed period or the number of tests processed by the instrument, which will de dictated by the workload.
81 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed on the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures from information provided by the manufacturer. The manufacturer or supplier will specify the maintenance support they will supply.
82 | Type of service contract (including leasing or reagent rental) | The contract must cover the costs of labour, repair, spare parts, shipping, logistics and training. Possible lending of smaller devices
83 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
84 | Availability of technical support in the country | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
85 | Estimated lifespan | 5 years after the start of use or installation, on the condition that the regular maintenance and inspection specified in the instruction manual has been provided, necessary parts have been replaced and repairs or maintenance measures performed under normal use and operating conditions.
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subjected to the contract is required.
88 | WHO prequalification (PQ) status | The automated method format of this assay is not within the scope of WHO PQ as of July 2022. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 12
Treponema pallidum particle agglutination (TPPA) test
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51800/ Treponema pallidum total antibody IVD, kit, agglutination
3 | EMDN code/ Term name | W0105010399/ SYPHILIS REAGENTS OTHER
4 | Generic name | T. pallidum particle agglutination test
5 | Specific type or variation (optional) | Qualitative or semi-quantitative assay
6 | Alternative name/s (optional)
7 | Keywords (optional) | T. pallidum particle agglutination test, TPPA, particle agglutination assay, syphilis, treponemal antibody detection
Intended use
8 | Detection target | Anti-treponemal antibodies (IgG and IgM)
9 | Test purpose | To confirm syphilis infection and diagnose early and late syphilis infection
10 | Specific disorder / condition | Syphilis, sexually and congenitally transmitted infections
11 | Test format | Particle agglutination assay (manual method)
12 | Specimen type | Serum (preferred), plasma
13 | Testing population | Patients with suspected syphilis include blood donors, sexually active people including men who have sex with men, sex workers, attendees at STI clinic and services. Also, people living with HIV and pregnant women.
14 | Intended users | Trained laboratory scientists, technologists, technicians and specialists with expertise
15 | Level of the health-care system | For use in clinical laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 99%
17 | Clinical specificity | ≥ 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Not applicable
20 | Rate of invalidity, error, unreturnable | < 5%
21 | Precision (repeatability/reproducibility) | Not applicable
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Particle agglutination assay in which gelatin particles coated with purified pathogenic T. pallidum (Nichols strain) are agglutinated in the presence of antibodies to T. pallidum in human serum or plasma
24 | Specimen(s) stability | Specimens may be stored at 2–8 °C for 24 h. For prolonged storage, serum and plasma may be stored at –20 °C. Avoid successive freezing and thawing.
25 | Specimen(s) volume | ≤ 50 µL
26 | Type of result | Qualitative or semi-quantitative assay (positive, negative or indeterminate)
27 | Time to result | 120–180 min
28 | End-point stability | Not applicable
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 48 specimens per day per operator
31 | Test limitations | The manufacturer should state any test limitations, i.e., conditions that have not been validated with the test such as testing population, sample type, user.
32 | Internal quality control | Positive control
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight. Any specific requirement for reagents once opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
38 | Transport stability for specimen collection media (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C beyond 24 h
39 | Storage stability for specimen collection media (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C beyond 24 h
40 | In-use stability for specimen collection media (temperature and humidity) | At 2–8 °C for 24 h and at –20 °C beyond 24 h
41 | Transport stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
42 | Storage stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
43 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight. Any specific requirement for controls once opened should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents at time of manufacture (months) | At least 12 months after date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months’ shelf-life remaining when product arrives at point of use. See p. 16 of the Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument technical and operational characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Display parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability / Test menu | Not applicable
Infrastructure requirements
61 | Water requirements | No
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | Reconstituting solution, sample diluent, sensitized gelatin particles, unsensitized gelatin particles, positive control and droppers
66 | Reagent kit size (number of tests) | ≥ 100 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, “U” shaped microplate, diluters, calibrated pipette droppers, pipette tips, disposable reagent troughs are required.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection is required. For testing, a plate shaker, a plate viewer, a cover for the microplate, calibrated micropipettes,a timer, a micropipette's stand, a refrigerator, a freezer and an autoclave are required. If room temperature is > 35 °C, an air conditioner is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
73 | Material safety data sheet | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | Potential source of harm. Classification according to the Global Harmonized System, e.g. health hazard, physical hazard or environmental hazard. This is usually contained in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g., simple trash or incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of technical support in the country | A local authorized service provider or distributor is recommended.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subjected to the contract is required.
88 | WHO prequalification (PQ) status | The manual method format of this assay is not within the scope of WHO PQ as of July 2022. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 13 Auto
Treponema pallidum particle agglutination (TPPA) test |  | The technical specification below is for a particular test to be run on an automated analyser. As automated analysers are usually used to run various tests, selection of the equipment should take into account the other tests performed in the facility and not just that in the technical specification.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51800/ Treponema pallidum total antibody IVD, kit, agglutination
3 | EMDN code/ Term name | W0105010399/ SYPHILIS REAGENTS OTHER
4 | Generic name | T. pallidum particle agglutination test
5 | Specific type or variation (optional) | Qualitative
6 | Alternative name/s (optional)
7 | Keywords (optional) | T. pallidum particle agglutination test, TPPA, particle agglutination assay, syphilis, treponemal antibody detection
Intended use
8 | Detection target | Anti-treponemal antibodies (IgG and IgM)
9 | Test purpose | To confirm syphilis infection and diagnose early and late syphilis infection
10 | Specific disorder / condition | Syphilis, sexually and congenitally transmitted infections
11 | Test format | Particle agglutination assay (automated method). In general, the intended use of an automated T. pallidum particle agglutination test product available on the market is for screening blood donations. When countries develop their own EDL, they should note the intended use claimed by the manufacturer.
12 | Specimen type | Plasma and serum
13 | Testing population | Patients with suspected syphilis, including blood donors, sexually active people such as men who have sex with men, sex workers and attendees at STI clinics and services). Also, people living with HIV and pregnant women
14 | Intended users | Trained laboratory scientists, technologists, technicians and specialists with expertise
15 | Level of the health-care system | For use in clinical laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 99%
17 | Clinical specificity | ≥ 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | ≤ 0.1 IU/mL
20 | Rate of invalidity, error, unreturnable | No information
21 | Precision (repeatability/reproducibility) | Not applicable
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Particle-agglutination assay in which gelatin particles coated with purified pathogenic T. pallidum (Nichols strain) are agglutinated in the presence of antibodies to T. pallidum in human serum or plasma
24 | Specimen(s) stability | Specimens may be stored at 2–8 °C for 5 days. For prolonged storage, specimens may be stored at –20 °C. Avoid successive freezing and thawing.
25 | Specimen(s) volume | ≤ 20 µL
26 | Type of result | Qualitative (positive or negative)
27 | Time to result | ≤ 60 min
28 | End-point stability | Not applicable
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | 2400 specimens per 8-h day (maximum)
31 | Test limitations | The manufacturer should state any limitations.
32 | Internal quality control | Quality control material should be run as defined by the manufacturer.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of the analysers' measuring unit
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C
38 | Transport stability for specimen collection (temperature and humidity) | At 2–8 °C for 5 days and at –20 °C for > 5 days
39 | Storage stability for specimen collection (temperature and humidity) | At 2 - 8°C for 5 days and at -20°C for more than 5 days
40 | In-use stability for specimen collection (temperature and humidity) | At 2–8 °C for 5 days and at –20 °C for > 5 days. Avoid freezing and thawing specimens.
41 | Transport stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
42 | Storage stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
43 | On-board stability of controls/calibrators (temperature and humidity) | At the temperature of the analysers' measuring unit
44 | In-use stability of controls/calibrators (temperature and humidity) | At 2–8 °C
45 | Shelf-life of kit/reagents at time of manufacture (months) | ≥ 12 months after date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months’ shelf-life remaining when product arrives at the point of use. See p. 16 of the Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument technical and operational characteristics
47 | Size of device (height x width x depth) | Depends on the specific brand or model.
48 | Weight of the device (kg) | Depends on the specific brand or model.
49 | Power requirements | 100–240 VAC (± 10%)/ 50–60 Hz
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Uninterrupted power source required
53 | Operating conditions (temperature and humidity) | Operating temperature range: 18–30 °C. Humidity: ≤ 80%
54 | User interface | Simple test menu, simple PC as provided by manufacturer, option for connection to laboratory computer and middle-ware
55 | Displayed parameters | Digital display of test results, electronic and printed readout
56 | Display languages | English or local language
57 | Built-in memory storage capacity | Ideally, > 5000 tests
58 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programme database to allow information-sharing. Capability for data transmission via, e.g., USB, electronic health records and health programme databases. Electronic linkage to manufacturer for data updates and instrument monitoring
59 | Open or closed system | Closed system
60 | Multi-disease testing capability / Test menu | Analysis of multiple analytes
Infrastructure requirements
61 | Water requirements | Deionized water. See the analyser operational manual or instructions for use of kit/reagents.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | Reconstituting solution, sample diluent and sensitized particles reagent
66 | Reagent kit size (number of tests) | ≥ 1900 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, it depends on the test/reagents and please see the package insert or instruction for use of the test/reagents. In general, reactive and non reactive controls, calibrators, microplates (if the kit/reagents does not contains), specimen collection tubes, pipette tips are required.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette's stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection is required. For testing an analyser (for the kits or reagent, if specified), a timer, pipettes (1.0 mL, 5.0 mL and 10.0 mL), calibrated micropipettes, a micropipette stand, a refrigerator, a freezer and an autoclave are required. Please see the package insert or instruction for use of the test/reagents. If the room temperature goes over the upper limit of the temperature for operating condition, an air conditioner is required.
69 | Spare parts (if relevant) | Basic spares that customers can replace themselves during maintenance. A list of spares should be provided by the manufacturer.
Documentation
70 | Instrument operator manual | Operating and service manuals in local language(s), including lists of important spares and accessories with their reference numbers and list of equipment and procedures required for calibration and routine maintenance. Documentation must also contain recommended procedures for disposal and any probable hazards to the environment and/or community.
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that confirms that a regulated product meets its product specifications. it should contain the actual results obtained during final quality control for lot release for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | The manufacturer should clearly document potential sources of harm on kit inserts, which should be classified according to the Global Harmonized System, e.g.,. health hazard, physical hazard or environmental hazard. More details will usually be provided on the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g., simple waste disposal, hazardous waste, incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range must be assessed.
77 | Requirements for installation and calibration | The manufacturer is responsible for installation and initial setup of the analyser; the analyser should be in good working order, calibrated and quality controlled before being handed over to users.
78 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuing training should be offered if appropriate.
Warranty and maintenance
79 | Warranty | The manufacturer must provide a minimum 1-year warranty from the date of handing over; the date of commencement, duration of the warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty must be specified.
80 | Preventive maintenance | The minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be based on either a fixed period or the number of tests processed by the instrument, which will de dictated by the workload.
81 | Corrective maintenance | Specific requirements for calibration and test procedures must be detailed on the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures from information provided by the manufacturer. The manufacturer or supplier will specify the maintenance support they will supply.
82 | Type of service contract (including leasing or reagent rental) | The contract must cover the costs of labour, repair, spare parts, shipping, logistics and training. Possible lending of smaller devices
83 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
84 | Availability of technical support in the country | Required: The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
85 | Estimated lifespan | 5 years after the start of use or installation, on the condition that the regular maintenance and inspection specified in the instruction manual has been provided, necessary parts have been replaced and repairs or maintenance measures performed under normal use and operating conditions.
Quality and registration
86 | Global regulatory approvals | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements for the version to be procured and subjected to the contract is required.
88 | WHO prequalification (PQ) status | The automated method format of this assay was not within the scope of WHO PQ as of July 2022. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 14
Antibodies to Treponema pallidum
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51798/ Treponema pallidum total antibody IVD, kit, enzyme immunoassay (EIA)
3 | EMDN code/ Term name | W0105010302/ SYPHILIS ANTIBODY ASSAYS TOTA
4 | Generic name | Antibodies to Treponema pallidum, immunoassay
5 | Specific type or variation (optional) | Qualitative assay
6 | Alternative name/s (optional)
7 | Keywords (optional) | T. pallidum, syphilis, antibodies to T. pallidum, microtitre plate enzyme immunoassay
Intended use
8 | Detection target | Anti-treponemal antibodies
9 | Test purpose | To diagnosis or to aid in the diagnosis of syphilis
10 | Specific disorder / condition | Sexually, blood transfusion- or congenitally transmitted infections
11 | Test format | Immunoassay (manual method)
12 | Specimen type | Plasma and serum
13 | Testing population | Adults, including pregnant women
14 | Intended users | Trained laboratory scientists, technologists, technicians and specialists with expertise
15 | Level of the health-care system | For use in blood screening laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 98%
17 | Clinical specificity | ≥ 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Not applicable
20 | Rate of invalidity, error and unreturnable | < 5%
21 | Precision (repeatability/reproducibility) | With reactive specimens: Intra-assay precision (repeatability): coefficient of variation ≤ 10% inter-assay precision (reproducibility): coefficient of variation ≤ 20%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Enzyme immunoassay, microwells coated with T. pallidum antigen, specific for T. pallidum
24 | Specimen(s) stability | Specimens can be stored at 2–8 °C for 2 days. For prolonged storage, plasma and serum specimens can be stored at – 20 °C.
25 | Specimen(s) volume | ≤ 100 µL
26 | Type of result | Qualitative (reactive or non-reactive)
27 | Time to result | 3–4 h
28 | End-point stability | Fixed reading time. Follow the manufacturer's instructions on use or package insert.
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | ≥ 40 specimens per day per operator
31 | Test limitations | The manufacturer should state any limitations.
32 | Internal quality control | Antibodies to T. pallidum, positive control, negative control and blank control
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
36 | On-board stability of kit/reagents (temperature and humidity) | Not applicable
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C protected against humidity and direct sunlight. Any specific requirement once reagents are opened should be clearly stated.
38 | Transport stability for specimen collection (temperature and humidity) | Whole blood specimen at 2–8 °C for 8 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 48 h Plasma and serum at – 20 °C for > 48 h Avoid multiple freeze–thaw cycles.
39 | Storage stability for specimen collection (temperature and humidity) | Whole blood specimen at 2–8 °C for 8 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 48 h Plasma and serum at – 20 °C for > 48 h Avoid multiple freeze–thaw cycles.
40 | In-use stability for specimen collection (temperature and humidity) | Whole blood specimen at 2–8 °C for 8 h (Do not freeze whole blood samples.) Plasma and serum at 2–8 °C for 48 h Plasma and serum at – 20 °C for more 48 h Avoid multiple freeze–thaw cycles.
41 | Transport stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
42 | Storage stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight
43 | On-board stability of controls/calibrators (temperature and humidity) | Not applicable
44 | In-use stability of controls/calibrators (temperature and humidity) | At 2–8 °C, protected against humidity and direct sunlight. Any specific requirement once controls are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents at the time of manufacture (months) | ≥ 12 months for kits/reagents and 6 months for controls/calibrators from the date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months’ shelf-life remaining when product arrives at point of use. See p. 16 of Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Not applicable
48 | Weight of the device (kg) | Not applicable
49 | Power requirements | Not applicable
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Not applicable
53 | Operating conditions (temperature and humidity) | Not applicable
54 | User interface | Not applicable
55 | Display parameters | Not applicable
56 | Display languages | Not applicable
57 | Built-in memory storage capacity | Not applicable
58 | Diagnostic connectivity | Not applicable
59 | Open or closed system | Not applicable
60 | Multi-disease testing capability / Test menu | Not applicable
Infrastructure requirements
61 | Water requirements | Distilled or deionized water with adequate quality for enzyme immunoassays and enzyme-linked immunosorbent assays. See the instruction for use or package insert of kit or reagents.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | Coated microwells or coated microwell strips, calibrator, substrate, positive and negative controls, conjugate, wash buffer and stop solution
66 | Reagent kit size (number of tests) | ≥ 96 tests
67 | Consumables required but not provided in the test kit | For venous whole blood collection, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips, containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, absorbent paper tissues and some reagents that the kits do not contain are required. See the package insert or instruction for use of the kit or reagent.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette's stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection is required. For testing, a microwell plate or strip reader, an incubator or water bath, a heating block, a microwell plate or strip washer, a vacuum pump, a vortex, calibrated single and multi-channel micropipettes, an electronic scale with hoods and a cover for calibrating the micropipettes, a micropipette stand, a timer, a cover for plate, a refrigerator, a freezer and an autoclave are required. If the room temperature exceeds the upper limit of the operating temperature, an air conditioner is required.
69 | Spare parts (if relevant) | Not applicable
Documentation
70 | Instrument operator manual | Not applicable
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | Required. A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | Potential source of harm classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. Usually stated on the material safety data sheet
75 | Disposal requirements | Description of disposal techniques available according to the hazard classification, e.g., simple trash or incineration.
Training, installation and use
76 | Pre-installation requirements (if relevant) | Not applicable
77 | Requirements for installation and calibration | Not applicable
78 | Training of user/s (if relevant) | Not applicable
Warranty and maintenance
79 | Warranty | Not applicable
80 | Preventive maintenance | Not applicable
81 | Corrective maintenance | Not applicable
82 | Type of service contract (including leasing or reagent rental) | Not applicable
83 | Availability of spare parts post-warranty | Not applicable
84 | Availability of technical support in the country | A local authorized service provider or distributor is recommended.
Decommissioning
85 | Estimated lifespan | Not applicable
Quality and registration
86 | Global regulatory approval | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
88 | WHO prequalification (PQ) status | The manual method format of this assay is not within the scope of WHO PQ as of July 2022. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## 15 Auto
Antibodies to Treponema pallidum |  | The technical specification below is for a particular test to be run on an automated analyser. As automated analysers are usually used to run various tests, selection of the equipment should take into account the other tests performed in the facility and not just that in the technical specification.
IVD technical specifications |  | Description and examples
i | Version No. | One (1)
ii | Date of initial version | 2020-01-30
iii | Date of last modification | 2025-01-23
iv | Date of publication
v | Prepared by | World Health Organization
Name, category and coding
1 | WHO category / MTE code
2 | GMDN code/ Term name | 51798/ Treponema pallidum total antibody IVD, kit, enzyme immunoassay (EIA)
3 | EMDN code/ Term name | W0105010302/ SYPHILIS ANTIBODY ASSAYS TOTAL
4 | Generic name | Antibodies to Treponema pallidum, immunoassay
5 | Specific type or variation (optional) | Qualitative
6 | Alternative name/s (optional)
7 | Keywords (optional) | T. pallidum, syphilis, antibodies to T. pallidum, chemiluminescence immunoassay (CLIA)
Intended use
8 | Detection target | Anti-treponemal antibodies
9 | Test purpose | To diagnose or to aid in the diagnosis of syphilis
10 | Specific disorder / condition | Sexually, blood transfusion- or congenitally transmitted infections
11 | Test format | Immunoassay (automated method) The intended use for blood screening of the products for this assay with automated method may vary depending on the regulatory authorities. When countries develop their own EDL, they should take note of the intended use claimed by the manufacturer.
12 | Specimen type | Plasma and serum
13 | Testing population | Adults, including pregnant women
14 | Intended users | Trained laboratory scientists, technologists, technicians and specialists with expertise
15 | Level of the health-care system | For use in blood screening laboratories
Performance characteristics
16 | Clinical sensitivity | ≥ 99%
17 | Clinical specificity | ≥ 99%
18 | Analytical specificity | Assay does not cross-react with other human diseases relevant to the local setting or common interfering substances, which are described in technical specifications for submission for WHO prequalification – Diagnostic assessment: Syphilis rapid diagnostic test (WHO TSS-6 (https://apps.who.int/iris/bitstream/handle10665/277285/9789241515160-eng.pdf?us=1). If some cross-reactivity is present, manufacturers must state this as a limitation of their product, and a decision to use the test is made on the basis that any interference observed will not have a significant impact on the performance of the product.
19 | Limit of detection | Not applicable
20 | Rate of invalidity, error or unreturnable | Not applicable
21 | Precision (repeatability/reproducibility) | With reactive samples: intra-assay precision (repeatability): coefficient of variation ≤ 5% inter-assay precision (reproducibility): coefficient of variation ≤ 6%
22 | Trueness of measurement: bias | Not applicable
Technical and operational characteristics
23 | Principle of the assay | Chemiluminescent immunoassay
24 | Specimen(s) stability | Specimens can be stored at 2–8 °C for 4 days. For prolonged storage, specimens can be stored at –10 °C or colder.
25 | Specimen(s) volume | ≤ 150 µL
26 | Type of result | Qualitative
27 | Time to result | ≤ 30 min
28 | End-point stability | Not applicable
29 | Ease of use for point-of-care tests only: number of steps that require precision | Not applicable
30 | Specimen throughput per operator, per hour, per 8-h working day or per batched run | In a laboratory with a low sample volume: Specimens per day per analyser : ≤ 500 (Sample throughput per hour per analyser: ≤ 62) In a laboratory with a medium sample volume: Specimens per day per analyser: 500–1000 (Sample throughput per hour: 63–125) In a laboratory with a high sample volume: Specimens per day per analyser : ≥ 1000 (Sample throughput per hour : ≥ 126 )
31 | Test limitations | The manufacturer should state any limitations.
32 | Internal quality control | Quality control material should be run as defined by the manufacturer.
33 | Compatibility with external quality control material | Testing of known positive and known negative specimens should be possible as a form of external control. The manufacturer should state any external quality controls with which the IVD is not compatible. Ideally, each user should participate in an external quality proficiency programme.
34 | Transport stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected against humidity and all heat and light sources
35 | Storage stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected against humidity and all heat and light sources
36 | On-board stability of kit/reagents (temperature and humidity) | At the temperature of the analyser's measuring unit
37 | In-use stability of kit/reagents (temperature and humidity) | At 2–8 °C, in an upright position, protected against humidity and all heat and light sources. Any specific requirement for reagents once they are opened should be clearly stated. The instructions should clearly indicate the duration of stability of each reagent once opened.
38 | Transport stability of specimen (temperature and humidity) | At 2–8 °C for 4 days and at –20 °C or lower for > 48 h
39 | Storage stability of specimen (temperature and humidity) | At 2–8 °C for 4 days. At –20 °C or lower for > 4 days
40 | In-use stability of specimen | At 2–8 °C for 48 h and at –10 °C or lower for > 48 h
41 | Transport stability of controls/calibrators (temperature and humidity) | At 2–8 °C, in an upright position, protected against humidity and all heat and light sources
42 | Storage stability of controls/calibrators (temperature and humidity) | At 2–8 °C, in an upright position, protected against humidity and all heat and light sources
43 | On-board stability of controls/calibrators (temperature and humidity) | At the temperature of the analyser's measuring unit
44 | In-use stability of controls/calibrators (temperature and humidity) | At 2–8 °C, in an upright position, protected against humidity and all heat and light sources. Any specific requirement for controls or calibrators once opened should be clearly stated. The instructions should clearly indicate the duration of stability of each control once opened.
45 | Shelf-life of kit/reagents at time of manufacture (months) | At least 9 months for kits/reagents and 6 months for controls/calibrators from the date of manufacture
46 | Remaining shelf-life of kit/reagents upon delivery (months) | At least 6 months’ shelf-life remaining when the product arrives at the point of use. See p. 16 of the Guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/handle/10665/255577).
Instrument physical and technical characteristics
47 | Size of device (height x width x depth) | Depends on the specific brand or model and whether bench-top or floor-standing
48 | Weight of the device (kg) | Depends on the specific brand or model. Many analysers weigh 150–800 kg.
49 | Power requirements | 100–240 VAC / 50–60 Hz or 200–240 VAC / 50–60 Hz
50 | Time for battery to charge | Not applicable
51 | Battery duration | Not applicable
52 | Alternative charging options | Ideally, an uninterrupted power source
53 | Operating conditions (temperature and humidity) | Temperature range: 18–30 °C. Humidity: < 80%. Altitude: ≤ 2000 m
54 | User interface | Simple test menu, simple PC as provided by the manufacturer, option for connection to laboratory computer and middle-ware
55 | Display parameters | Digital display of test results, electronic and printed readout
56 | Display languages | English or local language
57 | Built-in memory storage capacity | Analyser for low-volume sample laboratory: ≥ 800 tests Analyser for medium-volume sample laboratory: ≥ 2,000 tests Analyser for high-volume sample laboratory: ≥ 10,000 tests
58 | Diagnostic connectivity | Data connectivity compatible with laboratory information system and programmed database to allow information-sharing. Capability for data transmission, e.g., via USB, electronic health records and health programme databases
59 | Open or closed system | Both open and closed systems available
60 | Multi-disease testing capability / Test menu | Multi-disease testing
Infrastructure requirements
61 | Water requirements | Distilled or deionized water with adequate quality for chemiluminescence immunoassays. See the instructions for use or package insert of kit or reagents.
62 | Electricity requirements | Yes
63 | Refrigeration or cold chain for storage of kit/reagents | Yes
64 | Refrigeration or cold chain for reconstituted reagents and controls | Yes
Accessories, consumables, spare parts and other components
65 | Kit component – test / reagents / consumables (if relevant) | In general, buffer/diluent, antigen and conjugate.
66 | Reagent kit size (number of tests) | ≥ 40 tests
67 | Consumables required but not provided in the test kit | For collecting venous whole blood, blood collection tubes and needles are required. For preparing plasma or serum from venous whole blood, pipette tips and containers for plasma or serum are required. Common consumables for blood collection (alcohol swabs, tissues, permanent markers/pens and biohazard bags and disposable gloves) are required. For testing, bleach or sodium hypochlorite, ethanol or denatured ethanol and pipette tips are required.
68 | Other auxiliary laboratory equipment required but not provided | For venous whole blood collection, a tourniquet and a blood collection tube stand are required. For serum, a centrifuge is required. For plasma or serum, a micropipette, a micropipette stand and a stand for containers of plasma or serum are required. Common equipment (such as a container for used needles or lancets) for venous whole blood collection is required. For testing, an analyser (if specified by the kits/reagent, the specified one is required), a timer, calibrated single and multi-channel micropipettes, a micropipette stand, a refrigerator, a freezer and an autoclave are required. Please see the package insert or instruction for use of the test/reagents. If the room temperature goes over the upper limit of the temperature for operating condition, an aircon is required.
69 | Spare parts (if relevant) | Basic spare parts that customers can replace themselves during maintenance. The manufacturer should supply a list of spare parts.
Documentation
70 | Instrument operator manual | Operating and service manuals in the local language, including lists of important spares and accessories with their reference numbers and a list of equipment and procedures required for calibration and routine maintenance. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
71 | Instructions for use | Instructions for use in the local language. Must conform to the regulatory version registered for sale and use in the country of supply.
72 | Certificate of analysis | A certificate of analysis is a document issued by the quality assurance department of the legal manufacturer that a regulated product meets its product specifications. It should contain the actual results obtained from final quality control for lot release performed for each lot of a product.
73 | Material safety data sheet (MSDS) | Required. The manufacturer should supply detailed documentation for each method, with comprehensive information on (i) the health effects of exposure to the product; (ii) hazard evaluation for handling, storage, use and disposal; (iii) measures to protect workers at risk of exposure; and (iv) emergency procedures.
Environmental and safety requirements
74 | Hazard classification | The manufacturer should clearly document potential sources of harm on kit inserts. Harm should be classified according to the Global Harmonized System, e.g., health hazard, physical hazard or environmental hazard. More details will usually be provided in the material safety data sheet.
75 | Disposal requirements | Description of disposal techniques available and/or required according to the hazard classification, e.g., simple waste disposal, hazardous waste, incineration
Training, installation and use
76 | Pre-installation requirements (if relevant) | Access, space, water, power, waste, refrigeration and room temperature range must be assessed.
77 | Requirements for installation and calibration | The manufacturer should be responsible for installation and initial setup of the analyser. The analyser should be in good working order, calibrated and quality-controlled before being handed over to users.
78 | Training of user/s (if relevant) | Manufacturers should provide detailed training in use and maintenance of the analyser before handing over the system. Continuous training should be offered if appropriate.
Warranty and maintenance
79 | Warranty | The manufacturer must provide a minimum 1-year warranty from the date of handover and specify the date of commencement, duration of warranty period, exclusions/inclusions and other conditions such as maintenance support during the warranty.
80 | Preventive maintenance | The minimal maintenance to be conducted by the user and by the supplier should be stated. The frequency of maintenance should be either a fixed period or according to the number of tests processed, which will be dictated by the workload.
81 | Corrective maintenance | Specific requirements for calibration and test procedures must be described in the kit inserts. Advanced maintenance tasks must be documented in standard operating procedures based on information provided by the manufacturer. The manufacturer or the supplier will specify the maintenance support they will supply.
82 | Type of service contract (including leasing or reagent rental) | Must cover costs of labour, repair, spare parts, shipping, logistics and training. Possible lending of small instruments
83 | Availability of spare parts post-warranty | Spare parts should be available for the life of the system, which should be a minimum of 5 years.
84 | Availability of technical support in the country | Required. The manufacturer or agreed distributor must provide adequate maintenance and repair support for the clinical requirements of the user.
Decommissioning
85 | Estimated lifespan | 5 years after the start of use/installation under the conditions of regular maintenance / inspection specified in the instruction manual, replacement of parts and necessary repairs or maintenance under normal use and operating conditions
Quality and registration
86 | Global regulatory approval | The supplier should state any regulatory approval of the product. For a product to be considered, it should have marketing authorization for supply in the jurisdiction of a well-established medical devices regulator and comply with the corresponding pre-market requirements as described in the WHO guidance for procurement of in vitro diagnostics and related laboratory items and equipment (https://apps.who.int/iris/bitstream/handle/10665/255577/9789241512558-eng.pdf?sequence=1&isAllowed=y). Well-recognized regulatory authorities include the five founding members of the International Medical Devices Regulatory Forum (previously the Global Harmonization Task Force): Australian Therapeutic Goods Administration, European Commission, Health Canada, Japan's Ministry of Health, Labour and Welfare and Pharmaceutical and Medical Devices Agency and US Food and Drug Administration.
87 | Regulatory version to be procured | Valid certification (including national registration) of compliance with regulatory requirements of the version to be procured and subject to the contract is required.
88 | WHO prequalification (PQ) status | The automated method format of this assay is not within the scope of WHO PQ as of July 2022. Prequalification is recommended If the test category is within the scope of WHO PQ. See WHO public reports for in vitro diagnostics (https://extranet.who.int/pqweb/vitro-diagnostics/prequalification-reports/whopr).
89 | Post-market surveillance | Provision of a summary of current post-market information by an applicant could indicate product performance and is a useful basis for a procurement decision. Whether the supplier is a manufacturer, an authorized representative or a distributor, each has an important role in ensuring timely feedback of complaints, identification of non-conformity and subsequent recall, correction or corrective actions. Contractual agreements should require the supplier to have procedures and agreements to support post-market activities in accordance with their role in a post-market system.
90 | Field safety collective action | It is recommended that manufacturers and providers have procedures for notifying any corrective action to final users. Such procedures should be declared in the contract.

## Acknowledgement and Purpose
Purpose and Aim
The WHO Model List of Essential In Vitro Diagnostics (EDL) is a list of recommended types of tests to support diagnosis of infectious diseases and non-communicable diseases at each level of the healthcare system. The EDL serves as an evidence-based reference point for countries to develop their own national lists to guide how they choose and use IVDs. The EDL is also published in an electronic format as eEDL. However, the EDL and national EDLs (NEDL) do not provide detailed specifications of in vitro diagnostics (IVDs) that are needed to guide procurement of specific IVDs. There is thus a need to develop detailed technical specification to help countries procure appropriate tests. Technical specifications for in vitro diagnostics (IVDs) constitute a set of pre-defined criteria and baseline requirements to ensure good quality, safety, performance and efficacy. Technical specifications for EDL will be also published in an electric format. The purpose of technical specifications for EDL is to help advance universal health coverage (UHC), address health emergencies and promote healthier populations by helping countries to select and procure appropriate IVD products. The aim of technical specifications for EDL is to provide detailed technical information to support decision-making regarding procurement of good quality and safe in vitro diagnostics to improve access to in vitro diagnostic testing in Low and Middle Income Countries (LMICs) in particular.
Acknowledgement list
Technical contribution: Nikhil Moorchung (WHO consultant, India), Emma English (WHO consultant, UK), Jitendra Kumar (WHO consultant, India), Olga Perovic (WHO consultant, South Africa), Ehab Saad (WHO consultant, South Africa), Francis Gabriel Moussy (WHO Headquarters), Ana Aceves Capri (WHO Headquarters), Naofumi Hashimoto (WHO Headquarters), Adriana Velazquez Berumen (WHO Headquarters), Teodora Wi (WHO Headquarters), Ute Ströher (WHO Headquarters), Robert Luo (WHO consultant, USA), Magdalena Barr-DiChiara (WHO consultant, Switzerland).
COI and DOI
All WHO consultants involved in these technical specifications provided DOIs and no conflicts were found.
Funder
Funding and in-kind contributions to this project were provided by the following organizations: • The UK Foreign, Commonwealth & Development Office (FCDO) • The Economic Research Institute for ASEAN and East Asia (ERIA)
Updating
Regular updating of the document will be conducted based on the availability of new knowledge and at least every 3 years.

## Process
Process
Technical specifications define the minimum requirements for the products to ensure good quality, safety and efficacy. The process to develop technical specifications consists of 6 steps as described below. •Step1: Review existing WHO guidance and specifications that support the procurement, donation or lease of good quality IVDs in order to develop a common template of generic specifications. The guidance documents analyzed were: Guidance for procurement of in vitro diagnostics and related laboratory items and equipment, Laboratory-based evaluation of rapid syphilis diagnostics: results from 8 SDI sites, Laboratory diagnosis of sexually transmitted infections, including human immunodeficiency virus, technical specifications for submission for WHO prequalification-Diagnostic assessment: TSS-6 Syphilis rapid diagnostic test, WHO EDL and WHO internal procurement templates. •Step2: WHO commissions an external team of consultants to develop technical specifications for the IVDs listed in the WHO-EDL. The process to develop these specifications includes: horizon scanning of existing tests and platforms; analysis of existing products in the market based on approval from stringent regulatory authorities; analysis of international, regional and country standards; review of WHO guidance and guidelines for all diseases included in the EDL, review of peer-reviewed publications, textbooks, and published qualitative and quantitative comparative studies of existing tests and platforms; comparative study of test specifications (i.e. accuracy, precision, positive & negative predictive value, likelihood ratios) from catalogues and previous tenders of UN agencies and Global Fund with generic formulation; comparison and review of the package inserts and user manuals of the products which were approved by stringent regulatory authorities; comparison and review of documents and data on the safety and performance of the approved products which were uploaded in the websites of stringent regulatory authorities; External Quality Assessment (EQA) data and technical expert opinion. Each template is drafted/reviewed by at least two qualified technical experts. All consultants provided Conflict of Interest declarations and no conflicts were found. •Step3: Internal review of all the draft technical specifications by the EDL secretariat. •Step4: Review of the draft technical specifications by the relevant disease programmes. •Step5: 3 weeks public consultation of the draft technical consultations (through posting on the WHO web site, sharing with relevant networks and Non States Actors (NSA) in official relations with WHO such as International Federation of Biomedical Laboratory Science (IFBLS). •Step6: Compile all feedback and finalize the technical specifications before adding to the electronic EDL.

## Evidence and Input
Evidence and Input
Sources for evidence: 1.Sexually Transmitted Diseases Diagnostics Initiative & UNDP/World Bank/WHO Special Programme for Research and Training in Tropical Diseases. (‎2003)‎. Laboratory-based evaluation of rapid syphilis diagnostics : results from 8 SDI sites. World Health Organization. https://apps.who.int/iris/handle/10665/67993 2.Unemo, Magnus, Ballard, Ronald, Ison, Catherine, Lewis, David, Ndowa, Francis. et al. (‎2013)‎. Laboratory diagnosis of sexually transmitted infections, including human immunodeficiency virus‎. World Health Organization. https://apps.who.int/iris/handle/10665/85343 3. World Health Organization. (‎2017)‎. WHO information note on the use of dual HIV/Syphilis Rapid Diagnostic Tests (‎RDT)‎. World Health Organization. https://apps.who.int/iris/handle/10665/252849 4.Technical specifications for submission for WHO prequalification -Diagnostic assessment: TSS-6 Syphilis rapid diagnostic test, World Health Organization, Geneva, 2018 https://apps.who.int/iris/bitstream/handle/10665/277285/9789241515160-eng.pdf?ua=1 5. World Health Organization. (‎2017)‎. Guidance for procurement of in vitro diagnostics and related laboratory items and equipment. World Health Organization. https://apps.who.int/iris/handle/10665/255577. 6.Ina U. Park, Anthony Tran, Lara Pereira and Yetunde Fakile. Sensitivity and Specificity of Treponemal-specific Tests for the Diagnosis of Syphilis. Clinical Infectious Diseases, cid 2020:71 (Suppl 1) S13-S20 7.Cadamuro J, Lippi G, von Meyer A, Ibarz M, van Dongen E; Lases5, Cornes M, Nybo M, Vermeersch P, Grankvist K, Guimaraes JT, Kristensen GBB, de la Salle B, Simundic AM. European survey on preanalytical sample handling - Part 2: Practices of European laboratories on monitoring and processing haemolytic, icteric and lipemic samples. On behalf of the European Federation of Clinical Chemistry and Laboratory Medicine (EFLM) Working Group for the Preanalytical Phase (WG-PRE). Biochem Med (Zagreb). 2019 Jun 15;29(2):020705. doi: 10.11613/BM.2019.020705. 8.Current ISO standards: ISO 15189:2012, https://www.iso.org/standard/56115.html 9.Kuan, C., Chen, C., Cheng, C. (2015). In-Vitro Diagnostic Devices: Introduction to Current Point-of-Care Diagnostic Devices. Germany: Springer International Publishing. 10.Ogrodnik, P. J. (2014). In Vitro Diagnostic devices: Class 1 Sterile and Measuring Devices. Netherlands: Elsevier Science. 11.Guidance for industry: sterile drug products produced by aseptic processing, current good manufacturing practice. (2004). United States: U.S. Department of Health and Human Services, Food and Drug Administration, Center for Drug Evaluation and Research. 12.FDA Regulatory Affairs: A Guide for Prescription Drugs, Medical Devices, and Biologics. (2008). United Kingdom: Taylor & Francis. 13.Managing Medical Devices Within a Regulatory Framework. (2016). Netherlands: Elsevier Science. 14.Tobin, J. J., Walsh, G. (2011). Medical Product Regulatory Affairs: Pharmaceuticals, Diagnostics, Medical Devices. Germany: Wiley. 15.Handbook of Medical Device Regulatory Affairs in Asia: Second Edition. (2018). Singapore: Jenny Stanford Publishing. 16.FDA Regulatory Affairs: Third Edition. (2014). United States: Taylor & Francis.Weinberg, S. (2009). Guidebook for Drug Regulatory Submissions. United Kingdom: Wiley.