## Foglio1
Cryostat for intraoperative frozen sections
i | Version No. | 1
ii | Date of initial version | December 2019
iii | Date of last modification
iv | Date of publication | February 2020
v | Completed / submitted by | WHO
Name, category and coding
1 | WHO Category / Code | Cryosurgery units, XD9LJ9
2 | Generic name | Cryosurgical Unit, Gynaecological
3 | Specific type or variation (optional)
4 | GMDN name
5 | GMDN code
6 | GMDN category
7 | UMDNS name
8 | UMDNS code
9 | UNSPS code (optional)
10 | Alternative name/s (optional) | Cryoprobes; Cryostats; Cryo Units; CSU; Probes,
11 | Alternative code/s (optional)
12 | Keywords (optional) | surgical, gynaecology, dermatology, cryotherapy, cryocautery, cryosurgery, “cryo”
13 | GMDN/UMDNS definition (optional)
Purpose of use
14 | Clinical or other purpose | Cryosurgical unit with a gas (carbon dioxide,CO2, or nitrous oxide, N2O) cooled or electrically-cooled cryoprobe intended to destroy tissue of the uterine cervix with the application of extreme cold.
15 | Level of use (if relevant) | Health posts, health centres, specialty clinics, hospitals
16 | Clinical department/ ward(if relevant) | Family medicine; gynecology; outpatient clinic; outreach; oncology; obstetrics; surgery; nursing services
17 | Overview of functional requirements | The metal probe head (cryotip) makes contact with epithelium to destroy tissue.
Technical characteristics
18 | Detailed requirements | • Cryosystems include the hand-held unit with use-specific cryotips: • Temperature at probe edge shall be no greater than -20 °C (can be colder, NOT warmer) • If gas-cooled: - either carbon dioxide (CO2) or nitrous oxide (N2O) is used - must be a closed system, whereby the cryogen circulates through probe head, and then back through shaft. • If electrically-cooled, an ethanol-based solution and electricity are necessary.
19 | Displayed parameters | N.A.
20 | User adjustable settings | Pressure gauge (for gas-cooled units), cryogun trigger
Physical/chemical characteristics
21 | Components(if relevant) | • Minimum of 2 cryotips are required: - One cryotip must be flat. - The second cryotip can be either flat or can have a gentle nipple extrusion not exceeding 5mm (so as to anchor in centre of cervix but not to ablate endocervix). - Cryotip diameters ranging from 17 mm to 23 mm - Biocompatible, material that will not adhere to cervix - Reusable and thus decontaminable • The overall length of the cryoshaft and cryotip assembly should be between 170 and 200 mm. Gas tank, pressure gauge, handheld unit with probe, scavenging/suction system if using N2O
22 | Mobility, portability(if relevant) | Can be portable or handheld.
23 | Raw Materials(if relevant) | N.A.
Utility requirements
24 | Electrical, water and/or gas supply (if relevant) | For gas-based: N2O or CO2 The N2O units should only be used with scavenging ability. For electrically cooled: • The unit is suggested to be connected to a continuous, reliable power source • Electrical source requirements (based on country/setting of use): o Amperage: ______; o Voltage: ______; o Plug type: ______
Accessories, consumables, spare parts, other components
25 | Accessories (if relevant) | • Speculum • Light source • Timer
26 | Sterilization process for accessories (if relevant) | - Probes should withstand repeated cycles of decontamination, done so according to manufacturers’ instructions. - Device handle, charging base or gas tank and accessories (whichever applicable) to cleaned and decontaminated after each use according to manufacturer's instructions.
27 | Consumables / reagents (if relevant) | Gas-based: N2O or CO2 gas Electrically cooled: ethanol-based solution
28 | Spare parts (if relevant) | Consider extra probes to account for decontamination cycles and patient scheduling. Note: probes and spares are not interchangeable between devices of different brands and models and can vary in their design and lifetime.
29 | Other components (if relevant) | N.A.
Packaging
30 | Sterility status on delivery (if relevant) | N.A.
31 | Shelf life (if relevant) | N.A.
32 | Transportation and storage (if relevant) | Storage area should be clean and dust-free, dry, cool, well-lit, ventilated and vermin-proof. Store device in original packaging on a shelf or on in a storage cabinet. Devices should be able to withstand storage temperatures ranging from 15°C to 30°C, relative humidity ≤ 85% for gas-based systems or ≤ 60% for electrically cooled systems (both non-condensing), and be protected from dripping water. For compressed gases in cylinders (for gas-based system), both storage and transport can pose as a risk. Transport and storage shall adhere to local regulations; however, gases should never be stored at temparatures in excess of 30°C.
33 | Labelling (if relevant) | Labelling on the primary packaging should include the name and/or trademark of the manufacturer and should adhere to the most current version of ISO 15223 – 1: Medical devices -- Symbols to be used with medical device labels, labelling and information to be supplied -- Part 1: General requirements.
Environmental requirements
34 | Context-dependent requirements | Environmental conditions vary globally and can be extreme; however, the following are tenable: - Operating temperature: 15°C to 35°C - Operating relative humidity: ≤ 85%, non-condensing It is the responsibility of the procurement body to ensure that the manufacturer’s recommended operation and storage conditions are respected. If the device requires that the operating and/or storage environment be climate-controlled, appropriate temperature and humidity control systems, including monitoring, should be applied to avoid premature material disintegration and/or device failure.
Training, installation and utilization
35 | Pre-installation requirements(if relevant) | N.A.
36 | Requirements for commissioning (if relevant) | Supply of either a gas (N2O or CO2) or ethanol-based solution a requesite for operations
37 | Training of user/s (if relevant) | Training of users in operation and basic maintenance shall be provided. Clinical staff training in cervical precancer lesion treatment guidelines and device use to be provided.
38 | User care(if relevant) | • The cryosurgical unit shall be decontaminated between patients. • Probes should be inspected regularly for mechanical integrity.
Warranty and maintenance
39 | Warranty | Minimum one year. Specific inclusions and exclusions to be listed. Contact details of manufacturer, supplier and local service agent to be provided.
40 | Maintenance tasks | Routine check of probe and device mechanical integrity recommended (refer to user manual). List shall be provided of equipment and procedures required for routine inspection. Advanced maintenance tasks required are not recommended to be performed in the hospital due to safety concern.
41 | Type of service contract | Costs and types of post-warranty service contract available shall be described (if applicable).
42 | Spare parts availability post-warranty | 8 years minimum, starting from installation/commisssioning
43 | Software / Hardware upgrade availability | N.A.
Documentation
44 | Documentation requirements | • Instructions for use and service manuals to be provided (including procedures for decontamination) • User language preference prioritized, English is mandatory. • Contact details of manufacturer, supplier and local service agent. • Certificate of calibration and inspection to be provided (if applicable). • List to be provided of equipment and procedures required for local calibration and routine maintenance. • List to be provided of common spares and accessories, with part numbers.
Decommissioning
45 | Estimated Life Span | 10 years
Safety and standards
46 | Standards, for the manufacturer and the equipment | US FDA: Device Class 2 EU: Class IIa
47 | Regulatory Approval / Certification | Compliance to (where applicable, but not limited to): - National Regulatory Agency/Authority (NRA) requirements compliance - Approval by regulatory body of country of manufacturer (if applicable) And at least one of: - FDA 510k clearance (US FDA) - CE mark (EU), with indication of Notifying Body (when applicable) - Other regulatory body in an IMDRF founding member country such as Australia, Canada, or Japan.
48 | International standards (for reference) | Compliant with active version of the following standards (or equivalent): General manufacturing: • ISO 13485: Medical Devices - Quality Management Systems - Requirements for Regulatory Purposes • ISO 14971: Medical Devices - Application of Risk Management to Medical Devices • ISO 15223-1: Medical devices - Symbols to be used with medical device labels, labelling and information to be supplied - Part 1: General requirements Safety and product standards for electrically-cooled systems: • IEC 60601-1 - Medical electrical equipment - Part 1: General requirements for basic safety and essential performance • IEC 60601-1-2: Medical electrical equipment - Part 1-2 General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests Safety and product standards for gas-based systems: • ISO 21969 High-pressure ﬂexible connections for use with medical gas systems Probe-specific requirements: • ISO 10993-1: Biological evaluation of medical devices Part 1: Evaluation and testing within a risk management process • ISO 10993-5: Biological evaluation of medical devices Part 5: Tests for in vitro cytotoxicity • ISO 10993-10: Biological evaluation of medical devices Part 10: Tests for irritation and skin sensitization
49 | Regional/local standards | Country-specific and regional standards may apply
50 | Regulations | US regulations: 21 CFR part 820 Quality System Regulation EU regulations: European Commission Regulation (EU) No. 2017/745 (replacing 93/42/EEC)