## Simulator newborn complete care
Simulator, newborn complete care
i | Version No.
ii | Date of initial version
iii | Date of last modification | October 2023
iv | Date of publication
v | Completed / submitted by | WHO working group
Name, category and coding
1 | WHO Category / Code | Under development
2 | Generic name | Simulator, newborn complete care
3 | Specific type or variation (optional) | Simulator, newborn complete care, light complexion. Simulator, newborn complete care, dark complexion
Purpose of use
4 | Clinical or other purpose | Newborn simulator for complete newborn care, including resuscitation and cord cutting.
5 | Level of use (if relevant) | General outpatient and outreach services for primary care; First-level (district) hospital services; Second-level and third-level hospital services and specialized outpatient services
6 | Clinical department/ ward(if relevant) | Outpatient care, Inpatient care, Labor and delivery services.
7 | Overview of functional requirements | Full term anatomically accurate newborn manikin with umbilical cord. Realistic mouth and nostrils.
Technical characteristics
10 | Detailed requirements | Light Source: The simulator should have a light source that emits the required wavelength for phototherapy and neonatal care simulation. Intensity Control: Adjustable light intensity settings to simulate different treatment scenarios and levels of neonatal jaundice. Simulated Conditions: The simulator should be able to mimic various neonatal conditions, including different levels of jaundice, skin colors, and other relevant visual indicators. Adjustable Light Area: The ability to control the size and coverage area of the light source to simulate specific treatment or care scenarios. Control Panel: An easy-to-use control panel that allows instructors or users to adjust light intensity and settings as needed. Realistic Appearance: The simulator should have a realistic appearance to create a lifelike training environment, including a model of a newborn with the necessary anatomical features. Safety Features: Safety mechanisms to ensure that the simulator does not pose risks to users or the simulated "patient." Portability: If used for training and simulation purposes, portability is essential for ease of setup and relocation.
11 | Displayed Alarms | N/A
12 | User adjustable settings | N/A
Physical/chemical characteristics
13 | Components(if relevant) | Realistic body color, head tilt/chin lift functionality to open the airway and the texture of the manikin should be as close to the feel of the baby/ adult skin as relevant. Complete unit to be easily washable with mild soap and water without bringing deterioration in the manikin, non-reactive to manufacturer-specified disinfectants/alcohol wipes.
14 | Mobility, portability(if relevant) | Portable, ideally realistic weight and size of a newborn baby
15 | Raw Materials(if relevant) | Latex-free manikins - ideally made of silicone rubber, if not, made of plastics. Hinges or joints if any should be of noncorrodible and durable metals/plastics or silicone rubber. Lubrication for joints or hinges if any should be of non-damaging substances and safe to use chemicals
Utility requirements
16 | Electrical, water and/or gas supply (if relevant) | Relevant fluid specifications for inflatable manikins (fluid type, quantity, temperature, etc) Optional: Electrical connections as recommended by the manufacturer of manikins for alarms, simulation devices and software
Accessories, consumables, spare parts, other components
17 | Accessories (if relevant) | Sheets for drying and wrapping a newborn. Head cap. Umbilical cord, which can be cut and attached several times for demonstration purposes. Umbilical ties. One suction device, portable, hand-held, manual to clear a newborn’s airway. Translucent, made of silicone or other similar material, which is easy to clean and can withstand sterilization. At least self-inflating neonatal resuscitation bag with translucent masks. At least shoulder bag for transport/storage.
18 | Sterilization process for accessories (if relevant) | Easily disinfectable by either rubbing vigorously with alcohol wipes or by use of chemical agents, as recommended by the manufacturer and in compliance with regional/organizational regulatory protocols for manikin disinfection. The product should be cleaned by soap and water prior to any disinfection methods.
19 | Consumables / reagents (if relevant) | Disposable airway systems and lungs, alcohol wipes/disinfectants for cleaning the manikin, lubricants (if any), puncture repair patches (for inflatable manikins).
20 | Spare parts (if relevant) | Detachable faces, airway systems
21 | Other components (if relevant) | Sheets for drying and wrapping a newborn. A head cap. Umbilical cord, that can be cut and attached several times for demonstration purposes. Umbilical ties. One suction device, portable, hand-held, manual to clear a newborn’s airway. Translucent, made of silicone or other similar material, which is easy to clean and can withstand sterilization. At least self-inflating neonatal resuscitation bag with translucent masks. At least shoulder bag for transport/storage.
Packaging
22 | Sterility status on delivery (if relevant) | N/A
23 | Shelf life (if relevant) | Shelf life ofat least 5 years on storage at 0˚C to +40˚C, relative humidity of 15 to 95%
24 | Transportation and storage (if relevant) | Capable of being transported/stored continuously in an ambient temperature of at least -30˚C to +60˚C and relative humidity of 15 to 95%. Capable of operating continuously in an ambient temperature of at least 0˚C to +40˚C and relative humidity of 15 to 95%
25 | Labelling (if relevant) | Labeling on the primary packaging to include: Name and/or trademark of the manufacturer. Model or product reference. Information for particular storage conditions (temperature, pressure, light, humidity). Manufacturer’s product reference Manufacture and expiry date, lot number Main characteristics (dimensions, capacity) Recommended storage conditions Information for handling
Environmental requirements
26 | Context-dependent requirements | Capable of being stored continuously in an ambient temperature of at least 0–50 °C and relative humidity of 15–85%, preferably 90%. Capable of operating continuously in an ambient temperature of at least 10–40 °C and relative humidity of 15–85%, preferably 90%.
Training, installation and utilization
27 | Pre-installation requirements(if relevant) | Not required.
28 | Requirements for commissioning (if relevant) | N/A
29 | Training of user/s (if relevant) | Source of useful information; Training of users in handling, disinfection, and basic maintenance should be provided: WHO Guidelines on basic newborn resuscitation https://apps.who.int/iris/handle/10665/75157 WHO technical specifications of neonatal resuscitation devices. https://apps.who.int/iris/handle/10665/206540 ASTM F1453 - 92 Standard Guide for Training and Evaluation of First Responders Who Provide Emergency Medical Care (USA). https://www.astm.org/Standards/F1453.htm
30 | User care(if relevant) | User manual for handling, instructions for disinfection and basic maintenance should be provided.
Warranty and maintenance
31 | Warranty | At least 2 years from date of purchase.
32 | Maintenance tasks | A list shall be provided of procedures required for routine maintenance.
33 | Type of service contract | N/A
34 | Spare parts availability post-warranty | 5 years from the date of purchase or more as per the manufacturer's offer. List of spares/ accessories, relevant part numbers, MOQ, and price list to be submitted along with the quotation.
35 | Software / Hardware upgrade availability | N/A
Documentation
36 | Documentation requirements | The instruction manual should include the Standard Operating Procedure for using the product, day-to-day maintenance, a maintenance schedule, storage condition, cleaning, troubleshooting, safety details, warranty details with a scope, and postal, electronic, and telephone contacts of authorized dealer/country representative as appropriate. Documentation must also show recommended procedures for disposal and any probable hazards to the environment and/or community.
Decommissioning
37 | Estimated Life Span | At least 5 years.
Safety and standards
38 | Standards, for the manufacturer and the equipment | N/A
39 | Regulatory Approval / Certification | Free sales certificate (FSC). Certificate for exportation of medical device provided by the authority in the manufacturing country. Proof of regulatory compliance, as appropriate, per the product’s risk classification (e.g. Food and Drug Administration [FDA] and/or Conformité Européenne [CE]).
40 | International standards | In the absence of a product standard for manikins, relevant International standards for testing of raw material and/or the finished product (or other internationally recognized or national equivalent standards) may be used to qualify the manikins. The qualification approach should be as per base guidelines given in section 2.4 of this document. Silicone rubber: The proof of raw material qualification certificate from a nationally or Internationally approved laboratory. Some standards that are suggested are: Plastics: ISO 11403-3: Plastics -- Acquisition and presentation of comparable multipoint data - Part 3: Environmental influences on properties. https://www.iso.org/standard/55462.html Weathering of adhesive joints: ISO 14615: Durability of structural adhesive joints - Exposure to humidity and temperature under load https://www.iso.org/standard/24971.html General compliance: ISO/IEC 17050-1: Conformity assessment. Supplier’s declaration of conformity. General requirements and regional regulatory guidelines should be complied with for waste disposal and recovery management. https://www.iso.org/standard/29373.html